Bend 1.26D103.3T

SN 16101

LCLS2 Barcode 4509

Alignment Engineering Group

Sept.26, 2016

TB4 & D

TB3 & C

Beam Line

TB1 & A

TB2 & B

Tooling ball values 1 to 4 to center of 1.0000 inch tooling ball, A to D to center of 0.3125 inch tooling ball, all units are inches.

| Tooling Ball | Z (in) | X (in) | Y (in) |
| --- | --- | --- | --- |
| TB1 | -51.2902 | -4.6818 | 5.3989 |
| TB2 | -51.292 | 4.6632 | 5.4082 |
| TB3 | 51.2865 | 4.6724 | 5.4056 |
| TB4 | 51.2884 | -4.6688 | 5.4044 |
| TBA | -51.2886 | -4.6797 | 4.7126 |
| TBB | -51.2878 | 4.6647 | 4.7195 |
| TBC | 51.2852 | 4.6723 | 4.7178 |
| TBD | 51.2881 | -4.6696 | 4.7164 |

-Constructed 6 planes,  top pole, bottom pole, +X side top and bottom pole, -X side top and bottom pole, upstream end of steel and downstream end of steel.

-Bisected two Y planes for Y = 0 plane , two X planes for X = 0 plane, and two Z planes for Z = 0 plane.

-Origin is the intersection of three planes.  Used the Y = 0 as primary plane, and X = 0 plane to clock the yaw angle.

-Measured both 1.0000 inch and 0.3125 inch tooling balls.

-Used AT401 for overall control.  Used 9 foot Edge Arm on each end of magnet tied to tracker control to measure upstream and downstream poles.

-Average distance from the top and bottom pole planes or a total gap of 1.2631”.