## Slide 1
This magnet was standardized to 380 A
during measurement. The installed power
supply is only capable of 375 A. The original
380 A standardize polynomial will be retained,
and the operational limits in EPICS will be
set to a maximum current of 375 A.
Note: the LCLS Control System requires all dipole magnets powered in series to have positive IVB polynomials (i.e. BDES[kG-m] > 0 gives IDES[A] > 0). To that end, this dipole’s IVB (which is properly negative) has had its sign reversed both in Oracle and in EPICS.

## Slide 2
engineering type: 1.69D6.28T
SLAC drawing number: SA-388-320-05
SLAC magnet bar-code: 4512
vendor serial number: 2
when measured: 23-Jun-2017 08:32:47 (23-Aug-2018 10:58:50)
when analyzed: 25-Aug-2022
MAD assignment: BCX12
BCX12: BL vs X @ Imain = 250 A
without vacuum chamber

## Slide 3
Without vacuum chamber
(wirevsx.r13)
p0 = 1.5267e-02
p1 = -7.0609e-04
p2 = -1.5153e-04
p3 = -2.5169e-07
p4 = 1.2125e-07
@ r= 40 mm
|b1/b0| = 0.028 [0.015] %
|b2/b0| = 0.242 [0.097] %
|b3/b0| = 0.016 [1.200] %
With vacuum chamber
(wirevsx.r14)
p0 = 1.3864e-03
p1 = 3.6025e-04
p2 = -1.9058e-04
p3 = -4.4282e-07
p4 = 5.3158e-08
@ r= 40 mm
|b1/b0| = 0.014 [0.015] %
|b2/b0| = 0.305 [0.097] %
|b3/b0| = 0.028 [1.200] %
Imain = 250 A
Itrim = 0
Imain | b1/b0 % | b2/b0 % | b3/b0 %
249.98 | 0.0269 | 0.1321 | 0.0301
380.00 | 0.0212 | 0.1284 | 0.0252
tolerance | 0.0150 | 0.0970 | 1.2000
Harmonic content vs Imain @ r = 4 cm
hardat.r32 (rotating coil)

## Slide 4
name: BCX12T
# of main coil turns: 30
# of trim coil turns: 15
ratio of turns (Nmain/Ntrim): 2.0000
ratio of slopes (measured): 2.1637
IMMO (amps): [-12,+12]
BMMO (main coil amps): [-5.5461,+5.5461]
Itrim = 0
Imain < 230 A
Imain = 0
wiredat.r29
Itrim = -1.427 A
main ramped down from 380 A to 0, then trim sets zero field