## Slide 1
Note: the LCLS Control System requires all dipole magnets powered in series to have positive IVB polynomials (i.e. BDES[kG-m] > 0 gives IDES[A] > 0). To that end, this dipole’s IVB (which is properly negative) has had its sign reversed both in Oracle and in EPICS.

## Slide 2
engineering type: 0.788D11.50
SLAC drawing number: SA-237-005-02-R2
SLAC magnet bar-code: 4523
vendor serial number: 25
when measured: 04-Jun-2018 10:29:18
when analyzed: 04-Sep-2022
MAD assignment: BCX361
I = 250 A

## Slide 3
quadrupole component
artificially removed
BCX361: ΔBL/BL0 vs X @ 250 A

## Slide 4
name: BCX361T
# of main coil turns: 48
# of trim coil turns: 82
ratio of turns (Nmain/Ntrim): 0.5854
ratio of slopes (measured): 0.5954
IMMO (amps): [-3,+3]
BMMO (main coil amps): [-5.0389,+5.0389]
Itrim = 0
Imain = 0
Itrim = -0.2581 A
main ramped down from 250 A to 0, then trim sets zero field
wiredat.ru2