## Slide 1
engineering type: 0.433D14-C (1.575D14-C)
SLAC drawing number: ASM-20171113-6552 Red_Gap CHG_11mm_9_5_19.pdf
SLAC magnet bar-code: 4551
vendor serial number: 2
when measured: 07-Oct-2019 17:28:05
when analyzed: 10-Oct-2019
MAD assignment: BCXXL2
BL (after stdz) = 18.14 G

## Slide 2
after standardize
Imain = 0
slope @ Itrim = 0 : -4.1387 G-m/A
after degauss
Imain = 0
slope @ Itrim = 0 : -3.8212 G-m/A

## Slide 3
name: BCXXL2_Trim
# of main coil turns: 378
# of trim coil turns: 10
ratio of turns (Nmain/Ntrim): 37.8
ratio of slopes (measured): 40.172
IMMO (amps): [-6,6]
BMMO (main coil amps): [-0.1493,+0.1493]
Imain = 0 (after STDZ)
wiredat.ru4
Itrim = 0

## Slide 4
ramp type = COSINE
ramp rate = 2 A/s
settle time = 5 s
remnant BL = -0.1250 +/- 0.1192 G-m
degauss time = 420 s
remnant BL: wiredat.ru6