## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 4590
vendor serial number: 1
when measured: 12-Aug-2021 15:03:09
when analyzed: 19-Aug-2021
MAD assignment: BCXCBX0
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -28.86 +/- 0.12 G-m
after degauss = -0.53 +/- 0.22 G-m …
pole-tip field @ magnet center B = 0.0 +/- 0.2 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.015 [0.500] %
|b2/b0| = 0.053 [2.000] %
Rotating-coil (r= 3 mm):
|b1/b0| = 0.0074 [0.500] %
|b2/b0| = 0.0086 [2.000] %
Harmonics @ r= 10 mm:
|b1/b0| = 0.003 [0.500] %
|b2/b0| = 0.063 [2.000] %
Rotating-coil (r= 3 mm):
|b1/b0| = 0.0035 [0.500] %
|b2/b0| = 0.0071 [2.000] %

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX0T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6709
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.2998,+0.2998]

## Slide 4
Imain = 0 (after degauss)
wiredat.ru7
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)

## Slide 5
bhvszdat.ru9