| Date: | 04/03/2017 | Number: | NCR- | 33594-0 |  |  | Page: | 1 of 1 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |

(Job#-Leg#-xxx, assigned by QA or Mfg Eng Job Lead)

| Originator: | Keith Caban | Company or SLAC Shop: | Swiss Screw |
| --- | --- | --- | --- |
| Phone: | X4022 | PO Number | 33594 |
| E-mail: | kcaban@slac.stanford.edu |  |  |

| Part Name: | Tooling Ball Socket | Quantity Impacted: | 500 |  |
| --- | --- | --- | --- | --- |
| Part Number: | PF-444-316-76 R1 | Serial No. (if applicable): |  |  |
| Additional Notes: | Contact : John Amann & Bill Misson |  |  |  |

| ITEM NO. | NONCONFORMING CHARACTERISTIC | DISPOSITION | REINSPECTION |  |
| --- | --- | --- | --- | --- |
| 1 | Magnetic Permeability not to exceed 1.05 @ 200 oersteds (All 8 AQL sampled all measured between 1.15 - 1.19) | USE AS-IS REWORK REJECT RETURN TO VENDOR | YES NO |  |
| 2 |  | USE AS-IS REWORK REJECT RETURN TO VENDOR | YES NO |  |
| 3 |  | USE AS-IS REWORK REJECT RETURN TO VENDOR | YES NO |  |

Use continuation sheet if necessary.  Attach red-lines, photos as necessary

| AGREEMENT: | CONCURRENCE: |  |  |
| --- | --- | --- | --- |
| Technical Lead Date | Quality Assurance Rep. Date | CAM or Project Manager | Director (High Risk NCRs Only)* |

* Risk based definition as set forth by Project’s Quality Assurance Plan; see SLAC-I-020-701-001,  NCR procedure.

Indicate all follow-on or related processes:	NONE	CAR	ECR	HOLD