{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":20,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":390667},null,{"$f":"'DF raw data'!G7","v":491273},null,{"$f":"'DF raw data'!L7","v":591603},null,{"$f":"'DF raw data'!Q7","v":791856},null,{"$f":"'DF raw data'!V7","v":2191950},null,null,{"$f":"B4","v":10},{"$f":"B21","v":0.40847330000000009},{"$f":"C21","v":-6.3525133633333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":13.979193400000002},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":11.560233356666668}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-13.570720100000001},{"$f":"E21","v":-4.9834300666666671},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":13.979193400000002},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":11.560233356666668}]
[-10,{"$f":"'DF raw data'!B9","v":-10.40813},{"$f":"'DF raw data'!C9","v":22.615561},{"$f":"'DF raw data'!G9","v":-20.547198999999999},{"$f":"'DF raw data'!H9","v":15.721500000000001},{"$f":"'DF raw data'!L9","v":-12.45412},{"$f":"'DF raw data'!M9","v":18.966721},{"$f":"'DF raw data'!Q9","v":-13.10886},{"$f":"'DF raw data'!R9","v":12.82024},{"$f":"'DF raw data'!V9","v":24.505019999999998},{"$f":"'DF raw data'!W9","v":47.107219999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":-2.1970434033333333},{"$f":"G21","v":-3.2623400666666669}]
[-8,{"$f":"'DF raw data'!B10","v":-16.926888999999999},{"$f":"'DF raw data'!C10","v":11.92151},{"$f":"'DF raw data'!G10","v":-11.1257},{"$f":"'DF raw data'!H10","v":6.5561999999999996},{"$f":"'DF raw data'!L10","v":-0.75678003000000005},{"$f":"'DF raw data'!M10","v":29.377399},{"$f":"'DF raw data'!Q10","v":-17.3123},{"$f":"'DF raw data'!R10","v":8.1581202000000008},{"$f":"'DF raw data'!V10","v":22.197889},{"$f":"'DF raw data'!W10","v":27.062380000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-3.3058099666666667},{"$f":"I21","v":5.2077199933333338},null,"Limits 10-30","+/-7.5","+/-7.5"]
[-6,{"$f":"'DF raw data'!B11","v":-3.4418600000000001},{"$f":"'DF raw data'!C11","v":11.724119999999999},{"$f":"'DF raw data'!G11","v":-0.33118001000000002},{"$f":"'DF raw data'!H11","v":15.324299999999999},{"$f":"'DF raw data'!L11","v":-16.654890000000002},{"$f":"'DF raw data'!M11","v":24.740860000000001},{"$f":"'DF raw data'!Q11","v":-4.1212400999999996},{"$f":"'DF raw data'!R11","v":20.436710000000001},{"$f":"'DF raw data'!V11","v":2.96984},{"$f":"'DF raw data'!W11","v":18.416018999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":-7.9413666333333337},{"$f":"K21","v":3.5549433333333327}]
[-4,{"$f":"'DF raw data'!B12","v":-3.5786199999999999},{"$f":"'DF raw data'!C12","v":-7.8539997E-2},{"$f":"'DF raw data'!G12","v":-3.1667700000000001},{"$f":"'DF raw data'!H12","v":3.0032301000000001},{"$f":"'DF raw data'!L12","v":-0.17577999999999999},{"$f":"'DF raw data'!M12","v":-11.795019999999999},{"$f":"'DF raw data'!Q12","v":4.2343402000000001},{"$f":"'DF raw data'!R12","v":-2.1530499000000001},{"$f":"'DF raw data'!V12","v":-4.2705798000000001},{"$f":"'DF raw data'!W12","v":4.1683101999999996}]
[-2,{"$f":"'DF raw data'!B13","v":-1.2199899999999999},{"$f":"'DF raw data'!C13","v":-0.72435998999999995},{"$f":"'DF raw data'!G13","v":-9.9345703000000007},{"$f":"'DF raw data'!H13","v":-4.0702400000000001},{"$f":"'DF raw data'!L13","v":-6.9236101999999997},{"$f":"'DF raw data'!M13","v":6.1471099999999996},{"$f":"'DF raw data'!Q13","v":-6.7789402000000001},{"$f":"'DF raw data'!R13","v":12.307399999999999},{"$f":"'DF raw data'!V13","v":-3.5833599999999999},{"$f":"'DF raw data'!W13","v":8.2629900000000003},null,"Limits","+/-17.5","+/-17.5"]
[0,{"$f":"'DF raw data'!B14","v":3.6421299},{"$f":"'DF raw data'!C14","v":-11.27694},{"$f":"'DF raw data'!G14","v":-10.04942},{"$f":"'DF raw data'!H14","v":-2.5625501000000002},{"$f":"'DF raw data'!L14","v":-0.11144},{"$f":"'DF raw data'!M14","v":-4.4263902000000002},{"$f":"'DF raw data'!Q14","v":4.7518000999999996},{"$f":"'DF raw data'!R14","v":2.5121000000000002},{"$f":"'DF raw data'!V14","v":-4.9184799000000003},{"$f":"'DF raw data'!W14","v":0.28033000000000002}]
[2,{"$f":"'DF raw data'!B15","v":-1.19672},{"$f":"'DF raw data'!C15","v":-7.0562401000000001},{"$f":"'DF raw data'!G15","v":-20.728169999999999},{"$f":"'DF raw data'!H15","v":-8.3175001000000002},{"$f":"'DF raw data'!L15","v":0.44391998999999999},{"$f":"'DF raw data'!M15","v":-11.50774},{"$f":"'DF raw data'!Q15","v":-7.8902897999999997},{"$f":"'DF raw data'!R15","v":0.80365998000000005},{"$f":"'DF raw data'!V15","v":-15.32226},{"$f":"'DF raw data'!W15","v":2.1215099999999998}]
[4,{"$f":"'DF raw data'!B16","v":-3.2936698999999998},{"$f":"'DF raw data'!C16","v":-7.9275197999999998},{"$f":"'DF raw data'!G16","v":4.2442197999999998},{"$f":"'DF raw data'!H16","v":3.4748499000000002},{"$f":"'DF raw data'!L16","v":-4.6249199000000001},{"$f":"'DF raw data'!M16","v":-4.7775401999999998},{"$f":"'DF raw data'!Q16","v":1.0359799999999999},{"$f":"'DF raw data'!R16","v":11.954000000000001},{"$f":"'DF raw data'!V16","v":-5.6610198},{"$f":"'DF raw data'!W16","v":5.9339399000000004}]
[6,{"$f":"'DF raw data'!B17","v":-13.29407},{"$f":"'DF raw data'!C17","v":-21.514299000000001},{"$f":"'DF raw data'!G17","v":-12.916090000000001},{"$f":"'DF raw data'!H17","v":-10.83672},{"$f":"'DF raw data'!L17","v":-8.1622400000000006},{"$f":"'DF raw data'!M17","v":-15.336130000000001},{"$f":"'DF raw data'!Q17","v":-10.105689999999999},{"$f":"'DF raw data'!R17","v":-0.83767002999999995},{"$f":"'DF raw data'!V17","v":-2.8315598999999998},{"$f":"'DF raw data'!W17","v":-0.37485998999999998}]
[8,{"$f":"'DF raw data'!B18","v":-7.6586800000000004},{"$f":"'DF raw data'!C18","v":-16.402149000000001},{"$f":"'DF raw data'!G18","v":-11.48851},{"$f":"'DF raw data'!H18","v":-12.71106},{"$f":"'DF raw data'!L18","v":-14.889430000000001},{"$f":"'DF raw data'!M18","v":3.9316800000000001},{"$f":"'DF raw data'!Q18","v":-12.34858},{"$f":"'DF raw data'!R18","v":5.4706602000000002},{"$f":"'DF raw data'!V18","v":-3.4389900999999998},{"$f":"'DF raw data'!W18","v":0.10459}]
[10,{"$f":"'DF raw data'!B19","v":-12.520820000000001},{"$f":"'DF raw data'!C19","v":-26.15579},{"$f":"'DF raw data'!G19","v":-8.9065799999999999},{"$f":"'DF raw data'!H19","v":-16.231929999999998},{"$f":"'DF raw data'!L19","v":-16.7773},{"$f":"'DF raw data'!M19","v":6.8986101},{"$f":"'DF raw data'!Q19","v":-15.303599999999999},{"$f":"'DF raw data'!R19","v":1.10812},{"$f":"'DF raw data'!V19","v":-6.9569302000000004},{"$f":"'DF raw data'!W19","v":-4.6035700000000004}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":89},null,{"$f":"'DF raw data'!G4","v":90},null,{"$f":"'DF raw data'!L4","v":92},null,{"$f":"'DF raw data'!Q4","v":93},null,{"$f":"'DF raw data'!V4","v":94}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":0.40847330000000009},{"$f":"AVERAGE(C11:C13)","v":-6.3525133633333333},-13.570720100000001,-4.9834300666666671,-2.1970434033333333,-3.2623400666666669,-3.3058099666666667,5.2077199933333338,-7.9413666333333337,3.5549433333333327]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":390667},null,{"$f":"'DF raw data'!G7","v":491273},null,{"$f":"'DF raw data'!L7","v":591603},null,{"$f":"'DF raw data'!Q7","v":791856},null,{"$f":"'DF raw data'!V7","v":2191950},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-1.0030247999999999},{"$f":"C21","v":-28.815317666666662},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":28.766172566666668},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":36.558177666666666}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-29.455866666666669},{"$f":"E21","v":-23.128001266666661},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":31.819073966666668},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":36.558177666666666}]
[-10,{"$f":"'DF raw data'!D9/100","v":-12.872961},{"$f":"'DF raw data'!E9/100","v":5.0389404000000004},{"$f":"'DF raw data'!I9/100","v":-12.338459},{"$f":"'DF raw data'!J9/100","v":-0.5947559},{"$f":"'DF raw data'!N9/100","v":-11.582898},{"$f":"'DF raw data'!O9/100","v":-4.0950635000000002},{"$f":"'DF raw data'!S9/100","v":-3.5398516999999998},{"$f":"'DF raw data'!T9/100","v":-9.0329101999999999},{"$f":"'DF raw data'!X9/100","v":-11.981812999999999},{"$f":"'DF raw data'!Y9/100","v":-12.00165},null,{"$f":"F4","v":20},{"$f":"F21","v":-4.6187458999999995},{"$f":"G21","v":-24.403287833333334}]
[-8,{"$f":"'DF raw data'!D10/100","v":-30.470127000000002},{"$f":"'DF raw data'!E10/100","v":-7.2257227000000004},{"$f":"'DF raw data'!I10/100","v":-26.729801999999999},{"$f":"'DF raw data'!J10/100","v":-20.620659},{"$f":"'DF raw data'!N10/100","v":-9.9188109999999998},{"$f":"'DF raw data'!O10/100","v":39.662692999999997},{"$f":"'DF raw data'!S10/100","v":-29.654342999999997},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":6.3503490999999999},{"$f":"'DF raw data'!Y10/100","v":-1.8446338},null,{"$f":"H4","v":30},{"$f":"H21","v":-0.68969410000000053},{"$f":"I21","v":7.7428600000000003},null,"Limits 10-30","+/-42.5","+/-42.5"]
[-6,{"$f":"'DF raw data'!D11/100","v":-5.7958660999999996},{"$f":"'DF raw data'!E11/100","v":-7.8565186000000002},{"$f":"'DF raw data'!I11/100","v":8.2660137999999996},{"$f":"'DF raw data'!J11/100","v":-0.10785159999999999},{"$f":"'DF raw data'!N11/100","v":-16.718173999999998},{"$f":"'DF raw data'!O11/100","v":15.152188000000001},{"$f":"'DF raw data'!S11/100","v":6.4366198999999993},2.20269,{"$f":"'DF raw data'!X11/100","v":3.0011669999999997},{"$f":"'DF raw data'!Y11/100","v":-6.4147265999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":2.3632072999999996},{"$f":"K21","v":-12.044088633333333}]
[-4,{"$f":"'DF raw data'!D12/100","v":-9.344154099999999},{"$f":"'DF raw data'!E12/100","v":-15.567407000000001},{"$f":"'DF raw data'!I12/100","v":-11.646913},{"$f":"'DF raw data'!J12/100","v":-23.725337},{"$f":"'DF raw data'!N12/100","v":-21.732637},{"$f":"'DF raw data'!O12/100","v":-30.458105},{"$f":"'DF raw data'!S12/100","v":1.3835480000000002},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-6.7832861000000007},{"$f":"'DF raw data'!Y12/100","v":-26.603164000000003}]
[-2,{"$f":"'DF raw data'!D13/100","v":-3.2025546000000005},{"$f":"'DF raw data'!E13/100","v":-18.489550999999999},{"$f":"'DF raw data'!I13/100","v":-17.114439999999998},{"$f":"'DF raw data'!J13/100","v":-9.0937988000000001},{"$f":"'DF raw data'!N13/100","v":-13.103218},{"$f":"'DF raw data'!O13/100","v":-3.0394385000000002},{"$f":"'DF raw data'!S13/100","v":-11.745631000000001},1.63822,{"$f":"'DF raw data'!X13/100","v":8.1464068999999988},{"$f":"'DF raw data'!Y13/100","v":4.7674950999999997},null,"Limits","+/-42.5","+/-42.5"]
[0,{"$f":"'DF raw data'!D14/100","v":6.4104199000000008},{"$f":"'DF raw data'!E14/100","v":-45.721176999999997},{"$f":"'DF raw data'!I14/100","v":-20.287520000000001},{"$f":"'DF raw data'!J14/100","v":-23.175614999999997},{"$f":"'DF raw data'!N14/100","v":-4.3039331000000001},{"$f":"'DF raw data'!O14/100","v":-40.272680999999999},{"$f":"'DF raw data'!S14/100","v":15.597130999999999},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":14.165804000000001},{"$f":"'DF raw data'!Y14/100","v":-18.345497999999999}]
[2,{"$f":"'DF raw data'!D15/100","v":-6.2169397000000002},{"$f":"'DF raw data'!E15/100","v":-22.235225},{"$f":"'DF raw data'!I15/100","v":-50.96564},{"$f":"'DF raw data'!J15/100","v":-37.11459},{"$f":"'DF raw data'!N15/100","v":3.5509134000000002},{"$f":"'DF raw data'!O15/100","v":-29.897743999999999},{"$f":"'DF raw data'!S15/100","v":-5.9205822999999995},16.2971,{"$f":"'DF raw data'!X15/100","v":-15.222589000000001},{"$f":"'DF raw data'!Y15/100","v":-22.554263000000002}]
[4,{"$f":"'DF raw data'!D16/100","v":-8.3029956000000009},{"$f":"'DF raw data'!E16/100","v":-37.105185999999996},{"$f":"'DF raw data'!I16/100","v":20.188856000000001},{"$f":"'DF raw data'!J16/100","v":-9.0064209000000002},{"$f":"'DF raw data'!N16/100","v":-18.072212999999998},{"$f":"'DF raw data'!O16/100","v":-26.643350000000002},{"$f":"'DF raw data'!S16/100","v":19.91611},10.6371,{"$f":"'DF raw data'!X16/100","v":12.71269},{"$f":"'DF raw data'!Y16/100","v":-10.894971}]
[6,{"$f":"'DF raw data'!D17/100","v":-6.1244323999999999},{"$f":"'DF raw data'!E17/100","v":-56.149009},{"$f":"'DF raw data'!I17/100","v":-19.100403},{"$f":"'DF raw data'!J17/100","v":-36.728847999999999},{"$f":"'DF raw data'!N17/100","v":-13.533798999999998},{"$f":"'DF raw data'!O17/100","v":-61.135863999999998},{"$f":"'DF raw data'!S17/100","v":-13.793089999999999},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":14.045864999999999},{"$f":"'DF raw data'!Y17/100","v":-24.995360999999999}]
[8,{"$f":"'DF raw data'!D18/100","v":-12.726510000000001},{"$f":"'DF raw data'!E18/100","v":-54.597344},{"$f":"'DF raw data'!I18/100","v":-22.322727},{"$f":"'DF raw data'!J18/100","v":-37.612189999999998},{"$f":"'DF raw data'!N18/100","v":-9.5279607999999989},{"$f":"'DF raw data'!O18/100","v":-19.733898999999997},{"$f":"'DF raw data'!S18/100","v":-15.392021},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":8.8761022999999994},{"$f":"'DF raw data'!Y18/100","v":-30.732833999999997}]
[10,{"$f":"'DF raw data'!D19/100","v":-24.896422999999999},{"$f":"'DF raw data'!E19/100","v":-68.608174000000005},{"$f":"'DF raw data'!I19/100","v":-9.4666625999999994},{"$f":"'DF raw data'!J19/100","v":-66.799306999999999},{"$f":"'DF raw data'!N19/100","v":-30.394940999999999},{"$f":"'DF raw data'!O19/100","v":-14.670592999999998},{"$f":"'DF raw data'!S19/100","v":-27.344050000000003},13.3657,{"$f":"'DF raw data'!X19/100","v":1.2590024},{"$f":"'DF raw data'!Y19/100","v":-39.305880999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":89},null,{"$f":"'DF raw data'!G4","v":90},null,{"$f":"'DF raw data'!L4","v":92},null,{"$f":"'DF raw data'!Q4","v":93},null,{"$f":"'DF raw data'!V4","v":94}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-1.0030247999999999},{"$f":"AVERAGE(C11:C13)","v":-28.815317666666662},-29.455866666666669,-23.128001266666661,-4.6187458999999995,-24.403287833333334,-0.68969410000000053,7.7428600000000003,2.3632072999999996,-12.044088633333333]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16334 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16334 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16334 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16334 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16334 100 mm gap"]
["Date:","6. feb 2017","09:34:25",null,null,"Date:","6. feb 2017","09:54:16",null,null,"Date:","6. feb 2017","10:22:58",null,null,"Date:","6. feb 2017","10:37:27",null,null,"Date:","6. feb 2017","11:11:16"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",89,null,null,null,"Indice:",90,null,null,null,"Indice:",92,null,null,null,"Indice:",93,null,null,null,"Indice:",94]
["Background Indice:",89,null,null,null,"Background Indice:",90,null,null,null,"Background Indice:",92,null,null,null,"Background Indice:",93,null,null,null,"Background Indice:",94]
["Temperature:",20.72,null,null,null,"Temperature:",20.7,null,null,null,"Temperature:",20.74,null,null,null,"Temperature:",20.76,null,null,null,"Temperature:",20.78]
["Encoder:",390667,null,null,null,"Encoder:",491273,null,null,null,"Encoder:",591603,null,null,null,"Encoder:",791856,null,null,null,"Encoder:",2191950]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-10.40813,22.615561,-1287.2961,503.89404000000002,-1,-20.547198999999999,15.721500000000001,-1233.8459,-59.475589999999997,-1,-12.45412,18.966721,-1158.2898,-409.50635,-1,-13.10886,12.82024,-353.98516999999998,-903.29102,-1,24.505019999999998,47.107219999999998,-1198.1813,-1200.165]
[-0.80000000999999998,-16.926888999999999,11.92151,-3047.0127000000002,-722.57227,-0.80000000999999998,-11.1257,6.5561999999999996,-2672.9802,-2062.0659000000001,-0.80000000999999998,-0.75678003000000005,29.377399,-991.88109999999995,3966.2692999999999,-0.80000000999999998,-17.3123,8.1581202000000008,-2965.4342999999999,-416.61475000000002,-0.80000000999999998,22.197889,27.062380000000001,635.03490999999997,-184.46338]
[-0.60000001999999997,-3.4418600000000001,11.724119999999999,-579.58660999999995,-785.65186000000006,-0.60000001999999997,-0.33118001000000002,15.324299999999999,826.60137999999995,-10.785159999999999,-0.60000001999999997,-16.654890000000002,24.740860000000001,-1671.8173999999999,1515.2188000000001,-0.60000001999999997,-4.1212400999999996,20.436710000000001,643.66198999999995,864.95410000000004,-0.60000001999999997,2.96984,18.416018999999999,300.11669999999998,-641.47266000000002]
[-0.40000001000000002,-3.5786199999999999,-7.8539997E-2,-934.41540999999995,-1556.7407000000001,-0.40000001000000002,-3.1667700000000001,3.0032301000000001,-1164.6913,-2372.5337,-0.40000001000000002,-0.17577999999999999,-11.795019999999999,-2173.2637,-3045.8105,-0.40000001000000002,4.2343402000000001,-2.1530499000000001,138.35480000000001,-3132.5342000000001,-0.40000001000000002,-4.2705798000000001,4.1683101999999996,-678.32861000000003,-2660.3164000000002]
[-0.2,-1.2199899999999999,-0.72435998999999995,-320.25546000000003,-1848.9550999999999,-0.2,-9.9345703000000007,-4.0702400000000001,-1711.444,-909.37987999999996,-0.2,-6.9236101999999997,6.1471099999999996,-1310.3217999999999,-303.94385,-0.2,-6.7789402000000001,12.307399999999999,-1174.5631000000001,724.59424000000001,-0.2,-3.5833599999999999,8.2629900000000003,814.64068999999995,476.74950999999999]
[0,3.6421299,-11.27694,641.04199000000006,-4572.1176999999998,0,-10.04942,-2.5625501000000002,-2028.752,-2317.5614999999998,0,-0.11144,-4.4263902000000002,-430.39330999999999,-4027.2680999999998,0,4.7518000999999996,2.5121000000000002,1559.7130999999999,-3112.1435999999999,0,-4.9184799000000003,0.28033000000000002,1416.5804000000001,-1834.5498]
[0.2,-1.19672,-7.0562401000000001,-621.69397000000004,-2223.5225,0.2,-20.728169999999999,-8.3175001000000002,-5096.5640000000003,-3711.4589999999998,0.2,0.44391998999999999,-11.50774,355.09134,-2989.7743999999998,0.2,-7.8902897999999997,0.80365998000000005,-592.05822999999998,-1990.3032000000001,0.2,-15.32226,2.1215099999999998,-1522.2589,-2255.4263000000001]
[0.40000001000000002,-3.2936698999999998,-7.9275197999999998,-830.29956000000004,-3710.5185999999999,0.40000001000000002,4.2442197999999998,3.4748499000000002,2018.8856000000001,-900.64209000000005,0.40000001000000002,-4.6249199000000001,-4.7775401999999998,-1807.2212999999999,-2664.335,0.40000001000000002,1.0359799999999999,11.954000000000001,1991.6110000000001,-1328.4848999999999,0.40000001000000002,-5.6610198,5.9339399000000004,1271.269,-1089.4971]
[0.60000001999999997,-13.29407,-21.514299000000001,-612.44323999999995,-5614.9008999999996,0.60000001999999997,-12.916090000000001,-10.83672,-1910.0402999999999,-3672.8847999999998,0.60000001999999997,-8.1622400000000006,-15.336130000000001,-1353.3798999999999,-6113.5864000000001,0.60000001999999997,-10.105689999999999,-0.83767002999999995,-1379.309,-1919.5492999999999,0.60000001999999997,-2.8315598999999998,-0.37485998999999998,1404.5864999999999,-2499.5360999999998]
[0.80000000999999998,-7.6586800000000004,-16.402149000000001,-1272.6510000000001,-5459.7344000000003,0.80000000999999998,-11.48851,-12.71106,-2232.2727,-3761.2190000000001,0.80000000999999998,-14.889430000000001,3.9316800000000001,-952.79607999999996,-1973.3898999999999,0.80000000999999998,-12.34858,5.4706602000000002,-1539.2021,-1606.4983,0.80000000999999998,-3.4389900999999998,0.10459,887.61023,-3073.2833999999998]
[1,-12.520820000000001,-26.15579,-2489.6423,-6860.8173999999999,1,-8.9065799999999999,-16.231929999999998,-946.66625999999997,-6679.9306999999999,1,-16.7773,6.8986101,-3039.4940999999999,-1467.0592999999999,1,-15.303599999999999,1.10812,-2734.4050000000002,-3563.8087999999998,1,-6.9569302000000004,-4.6035700000000004,125.90024,-3930.5880999999999]