{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":36,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":393685},null,{"$f":"'DF raw data'!G7","v":494274},null,{"$f":"'DF raw data'!L7","v":594599},null,{"$f":"'DF raw data'!Q7","v":794907},null,{"$f":"'DF raw data'!V7","v":2195000},null,null,{"$f":"B4","v":10},{"$f":"B21","v":3.5508000333333336},{"$f":"C21","v":-5.5949732999999995},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":4.6938364999999997},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":3.268833299666666}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":5.9735933666666661},{"$f":"E21","v":-4.8358766433333331},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":4.6938364999999997},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":8.4875800009999995}]
[-10,{"$f":"'DF raw data'!B9","v":20.665130999999999},{"$f":"'DF raw data'!C9","v":10.398350000000001},{"$f":"'DF raw data'!G9","v":20.894439999999999},{"$f":"'DF raw data'!H9","v":7.1504697999999998},{"$f":"'DF raw data'!L9","v":11.64385},{"$f":"'DF raw data'!M9","v":1.8824799999999999},{"$f":"'DF raw data'!Q9","v":12.586980000000001},{"$f":"'DF raw data'!R9","v":9.8557997000000004},{"$f":"'DF raw data'!V9","v":5.9211998000000001},{"$f":"'DF raw data'!W9","v":-4.5762299999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":8.2446365333333329},{"$f":"G21","v":-5.2791767333333333}]
[-8,{"$f":"'DF raw data'!B10","v":1.0530999999999999},{"$f":"'DF raw data'!C10","v":8.2679796000000003},{"$f":"'DF raw data'!G10","v":10.840260000000001},{"$f":"'DF raw data'!H10","v":4.7358899000000001},{"$f":"'DF raw data'!L10","v":15.20275},{"$f":"'DF raw data'!M10","v":6.9724202000000002},{"$f":"'DF raw data'!Q10","v":15.8843},{"$f":"'DF raw data'!R10","v":4.0632099999999998},{"$f":"'DF raw data'!V10","v":8.6791295999999996},{"$f":"'DF raw data'!W10","v":0.25101000000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":7.8172566333333338},{"$f":"I21","v":-2.3261400003333335},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":12.976610000000001},{"$f":"'DF raw data'!C11","v":3.0506201000000002},{"$f":"'DF raw data'!G11","v":15.149279999999999},{"$f":"'DF raw data'!H11","v":1.7048099999999999},{"$f":"'DF raw data'!L11","v":9.2420501999999995},{"$f":"'DF raw data'!M11","v":-2.6835499},{"$f":"'DF raw data'!Q11","v":22.084378999999998},{"$f":"'DF raw data'!R11","v":4.5042200000000001},{"$f":"'DF raw data'!V11","v":18.534969},{"$f":"'DF raw data'!W11","v":-3.5035501},null,{"$f":"J4","v":100},{"$f":"J21","v":5.6237800333333334},{"$f":"K21","v":2.8926067009999996},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-3.7654399999999999},{"$f":"'DF raw data'!C12","v":-0.35185},{"$f":"'DF raw data'!G12","v":8.4055900999999995},{"$f":"'DF raw data'!H12","v":4.1184602000000003},{"$f":"'DF raw data'!L12","v":10.655110000000001},{"$f":"'DF raw data'!M12","v":4.3629699000000004},{"$f":"'DF raw data'!Q12","v":16.464680000000001},{"$f":"'DF raw data'!R12","v":7.8439598000000004},{"$f":"'DF raw data'!V12","v":11.39208},{"$f":"'DF raw data'!W12","v":5.4384499000000002},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-0.29049000000000003},{"$f":"'DF raw data'!C13","v":-10.562340000000001},{"$f":"'DF raw data'!G13","v":7.28409},{"$f":"'DF raw data'!H13","v":-6.6057601000000004},{"$f":"'DF raw data'!L13","v":4.7923397999999997},{"$f":"'DF raw data'!M13","v":-3.0118699000000002},{"$f":"'DF raw data'!Q13","v":7.6083999000000002},{"$f":"'DF raw data'!R13","v":6.4539999000000001E-2},{"$f":"'DF raw data'!V13","v":1.58684},{"$f":"'DF raw data'!W13","v":9.0680002999999995E-2},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":4.5007200000000003},{"$f":"'DF raw data'!C14","v":-4.1078099999999997},{"$f":"'DF raw data'!G14","v":8.1981801999999995},{"$f":"'DF raw data'!H14","v":-0.87041002999999995},{"$f":"'DF raw data'!L14","v":9.7185898000000002},{"$f":"'DF raw data'!M14","v":-6.6921100999999998},{"$f":"'DF raw data'!Q14","v":12.128690000000001},{"$f":"'DF raw data'!R14","v":-6.8943000000000003},{"$f":"'DF raw data'!V14","v":9.4333200000000001},{"$f":"'DF raw data'!W14","v":10.73086},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":6.4421701000000002},{"$f":"'DF raw data'!C15","v":-2.1147699000000002},{"$f":"'DF raw data'!G15","v":2.4385099000000001},{"$f":"'DF raw data'!H15","v":-7.0314598000000004},{"$f":"'DF raw data'!L15","v":10.22298},{"$f":"'DF raw data'!M15","v":-6.1335502000000002},{"$f":"'DF raw data'!Q15","v":3.71468},{"$f":"'DF raw data'!R15","v":-0.14865999999999999},{"$f":"'DF raw data'!V15","v":5.8511800999999997},{"$f":"'DF raw data'!W15","v":-2.1437198999999998},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":7.7705102000000004},{"$f":"'DF raw data'!C16","v":-8.6805000000000003},{"$f":"'DF raw data'!G16","v":3.8210299000000001},{"$f":"'DF raw data'!H16","v":-3.6071501000000001},{"$f":"'DF raw data'!L16","v":13.58405},{"$f":"'DF raw data'!M16","v":-2.2965800999999999},{"$f":"'DF raw data'!Q16","v":12.600910000000001},{"$f":"'DF raw data'!R16","v":6.3321199000000004},{"$f":"'DF raw data'!V16","v":15.565659999999999},{"$f":"'DF raw data'!W16","v":6.8545799000000001}]
[6,{"$f":"'DF raw data'!B17","v":8.1126099000000007},{"$f":"'DF raw data'!C17","v":-10.83893},{"$f":"'DF raw data'!G17","v":1.6286499999999999},{"$f":"'DF raw data'!H17","v":-15.12337},{"$f":"'DF raw data'!L17","v":16.727409000000002},{"$f":"'DF raw data'!M17","v":-3.6581999999999999},{"$f":"'DF raw data'!Q17","v":17.127179999999999},{"$f":"'DF raw data'!R17","v":-8.5406198999999994},{"$f":"'DF raw data'!V17","v":10.01327},{"$f":"'DF raw data'!W17","v":8.8025702999999993}]
[8,{"$f":"'DF raw data'!B18","v":12.70984},{"$f":"'DF raw data'!C18","v":-11.583769999999999},{"$f":"'DF raw data'!G18","v":11.305429999999999},{"$f":"'DF raw data'!H18","v":-9.5385703999999993},{"$f":"'DF raw data'!L18","v":11.733140000000001},{"$f":"'DF raw data'!M18","v":-6.4678798000000004},{"$f":"'DF raw data'!Q18","v":11.03403},{"$f":"'DF raw data'!R18","v":-6.0744400000000001},{"$f":"'DF raw data'!V18","v":5.9071002000000004},{"$f":"'DF raw data'!W18","v":0.46665001}]
[10,{"$f":"'DF raw data'!B19","v":17.442419000000001},{"$f":"'DF raw data'!C19","v":-2.3392700999999998},{"$f":"'DF raw data'!G19","v":31.26906},{"$f":"'DF raw data'!H19","v":-0.88476997999999996},{"$f":"'DF raw data'!L19","v":12.43543},{"$f":"'DF raw data'!M19","v":5.3699002},{"$f":"'DF raw data'!Q19","v":12.80176},{"$f":"'DF raw data'!R19","v":0.82713002000000002},{"$f":"'DF raw data'!V19","v":9.4711504000000009},{"$f":"'DF raw data'!W19","v":10.10704}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":247},null,{"$f":"'DF raw data'!G4","v":248},null,{"$f":"'DF raw data'!L4","v":251},null,{"$f":"'DF raw data'!Q4","v":253},null,{"$f":"'DF raw data'!V4","v":254}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":3.5508000333333336},{"$f":"AVERAGE(C11:C13)","v":-5.5949732999999995},5.9735933666666661,-4.8358766433333331,8.2446365333333329,-5.2791767333333333,7.8172566333333338,-2.3261400003333335,5.6237800333333334,2.8926067009999996]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":393685},null,{"$f":"'DF raw data'!G7","v":494274},null,{"$f":"'DF raw data'!L7","v":594599},null,{"$f":"'DF raw data'!Q7","v":794907},null,{"$f":"'DF raw data'!V7","v":2195000},null,null,{"$f":"B4","v":10},{"$f":"B21","v":19.356675333333332},{"$f":"C21","v":2.5490820000000003},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":16.328198333333336},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":6.1732716666666665}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":25.108498333333333},{"$f":"E21","v":8.7223536666666668},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":16.328198333333336},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":7.4568196999999996}]
[-10,{"$f":"'DF raw data'!D9/100","v":45.224647999999995},{"$f":"'DF raw data'!E9/100","v":19.436738000000002},{"$f":"'DF raw data'!I9/100","v":41.697665999999998},{"$f":"'DF raw data'!J9/100","v":24.579639},{"$f":"'DF raw data'!N9/100","v":21.596195999999999},{"$f":"'DF raw data'!O9/100","v":18.474785000000001},{"$f":"'DF raw data'!S9/100","v":31.423359000000001},{"$f":"'DF raw data'!T9/100","v":17.217147999999998},{"$f":"'DF raw data'!X9/100","v":8.2407214},{"$f":"'DF raw data'!Y9/100","v":-6.4810938},null,{"$f":"F4","v":20},{"$f":"F21","v":30.039383000000001},{"$f":"G21","v":5.1583855333333331}]
[-8,{"$f":"'DF raw data'!D10/100","v":5.0268725999999999},{"$f":"'DF raw data'!E10/100","v":29.877168000000001},{"$f":"'DF raw data'!I10/100","v":28.298703999999997},{"$f":"'DF raw data'!J10/100","v":11.341171999999998},{"$f":"'DF raw data'!N10/100","v":37.220058999999999},{"$f":"'DF raw data'!O10/100","v":18.317236000000001},{"$f":"'DF raw data'!S10/100","v":39.738329999999998},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":29.615180999999996},{"$f":"'DF raw data'!Y10/100","v":0.27154301000000003},null,{"$f":"H4","v":30},{"$f":"H21","v":35.684873666666668},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":38.219814},{"$f":"'DF raw data'!E11/100","v":20.732246},{"$f":"'DF raw data'!I11/100","v":53.172274999999999},{"$f":"'DF raw data'!J11/100","v":15.910947},{"$f":"'DF raw data'!N11/100","v":34.322563000000002},{"$f":"'DF raw data'!O11/100","v":10.141104},{"$f":"'DF raw data'!S11/100","v":54.586620999999994},2.20269,{"$f":"'DF raw data'!X11/100","v":49.049579999999999},{"$f":"'DF raw data'!Y11/100","v":5.1972559},null,{"$f":"J4","v":100},{"$f":"J21","v":20.409622333333335},{"$f":"K21","v":1.2655339666666672}]
[-4,{"$f":"'DF raw data'!D12/100","v":-5.2252173000000006},{"$f":"'DF raw data'!E12/100","v":27.701132999999999},{"$f":"'DF raw data'!I12/100","v":26.091222999999999},{"$f":"'DF raw data'!J12/100","v":21.405479},{"$f":"'DF raw data'!N12/100","v":39.904302000000001},{"$f":"'DF raw data'!O12/100","v":16.311776999999999},{"$f":"'DF raw data'!S12/100","v":44.538950000000007},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":30.047719999999998},{"$f":"'DF raw data'!Y12/100","v":21.828838000000001}]
[-2,{"$f":"'DF raw data'!D13/100","v":13.134808},{"$f":"'DF raw data'!E13/100","v":-24.558145},{"$f":"'DF raw data'!I13/100","v":14.394021},{"$f":"'DF raw data'!J13/100","v":7.7324805000000003},{"$f":"'DF raw data'!N13/100","v":20.850630000000002},{"$f":"'DF raw data'!O13/100","v":-8.1141503999999998},{"$f":"'DF raw data'!S13/100","v":41.892479999999999},1.63822,{"$f":"'DF raw data'!X13/100","v":13.404283},{"$f":"'DF raw data'!Y13/100","v":-6.7974219000000007},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":22.791872999999999},{"$f":"'DF raw data'!E14/100","v":19.092188},{"$f":"'DF raw data'!I14/100","v":32.103729999999999},{"$f":"'DF raw data'!J14/100","v":13.145977},{"$f":"'DF raw data'!N14/100","v":37.500576000000002},{"$f":"'DF raw data'!O14/100","v":14.164306999999999},{"$f":"'DF raw data'!S14/100","v":47.544013999999997},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":29.476558000000001},{"$f":"'DF raw data'!Y14/100","v":15.436309000000001},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":22.143345},{"$f":"'DF raw data'!E15/100","v":13.113203},{"$f":"'DF raw data'!I15/100","v":28.827743999999999},{"$f":"'DF raw data'!J15/100","v":5.2886035000000007},{"$f":"'DF raw data'!N15/100","v":31.766943000000001},{"$f":"'DF raw data'!O15/100","v":9.4250000000000007},{"$f":"'DF raw data'!S15/100","v":17.618127000000001},16.2971,{"$f":"'DF raw data'!X15/100","v":18.348026000000001},{"$f":"'DF raw data'!Y15/100","v":-4.8422852000000001},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":21.960228999999998},{"$f":"'DF raw data'!E16/100","v":-9.8557421999999999},{"$f":"'DF raw data'!I16/100","v":18.784445999999999},{"$f":"'DF raw data'!J16/100","v":28.858906000000001},{"$f":"'DF raw data'!N16/100","v":20.628585999999999},{"$f":"'DF raw data'!O16/100","v":14.002109000000001},{"$f":"'DF raw data'!S16/100","v":23.099944000000001},10.6371,{"$f":"'DF raw data'!X16/100","v":20.786763000000001},{"$f":"'DF raw data'!Y16/100","v":24.675801}]
[6,{"$f":"'DF raw data'!D17/100","v":34.155291000000005},{"$f":"'DF raw data'!E17/100","v":2.1002930000000002},{"$f":"'DF raw data'!I17/100","v":37.151643},{"$f":"'DF raw data'!J17/100","v":-1.7227636999999998},{"$f":"'DF raw data'!N17/100","v":42.095420000000004},{"$f":"'DF raw data'!O17/100","v":10.761562},{"$f":"'DF raw data'!S17/100","v":34.099299000000002},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":23.440547000000002},{"$f":"'DF raw data'!Y17/100","v":2.4190331999999999}]
[8,{"$f":"'DF raw data'!D18/100","v":21.653319999999997},{"$f":"'DF raw data'!E18/100","v":5.1042284999999996},{"$f":"'DF raw data'!I18/100","v":42.128368999999999},{"$f":"'DF raw data'!J18/100","v":-11.445898},{"$f":"'DF raw data'!N18/100","v":28.039843999999999},{"$f":"'DF raw data'!O18/100","v":-10.068769999999999},{"$f":"'DF raw data'!S18/100","v":27.803247000000002},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":20.377493999999999},{"$f":"'DF raw data'!Y18/100","v":-5.4884569999999995}]
[10,{"$f":"'DF raw data'!D19/100","v":43.614106},{"$f":"'DF raw data'!E19/100","v":17.023115000000001},{"$f":"'DF raw data'!I19/100","v":54.686099000000006},{"$f":"'DF raw data'!J19/100","v":39.003681999999998},{"$f":"'DF raw data'!N19/100","v":39.212868999999998},{"$f":"'DF raw data'!O19/100","v":12.472871},{"$f":"'DF raw data'!S19/100","v":46.179589999999997},13.3657,{"$f":"'DF raw data'!X19/100","v":26.880531999999999},{"$f":"'DF raw data'!Y19/100","v":24.736553000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":247},null,{"$f":"'DF raw data'!G4","v":248},null,{"$f":"'DF raw data'!L4","v":251},null,{"$f":"'DF raw data'!Q4","v":253},null,{"$f":"'DF raw data'!V4","v":254}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":19.356675333333332},{"$f":"AVERAGE(C11:C13)","v":2.5490820000000003},25.108498333333333,8.7223536666666668,30.039383000000001,5.1583855333333331,35.684873666666668,7.7428600000000003,20.409622333333335,1.2655339666666672]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16336 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16336 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16336 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16336 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16336 100 mm gap"]
["Date:","10. mar 2017","14:23:20",null,null,"Date:","10. mar 2017","14:37:33",null,null,"Date:","10. mar 2017","15:23:33",null,null,"Date:","10. mar 2017","15:49:58",null,null,"Date:","10. mar 2017","16:04:48"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",247,null,null,null,"Indice:",248,null,null,null,"Indice:",251,null,null,null,"Indice:",253,null,null,null,"Indice:",254]
["Background Indice:",247,null,null,null,"Background Indice:",248,null,null,null,"Background Indice:",251,null,null,null,"Background Indice:",253,null,null,null,"Background Indice:",254]
["Temperature:",21.97,null,null,null,"Temperature:",21.97,null,null,null,"Temperature:",21.95,null,null,null,"Temperature:",21.93,null,null,null,"Temperature:",21.97]
["Encoder:",393685,null,null,null,"Encoder:",494274,null,null,null,"Encoder:",594599,null,null,null,"Encoder:",794907,null,null,null,"Encoder:",2195000]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,20.665130999999999,10.398350000000001,4522.4647999999997,1943.6738,-1,20.894439999999999,7.1504697999999998,4169.7665999999999,2457.9639000000002,-1,11.64385,1.8824799999999999,2159.6196,1847.4784999999999,-1,12.586980000000001,9.8557997000000004,3142.3359,1721.7148,-1,5.9211998000000001,-4.5762299999999998,824.07213999999999,-648.10937999999999]
[-0.80000000999999998,1.0530999999999999,8.2679796000000003,502.68725999999998,2987.7168000000001,-0.80000000999999998,10.840260000000001,4.7358899000000001,2829.8703999999998,1134.1171999999999,-0.80000000999999998,15.20275,6.9724202000000002,3722.0059000000001,1831.7236,-0.80000000999999998,15.8843,4.0632099999999998,3973.8330000000001,3200.7988,-0.80000000999999998,8.6791295999999996,0.25101000000000001,2961.5180999999998,27.154301]
[-0.60000001999999997,12.976610000000001,3.0506201000000002,3821.9814000000001,2073.2246,-0.60000001999999997,15.149279999999999,1.7048099999999999,5317.2275,1591.0947000000001,-0.60000001999999997,9.2420501999999995,-2.6835499,3432.2563,1014.1104,-0.60000001999999997,22.084378999999998,4.5042200000000001,5458.6620999999996,732.24315999999999,-0.60000001999999997,18.534969,-3.5035501,4904.9579999999996,519.72559000000001]
[-0.40000001000000002,-3.7654399999999999,-0.35185,-522.52173000000005,2770.1133,-0.40000001000000002,8.4055900999999995,4.1184602000000003,2609.1223,2140.5479,-0.40000001000000002,10.655110000000001,4.3629699000000004,3990.4301999999998,1631.1777,-0.40000001000000002,16.464680000000001,7.8439598000000004,4453.8950000000004,5690.0879000000004,-0.40000001000000002,11.39208,5.4384499000000002,3004.7719999999999,2182.8838000000001]
[-0.2,-0.29049000000000003,-10.562340000000001,1313.4808,-2455.8145,-0.2,7.28409,-6.6057601000000004,1439.4021,773.24805000000003,-0.2,4.7923397999999997,-3.0118699000000002,2085.0630000000001,-811.41503999999998,-0.2,7.6083999000000002,6.4539999000000001E-2,4189.2479999999996,-167.53809000000001,-0.2,1.58684,9.0680002999999995E-2,1340.4283,-679.74219000000005]
[0,4.5007200000000003,-4.1078099999999997,2279.1873000000001,1909.2188000000001,0,8.1981801999999995,-0.87041002999999995,3210.373,1314.5977,0,9.7185898000000002,-6.6921100999999998,3750.0576000000001,1416.4306999999999,0,12.128690000000001,-6.8943000000000003,4754.4013999999997,-565.47753999999998,0,9.4333200000000001,10.73086,2947.6558,1543.6309000000001]
[0.2,6.4421701000000002,-2.1147699000000002,2214.3344999999999,1311.3203000000001,0.2,2.4385099000000001,-7.0314598000000004,2882.7743999999998,528.86035000000004,0.2,10.22298,-6.1335502000000002,3176.6943000000001,942.5,0.2,3.71468,-0.14865999999999999,1761.8126999999999,599.34569999999997,0.2,5.8511800999999997,-2.1437198999999998,1834.8026,-484.22852]
[0.40000001000000002,7.7705102000000004,-8.6805000000000003,2196.0228999999999,-985.57421999999997,0.40000001000000002,3.8210299000000001,-3.6071501000000001,1878.4446,2885.8906000000002,0.40000001000000002,13.58405,-2.2965800999999999,2062.8586,1400.2109,0.40000001000000002,12.600910000000001,6.3321199000000004,2309.9944,2644.9960999999998,0.40000001000000002,15.565659999999999,6.8545799000000001,2078.6763000000001,2467.5801000000001]
[0.60000001999999997,8.1126099000000007,-10.83893,3415.5291000000002,210.02930000000001,0.60000001999999997,1.6286499999999999,-15.12337,3715.1642999999999,-172.27636999999999,0.60000001999999997,16.727409000000002,-3.6581999999999999,4209.5420000000004,1076.1561999999999,0.60000001999999997,17.127179999999999,-8.5406198999999994,3409.9299000000001,229.46190999999999,0.60000001999999997,10.01327,8.8025702999999993,2344.0547000000001,241.90332000000001]
[0.80000000999999998,12.70984,-11.583769999999999,2165.3319999999999,510.42284999999998,0.80000000999999998,11.305429999999999,-9.5385703999999993,4212.8369000000002,-1144.5898,0.80000000999999998,11.733140000000001,-6.4678798000000004,2803.9843999999998,-1006.877,0.80000000999999998,11.03403,-6.0744400000000001,2780.3247000000001,-1455.25,0.80000000999999998,5.9071002000000004,0.46665001,2037.7493999999999,-548.84569999999997]
[1,17.442419000000001,-2.3392700999999998,4361.4106000000002,1702.3115,1,31.26906,-0.88476997999999996,5468.6099000000004,3900.3681999999999,1,12.43543,5.3699002,3921.2869000000001,1247.2871,1,12.80176,0.82713002000000002,4617.9589999999998,1876.9277,1,9.4711504000000009,10.10704,2688.0531999999998,2473.6552999999999]