{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":41,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":383465},null,{"$f":"'DF raw data'!G7","v":484056},null,{"$f":"'DF raw data'!L7","v":584393},null,{"$f":"'DF raw data'!Q7","v":784663},null,{"$f":"'DF raw data'!V7","v":2184611},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-1.7508366900000001},{"$f":"C21","v":-5.1006766666666659},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":10.96684321},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":5.2068331099999998}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-5.2938499999999999},{"$f":"E21","v":-8.0916996433333335},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":10.96684321},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":6.0383863099999999}]
[-10,{"$f":"'DF raw data'!B9","v":-14.925700000000001},{"$f":"'DF raw data'!C9","v":-16.470269999999999},{"$f":"'DF raw data'!G9","v":-15.65727},{"$f":"'DF raw data'!H9","v":-18.70186},{"$f":"'DF raw data'!L9","v":-18.680599000000001},{"$f":"'DF raw data'!M9","v":-22.517589999999998},{"$f":"'DF raw data'!Q9","v":-27.23048},{"$f":"'DF raw data'!R9","v":-20.292601000000001},{"$f":"'DF raw data'!V9","v":-10.649330000000001},{"$f":"'DF raw data'!W9","v":-5.1796597999999996},null,{"$f":"F4","v":20},{"$f":"F21","v":-9.2705033333333322},{"$f":"G21","v":-2.8848665333333336}]
[-8,{"$f":"'DF raw data'!B10","v":-4.9471002000000004},{"$f":"'DF raw data'!C10","v":-19.356339999999999},{"$f":"'DF raw data'!G10","v":1.53071},{"$f":"'DF raw data'!H10","v":-11.906359999999999},{"$f":"'DF raw data'!L10","v":-16.434111000000001},{"$f":"'DF raw data'!M10","v":-13.89709},{"$f":"'DF raw data'!Q10","v":-19.225629999999999},{"$f":"'DF raw data'!R10","v":-6.1830100999999997},{"$f":"'DF raw data'!V10","v":-8.8945503000000006},{"$f":"'DF raw data'!W10","v":4.7195201000000004},null,{"$f":"H4","v":30},{"$f":"H21","v":-12.7176799},{"$f":"I21","v":-4.0394666673333335},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":5.6953000999999999},{"$f":"'DF raw data'!C11","v":-24.080959},{"$f":"'DF raw data'!G11","v":-7.3682499000000004},{"$f":"'DF raw data'!H11","v":-14.943020000000001},{"$f":"'DF raw data'!L11","v":-14.284940000000001},{"$f":"'DF raw data'!M11","v":-13.8765},{"$f":"'DF raw data'!Q11","v":-19.552420000000001},{"$f":"'DF raw data'!R11","v":-5.6349100999999999},{"$f":"'DF raw data'!V11","v":-2.1487701000000001},{"$f":"'DF raw data'!W11","v":5.5802202000000003},null,{"$f":"J4","v":100},{"$f":"J21","v":-10.598473366666667},{"$f":"K21","v":-2.0533133333333335},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-9.8719999000000003E-2},{"$f":"'DF raw data'!C12","v":-11.667529999999999},{"$f":"'DF raw data'!G12","v":0.42501},{"$f":"'DF raw data'!H12","v":-3.1494401000000001},{"$f":"'DF raw data'!L12","v":-8.4636601999999996},{"$f":"'DF raw data'!M12","v":-2.4711498999999999},{"$f":"'DF raw data'!Q12","v":-9.7861404000000007},{"$f":"'DF raw data'!R12","v":-3.94998},{"$f":"'DF raw data'!V12","v":-5.9657698000000003},{"$f":"'DF raw data'!W12","v":4.4806999999999997},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-3.1478701},{"$f":"'DF raw data'!C13","v":-1.70618},{"$f":"'DF raw data'!G13","v":-7.1217299000000001},{"$f":"'DF raw data'!H13","v":-6.9072098999999998},{"$f":"'DF raw data'!L13","v":-5.3107699999999998},{"$f":"'DF raw data'!M13","v":-7.0240698000000004},{"$f":"'DF raw data'!Q13","v":-13.7799},{"$f":"'DF raw data'!R13","v":-6.6956401000000003},{"$f":"'DF raw data'!V13","v":-12.675079999999999},{"$f":"'DF raw data'!W13","v":-3.6101401000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-1.4948900000000001},{"$f":"'DF raw data'!C14","v":-15.067119999999999},{"$f":"'DF raw data'!G14","v":2.9874599000000002},{"$f":"'DF raw data'!H14","v":-16.392229},{"$f":"'DF raw data'!L14","v":-5.5382999999999996},{"$f":"'DF raw data'!M14","v":-8.9965200000000003},{"$f":"'DF raw data'!Q14","v":-15.10263},{"$f":"'DF raw data'!R14","v":-5.4806198999999998},{"$f":"'DF raw data'!V14","v":-16.51247},{"$f":"'DF raw data'!W14","v":-3.3754699000000001},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-0.60974996999999997},{"$f":"'DF raw data'!C15","v":1.4712700000000001},{"$f":"'DF raw data'!G15","v":-11.74728},{"$f":"'DF raw data'!H15","v":-0.97566003000000001},{"$f":"'DF raw data'!L15","v":-16.962440000000001},{"$f":"'DF raw data'!M15","v":7.3659901999999997},{"$f":"'DF raw data'!Q15","v":-9.2705096999999999},{"$f":"'DF raw data'!R15","v":5.7859998000000003E-2},{"$f":"'DF raw data'!V15","v":-2.6078701},{"$f":"'DF raw data'!W15","v":0.82567000000000002},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-7.75875},{"$f":"'DF raw data'!C16","v":-2.6644899999999998},{"$f":"'DF raw data'!G16","v":2.25088},{"$f":"'DF raw data'!H16","v":2.6684101},{"$f":"'DF raw data'!L16","v":-11.684430000000001},{"$f":"'DF raw data'!M16","v":-7.7821999000000002},{"$f":"'DF raw data'!Q16","v":-10.05625},{"$f":"'DF raw data'!R16","v":-1.9093699},{"$f":"'DF raw data'!V16","v":-5.7650699999999997},{"$f":"'DF raw data'!W16","v":-3.4578500000000001}]
[6,{"$f":"'DF raw data'!B17","v":-9.5122804999999993},{"$f":"'DF raw data'!C17","v":7.9639901999999996},{"$f":"'DF raw data'!G17","v":-7.3312802000000001},{"$f":"'DF raw data'!H17","v":16.246599},{"$f":"'DF raw data'!L17","v":-20.671859999999999},{"$f":"'DF raw data'!M17","v":3.0532601000000001},{"$f":"'DF raw data'!Q17","v":-12.072760000000001},{"$f":"'DF raw data'!R17","v":-4.4616799},{"$f":"'DF raw data'!V17","v":-1.3936200000000001},{"$f":"'DF raw data'!W17","v":0.26601999999999998}]
[8,{"$f":"'DF raw data'!B18","v":-12.009639999999999},{"$f":"'DF raw data'!C18","v":7.5310302},{"$f":"'DF raw data'!G18","v":-15.629659999999999},{"$f":"'DF raw data'!H18","v":14.25418},{"$f":"'DF raw data'!L18","v":-11.753489999999999},{"$f":"'DF raw data'!M18","v":1.8711599999999999},{"$f":"'DF raw data'!Q18","v":-15.22044},{"$f":"'DF raw data'!R18","v":1.1766700000000001},{"$f":"'DF raw data'!V18","v":-0.22341},{"$f":"'DF raw data'!W18","v":15.91905}]
[10,{"$f":"'DF raw data'!B19","v":-19.062940999999999},{"$f":"'DF raw data'!C19","v":13.091659999999999},{"$f":"'DF raw data'!G19","v":-18.696421000000001},{"$f":"'DF raw data'!H19","v":4.2351298000000002},{"$f":"'DF raw data'!L19","v":-15.38846},{"$f":"'DF raw data'!M19","v":1.0077799999999999},{"$f":"'DF raw data'!Q19","v":-11.22372},{"$f":"'DF raw data'!R19","v":2.1614201},{"$f":"'DF raw data'!V19","v":-9.5634297999999998},{"$f":"'DF raw data'!W19","v":10.89509}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":263},null,{"$f":"'DF raw data'!G4","v":264},null,{"$f":"'DF raw data'!L4","v":265},null,{"$f":"'DF raw data'!Q4","v":267},null,{"$f":"'DF raw data'!V4","v":268}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-1.7508366900000001},{"$f":"AVERAGE(C11:C13)","v":-5.1006766666666659},-5.2938499999999999,-8.0916996433333335,-9.2705033333333322,-2.8848665333333336,-12.7176799,-4.0394666673333335,-10.598473366666667,-2.0533133333333335]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":383465},null,{"$f":"'DF raw data'!G7","v":484056},null,{"$f":"'DF raw data'!L7","v":584393},null,{"$f":"'DF raw data'!Q7","v":784663},null,{"$f":"'DF raw data'!V7","v":2184611},null,null,{"$f":"B4","v":10},{"$f":"B21","v":1.8970686000000005},{"$f":"C21","v":6.7799740000000002},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":33.825416599999997},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":17.679036799999999}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":17.895627666666666},{"$f":"E21","v":-4.4046941333333329},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":33.825416599999997},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":17.679036799999999}]
[-10,{"$f":"'DF raw data'!D9/100","v":-33.243285999999998},{"$f":"'DF raw data'!E9/100","v":1.9062109},{"$f":"'DF raw data'!I9/100","v":-25.657548999999999},{"$f":"'DF raw data'!J9/100","v":-16.607227000000002},{"$f":"'DF raw data'!N9/100","v":-29.286895000000001},{"$f":"'DF raw data'!O9/100","v":-9.5703809},{"$f":"'DF raw data'!S9/100","v":-31.615576000000001},{"$f":"'DF raw data'!T9/100","v":-27.565781000000001},{"$f":"'DF raw data'!X9/100","v":-15.071949},{"$f":"'DF raw data'!Y9/100","v":-10.475653999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":1.5846964999999997},{"$f":"G21","v":13.274342666666668}]
[-8,{"$f":"'DF raw data'!D10/100","v":-13.739068999999999},{"$f":"'DF raw data'!E10/100","v":2.0804784999999999},{"$f":"'DF raw data'!I10/100","v":1.0574683},{"$f":"'DF raw data'!J10/100","v":-7.9238964999999997},{"$f":"'DF raw data'!N10/100","v":-23.442612},{"$f":"'DF raw data'!O10/100","v":-6.7548800000000006E-2},{"$f":"'DF raw data'!S10/100","v":-31.354281999999998},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":-6.2787585000000004},{"$f":"'DF raw data'!Y10/100","v":10.435840000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-15.929788933333334},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":8.1517590000000002},{"$f":"'DF raw data'!E11/100","v":-18.603584000000001},{"$f":"'DF raw data'!I11/100","v":3.1749817},{"$f":"'DF raw data'!J11/100","v":-13.926269999999999},{"$f":"'DF raw data'!N11/100","v":12.12032},{"$f":"'DF raw data'!O11/100","v":-14.632070000000001},{"$f":"'DF raw data'!S11/100","v":-27.264628999999999},2.20269,{"$f":"'DF raw data'!X11/100","v":9.5631444999999999},{"$f":"'DF raw data'!Y11/100","v":10.671904},null,{"$f":"J4","v":100},{"$f":"J21","v":-6.5490439},{"$f":"K21","v":5.8949121333333325}]
[-4,{"$f":"'DF raw data'!D12/100","v":1.8024853999999999},{"$f":"'DF raw data'!E12/100","v":2.9325098000000001},{"$f":"'DF raw data'!I12/100","v":2.3208972000000001},{"$f":"'DF raw data'!J12/100","v":11.781181999999999},{"$f":"'DF raw data'!N12/100","v":0.38122559},{"$f":"'DF raw data'!O12/100","v":35.688613000000004},{"$f":"'DF raw data'!S12/100","v":-7.3045812999999997},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-3.0889929},{"$f":"'DF raw data'!Y12/100","v":11.240693}]
[-2,{"$f":"'DF raw data'!D13/100","v":-7.3825531},{"$f":"'DF raw data'!E13/100","v":-19.000498},{"$f":"'DF raw data'!I13/100","v":10.751505999999999},{"$f":"'DF raw data'!J13/100","v":2.1325878999999999},{"$f":"'DF raw data'!N13/100","v":14.222467999999999},{"$f":"'DF raw data'!O13/100","v":3.6660840000000001},{"$f":"'DF raw data'!S13/100","v":-34.34272},1.63822,{"$f":"'DF raw data'!X13/100","v":-12.449155000000001},{"$f":"'DF raw data'!Y13/100","v":-6.1876660000000001},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":1.9150658999999999},{"$f":"'DF raw data'!E14/100","v":13.223369},{"$f":"'DF raw data'!I14/100","v":44.883739999999996},{"$f":"'DF raw data'!J14/100","v":-24.664932},{"$f":"'DF raw data'!N14/100","v":6.0955724999999994},{"$f":"'DF raw data'!O14/100","v":10.941592},{"$f":"'DF raw data'!S14/100","v":-12.531929999999999},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":-9.5478943000000012},{"$f":"'DF raw data'!Y14/100","v":1.1330663999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":11.158693000000001},{"$f":"'DF raw data'!E15/100","v":26.117051},{"$f":"'DF raw data'!I15/100","v":-1.9483629999999998},{"$f":"'DF raw data'!J15/100","v":9.3182617000000008},{"$f":"'DF raw data'!N15/100","v":-15.563950999999999},{"$f":"'DF raw data'!O15/100","v":25.215351999999999},{"$f":"'DF raw data'!S15/100","v":-0.91471680000000011},16.2971,{"$f":"'DF raw data'!X15/100","v":2.3499175999999999},{"$f":"'DF raw data'!Y15/100","v":22.739335999999998},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-14.528483000000001},{"$f":"'DF raw data'!E16/100","v":3.2032227},{"$f":"'DF raw data'!I16/100","v":34.585619999999999},{"$f":"'DF raw data'!J16/100","v":18.510497999999998},{"$f":"'DF raw data'!N16/100","v":-25.148386000000002},{"$f":"'DF raw data'!O16/100","v":15.38374},{"$f":"'DF raw data'!S16/100","v":-29.500320000000002},10.6371,{"$f":"'DF raw data'!X16/100","v":-4.0017009999999997},{"$f":"'DF raw data'!Y16/100","v":14.117246}]
[6,{"$f":"'DF raw data'!D17/100","v":13.62391},{"$f":"'DF raw data'!E17/100","v":27.029121},{"$f":"'DF raw data'!I17/100","v":-25.100210000000001},{"$f":"'DF raw data'!J17/100","v":33.936718999999997},{"$f":"'DF raw data'!N17/100","v":-31.025027000000001},{"$f":"'DF raw data'!O17/100","v":9.3358594000000004},{"$f":"'DF raw data'!S17/100","v":-11.399165},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":4.4166748},{"$f":"'DF raw data'!Y17/100","v":14.921425999999999}]
[8,{"$f":"'DF raw data'!D18/100","v":-10.842563},{"$f":"'DF raw data'!E18/100","v":12.742100000000001},{"$f":"'DF raw data'!I18/100","v":-28.367411999999998},{"$f":"'DF raw data'!J18/100","v":6.0408007999999995},{"$f":"'DF raw data'!N18/100","v":-7.0955944999999998},{"$f":"'DF raw data'!O18/100","v":9.8817187999999998},{"$f":"'DF raw data'!S18/100","v":-11.641592000000001},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":9.1466118999999999},{"$f":"'DF raw data'!Y18/100","v":23.036152000000001}]
[10,{"$f":"'DF raw data'!D19/100","v":-24.450075999999999},{"$f":"'DF raw data'!E19/100","v":33.807235999999996},{"$f":"'DF raw data'!I19/100","v":3.6512340999999999},{"$f":"'DF raw data'!J19/100","v":32.074140999999997},{"$f":"'DF raw data'!N19/100","v":-14.980383},{"$f":"'DF raw data'!O19/100","v":0.15447269999999999},{"$f":"'DF raw data'!S19/100","v":-8.9668439000000006},13.3657,{"$f":"'DF raw data'!X19/100","v":-0.85736510999999993},{"$f":"'DF raw data'!Y19/100","v":20.785010000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":263},null,{"$f":"'DF raw data'!G4","v":264},null,{"$f":"'DF raw data'!L4","v":265},null,{"$f":"'DF raw data'!Q4","v":267},null,{"$f":"'DF raw data'!V4","v":268}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":1.8970686000000005},{"$f":"AVERAGE(C11:C13)","v":6.7799740000000002},17.895627666666666,-4.4046941333333329,1.5846964999999997,13.274342666666668,-15.929788933333334,7.7428600000000003,-6.5490439,5.8949121333333325]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16337 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16337 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16337 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16337 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16337 100 mm gap"]
["Date:","14. mar 2017","12:54:02",null,null,"Date:","14. mar 2017","13:08:57",null,null,"Date:","14. mar 2017","13:23:18",null,null,"Date:","14. mar 2017","13:53:37",null,null,"Date:","14. mar 2017","14:06:55"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",263,null,null,null,"Indice:",264,null,null,null,"Indice:",265,null,null,null,"Indice:",267,null,null,null,"Indice:",268]
["Background Indice:",263,null,null,null,"Background Indice:",264,null,null,null,"Background Indice:",265,null,null,null,"Background Indice:",267,null,null,null,"Background Indice:",268]
["Temperature:",22.01,null,null,null,"Temperature:",21.99,null,null,null,"Temperature:",22.01,null,null,null,"Temperature:",22.04,null,null,null,"Temperature:",22.02]
["Encoder:",383465,null,null,null,"Encoder:",484056,null,null,null,"Encoder:",584393,null,null,null,"Encoder:",784663,null,null,null,"Encoder:",2184611]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-14.925700000000001,-16.470269999999999,-3324.3285999999998,190.62109000000001,-1,-15.65727,-18.70186,-2565.7548999999999,-1660.7227,-1,-18.680599000000001,-22.517589999999998,-2928.6895,-957.03809000000001,-1,-27.23048,-20.292601000000001,-3161.5576000000001,-2756.5781000000002,-1,-10.649330000000001,-5.1796597999999996,-1507.1949,-1047.5654]
[-0.80000000999999998,-4.9471002000000004,-19.356339999999999,-1373.9069,208.04785000000001,-0.80000000999999998,1.53071,-11.906359999999999,105.74683,-792.38964999999996,-0.80000000999999998,-16.434111000000001,-13.89709,-2344.2611999999999,-6.75488,-0.80000000999999998,-19.225629999999999,-6.1830100999999997,-3135.4281999999998,-2759.5576000000001,-0.80000000999999998,-8.8945503000000006,4.7195201000000004,-627.87585000000001,1043.5840000000001]
[-0.60000001999999997,5.6953000999999999,-24.080959,815.17589999999996,-1860.3584000000001,-0.60000001999999997,-7.3682499000000004,-14.943020000000001,317.49817000000002,-1392.627,-0.60000001999999997,-14.284940000000001,-13.8765,1212.0319999999999,-1463.2070000000001,-0.60000001999999997,-19.552420000000001,-5.6349100999999999,-2726.4629,-400.14940999999999,-0.60000001999999997,-2.1487701000000001,5.5802202000000003,956.31444999999997,1067.1904]
[-0.40000001000000002,-9.8719999000000003E-2,-11.667529999999999,180.24853999999999,293.25098000000003,-0.40000001000000002,0.42501,-3.1494401000000001,232.08972,1178.1181999999999,-0.40000001000000002,-8.4636601999999996,-2.4711498999999999,38.122559000000003,3568.8613,-0.40000001000000002,-9.7861404000000007,-3.94998,-730.45812999999998,699.61132999999995,-0.40000001000000002,-5.9657698000000003,4.4806999999999997,-308.89929000000001,1124.0693000000001]
[-0.2,-3.1478701,-1.70618,-738.25531000000001,-1900.0498,-0.2,-7.1217299000000001,-6.9072098999999998,1075.1505999999999,213.25879,-0.2,-5.3107699999999998,-7.0240698000000004,1422.2467999999999,366.60840000000002,-0.2,-13.7799,-6.6956401000000003,-3434.2719999999999,-1112.7188000000001,-0.2,-12.675079999999999,-3.6101401000000002,-1244.9155000000001,-618.76660000000004]
[0,-1.4948900000000001,-15.067119999999999,191.50658999999999,1322.3369,0,2.9874599000000002,-16.392229,4488.3739999999998,-2466.4931999999999,0,-5.5382999999999996,-8.9965200000000003,609.55724999999995,1094.1592000000001,0,-15.10263,-5.4806198999999998,-1253.193,-294.27832000000001,0,-16.51247,-3.3754699000000001,-954.78943000000004,113.30664]
[0.2,-0.60974996999999997,1.4712700000000001,1115.8693000000001,2611.7051000000001,0.2,-11.74728,-0.97566003000000001,-194.83629999999999,931.82617000000005,0.2,-16.962440000000001,7.3659901999999997,-1556.3951,2521.5351999999998,0.2,-9.2705096999999999,5.7859998000000003E-2,-91.471680000000006,1685.8613,0.2,-2.6078701,0.82567000000000002,234.99176,2273.9335999999998]
[0.40000001000000002,-7.75875,-2.6644899999999998,-1452.8483000000001,320.32227,0.40000001000000002,2.25088,2.6684101,3458.5619999999999,1851.0498,0.40000001000000002,-11.684430000000001,-7.7821999000000002,-2514.8386,1538.374,0.40000001000000002,-10.05625,-1.9093699,-2950.0320000000002,658.72167999999999,0.40000001000000002,-5.7650699999999997,-3.4578500000000001,-400.17009999999999,1411.7246]
[0.60000001999999997,-9.5122804999999993,7.9639901999999996,1362.3910000000001,2702.9121,0.60000001999999997,-7.3312802000000001,16.246599,-2510.0210000000002,3393.6718999999998,0.60000001999999997,-20.671859999999999,3.0532601000000001,-3102.5027,933.58594000000005,0.60000001999999997,-12.072760000000001,-4.4616799,-1139.9165,1280.6044999999999,0.60000001999999997,-1.3936200000000001,0.26601999999999998,441.66748000000001,1492.1425999999999]
[0.80000000999999998,-12.009639999999999,7.5310302,-1084.2563,1274.21,0.80000000999999998,-15.629659999999999,14.25418,-2836.7411999999999,604.08007999999995,0.80000000999999998,-11.753489999999999,1.8711599999999999,-709.55944999999997,988.17187999999999,0.80000000999999998,-15.22044,1.1766700000000001,-1164.1592000000001,-89.928711000000007,0.80000000999999998,-0.22341,15.91905,914.66119000000003,2303.6152000000002]
[1,-19.062940999999999,13.091659999999999,-2445.0075999999999,3380.7235999999998,1,-18.696421000000001,4.2351298000000002,365.12340999999998,3207.4141,1,-15.38846,1.0077799999999999,-1498.0382999999999,15.44727,1,-11.22372,2.1614201,-896.68439000000001,1540.2665999999999,1,-9.5634297999999998,10.89509,-85.736510999999993,2078.5010000000002]