{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":61,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":390844},null,{"$f":"'DF raw data'!G7","v":491397},null,{"$f":"'DF raw data'!L7","v":591770},null,{"$f":"'DF raw data'!Q7","v":792001},null,{"$f":"'DF raw data'!V7","v":2191994},null,null,{"$f":"B4","v":10},{"$f":"B21","v":8.6326066666666676},{"$f":"C21","v":-2.4694066766666665},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":6.4155267000000009},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":7.6024367366666663}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":5.2768699666666663},{"$f":"E21","v":1.6201132966666669},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":6.4155267000000009},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":7.6024367366666663}]
[-10,{"$f":"'DF raw data'!B9","v":-6.8290003000000002E-2},{"$f":"'DF raw data'!C9","v":-19.517638999999999},{"$f":"'DF raw data'!G9","v":-2.44028},{"$f":"'DF raw data'!H9","v":-12.52792},{"$f":"'DF raw data'!L9","v":-4.4517398000000004},{"$f":"'DF raw data'!M9","v":-5.9031700999999996},{"$f":"'DF raw data'!Q9","v":-5.8045998000000001},{"$f":"'DF raw data'!R9","v":-3.5119699999999998},{"$f":"'DF raw data'!V9","v":-0.19878000000000001},{"$f":"'DF raw data'!W9","v":12.60059},null,{"$f":"F4","v":20},{"$f":"F21","v":6.2008966666666678},{"$f":"G21","v":5.1330300600000003}]
[-8,{"$f":"'DF raw data'!B10","v":2.8762400000000001},{"$f":"'DF raw data'!C10","v":-15.25848},{"$f":"'DF raw data'!G10","v":5.0292500999999996},{"$f":"'DF raw data'!H10","v":-12.528840000000001},{"$f":"'DF raw data'!L10","v":1.3040299},{"$f":"'DF raw data'!M10","v":-3.0804501000000002},{"$f":"'DF raw data'!Q10","v":16.958919999999999},{"$f":"'DF raw data'!R10","v":-1.68808},{"$f":"'DF raw data'!V10","v":7.0257000999999999},{"$f":"'DF raw data'!W10","v":2.0891299000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":2.2170799666666667},{"$f":"I21","v":1.619469966666667},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":5.4912200000000002},{"$f":"'DF raw data'!C11","v":-22.148340000000001},{"$f":"'DF raw data'!G11","v":8.4145097999999994},{"$f":"'DF raw data'!H11","v":-6.4864302},{"$f":"'DF raw data'!L11","v":2.02915},{"$f":"'DF raw data'!M11","v":-7.2805499999999999},{"$f":"'DF raw data'!Q11","v":-2.43648},{"$f":"'DF raw data'!R11","v":-4.28688},{"$f":"'DF raw data'!V11","v":4.1708698000000002},{"$f":"'DF raw data'!W11","v":-1.4681200000000001},null,{"$f":"J4","v":100},{"$f":"J21","v":3.5000933033333332},{"$f":"K21","v":0.50525660000000017},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-4.0674600999999999},{"$f":"'DF raw data'!C12","v":-6.6291599000000003},{"$f":"'DF raw data'!G12","v":5.3109598},{"$f":"'DF raw data'!H12","v":-5.0586801000000001},{"$f":"'DF raw data'!L12","v":0.10771},{"$f":"'DF raw data'!M12","v":4.5401201000000002},{"$f":"'DF raw data'!Q12","v":-6.8024402000000004},{"$f":"'DF raw data'!R12","v":6.9815202000000003},{"$f":"'DF raw data'!V12","v":-4.7411798999999997},{"$f":"'DF raw data'!W12","v":-0.23544000000000001},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":8.4053602000000005},{"$f":"'DF raw data'!C13","v":-6.9767999999999999},{"$f":"'DF raw data'!G13","v":6.6448897999999996},{"$f":"'DF raw data'!H13","v":4.8124199000000001},{"$f":"'DF raw data'!L13","v":1.64649},{"$f":"'DF raw data'!M13","v":-0.58117001999999995},{"$f":"'DF raw data'!Q13","v":-1.7440899999999999},{"$f":"'DF raw data'!R13","v":3.3041100999999999},{"$f":"'DF raw data'!V13","v":0.61342001000000002},{"$f":"'DF raw data'!W13","v":-4.3091401999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":9.7709799000000004},{"$f":"'DF raw data'!C14","v":-0.83595001999999996},{"$f":"'DF raw data'!G14","v":10.281890000000001},{"$f":"'DF raw data'!H14","v":-0.63569998999999999},{"$f":"'DF raw data'!L14","v":13.842790000000001},{"$f":"'DF raw data'!M14","v":5.1065402000000004},{"$f":"'DF raw data'!Q14","v":3.5578899000000002},{"$f":"'DF raw data'!R14","v":7.0382198999999996},{"$f":"'DF raw data'!V14","v":5.52597},{"$f":"'DF raw data'!W14","v":3.9647100000000002},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":7.7214799000000003},{"$f":"'DF raw data'!C15","v":0.40452999000000001},{"$f":"'DF raw data'!G15","v":-1.0961699},{"$f":"'DF raw data'!H15","v":0.68361998000000002},{"$f":"'DF raw data'!L15","v":3.11341},{"$f":"'DF raw data'!M15","v":10.87372},{"$f":"'DF raw data'!Q15","v":4.83744},{"$f":"'DF raw data'!R15","v":-5.4839200999999997},{"$f":"'DF raw data'!V15","v":4.3608899000000001},{"$f":"'DF raw data'!W15","v":1.8602000000000001},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":3.5197699},{"$f":"'DF raw data'!C16","v":8.5720997000000008},{"$f":"'DF raw data'!G16","v":13.760630000000001},{"$f":"'DF raw data'!H16","v":-5.3653301999999998},{"$f":"'DF raw data'!L16","v":12.694559999999999},{"$f":"'DF raw data'!M16","v":12.10873},{"$f":"'DF raw data'!Q16","v":9.6094503000000007},{"$f":"'DF raw data'!R16","v":5.4543900000000001},{"$f":"'DF raw data'!V16","v":6.2482100000000003},{"$f":"'DF raw data'!W16","v":6.6396198000000002}]
[6,{"$f":"'DF raw data'!B17","v":5.6844001000000004},{"$f":"'DF raw data'!C17","v":11.212249999999999},{"$f":"'DF raw data'!G17","v":8.6774301999999999},{"$f":"'DF raw data'!H17","v":6.7131701000000001},{"$f":"'DF raw data'!L17","v":17.488689000000001},{"$f":"'DF raw data'!M17","v":7.4182800999999996},{"$f":"'DF raw data'!Q17","v":14.684419999999999},{"$f":"'DF raw data'!R17","v":5.9612198000000003},{"$f":"'DF raw data'!V17","v":1.9536800000000001},{"$f":"'DF raw data'!W17","v":-6.1159401000000004}]
[8,{"$f":"'DF raw data'!B18","v":4.7358098000000002},{"$f":"'DF raw data'!C18","v":17.431141},{"$f":"'DF raw data'!G18","v":0.24697},{"$f":"'DF raw data'!H18","v":15.20241},{"$f":"'DF raw data'!L18","v":3.7042500999999999},{"$f":"'DF raw data'!M18","v":8.3168296999999995},{"$f":"'DF raw data'!Q18","v":9.0475396999999997},{"$f":"'DF raw data'!R18","v":14.99663},{"$f":"'DF raw data'!V18","v":2.4740099999999998},{"$f":"'DF raw data'!W18","v":-7.6116900000000003}]
[10,{"$f":"'DF raw data'!B19","v":16.159089999999999},{"$f":"'DF raw data'!C19","v":25.608179},{"$f":"'DF raw data'!G19","v":15.22213},{"$f":"'DF raw data'!H19","v":24.848551},{"$f":"'DF raw data'!L19","v":21.175899999999999},{"$f":"'DF raw data'!M19","v":12.63808},{"$f":"'DF raw data'!Q19","v":16.217251000000001},{"$f":"'DF raw data'!R19","v":5.4799699999999998},{"$f":"'DF raw data'!V19","v":-3.4462199},{"$f":"'DF raw data'!W19","v":-6.8116101999999996}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":303},null,{"$f":"'DF raw data'!G4","v":304},null,{"$f":"'DF raw data'!L4","v":305},null,{"$f":"'DF raw data'!Q4","v":306},null,{"$f":"'DF raw data'!V4","v":307}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":8.6326066666666676},{"$f":"AVERAGE(C11:C13)","v":-2.4694066766666665},5.2768699666666663,1.6201132966666669,6.2008966666666678,5.1330300600000003,2.2170799666666667,1.619469966666667,3.5000933033333332,0.50525660000000017]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":390844},null,{"$f":"'DF raw data'!G7","v":491397},null,{"$f":"'DF raw data'!L7","v":591770},null,{"$f":"'DF raw data'!Q7","v":792001},null,{"$f":"'DF raw data'!V7","v":2191994},null,null,{"$f":"B4","v":10},{"$f":"B21","v":15.110695166666666},{"$f":"C21","v":0.30758800000000025},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":16.820699433333331},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":13.188626233333332}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":6.146857999999999},{"$f":"E21","v":3.0733103333333331},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":16.820699433333331},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":18.125934433333335}]
[-10,{"$f":"'DF raw data'!D9/100","v":-3.3911254999999998},{"$f":"'DF raw data'!E9/100","v":-32.102676000000002},{"$f":"'DF raw data'!I9/100","v":12.815769},{"$f":"'DF raw data'!J9/100","v":-20.721093999999997},{"$f":"'DF raw data'!N9/100","v":-14.463363999999999},{"$f":"'DF raw data'!O9/100","v":-6.7573242000000002},{"$f":"'DF raw data'!S9/100","v":-51.192758999999995},{"$f":"'DF raw data'!T9/100","v":3.0744823999999999},{"$f":"'DF raw data'!X9/100","v":7.4748785},{"$f":"'DF raw data'!Y9/100","v":12.904033},null,{"$f":"F4","v":20},{"$f":"F21","v":8.7852754666666666},{"$f":"G21","v":13.496214233333333}]
[-8,{"$f":"'DF raw data'!D10/100","v":-0.40658810000000001},{"$f":"'DF raw data'!E10/100","v":-16.619140999999999},{"$f":"'DF raw data'!I10/100","v":0.45365601},{"$f":"'DF raw data'!J10/100","v":-9.7403417999999995},{"$f":"'DF raw data'!N10/100","v":-0.77520142000000003},{"$f":"'DF raw data'!O10/100","v":3.2696289000000003},{"$f":"'DF raw data'!S10/100","v":-1.2274236999999999},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":-3.9228778000000002},{"$f":"'DF raw data'!Y10/100","v":18.085546999999998},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.710004266666667},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":13.869584},{"$f":"'DF raw data'!E11/100","v":-24.337372999999999},{"$f":"'DF raw data'!I11/100","v":18.238972},{"$f":"'DF raw data'!J11/100","v":6.2676464999999997},{"$f":"'DF raw data'!N11/100","v":-3.4878680000000002},{"$f":"'DF raw data'!O11/100","v":-2.5225},{"$f":"'DF raw data'!S11/100","v":6.7566742},2.20269,{"$f":"'DF raw data'!X11/100","v":16.015917999999999},{"$f":"'DF raw data'!Y11/100","v":-1.45215E-2},null,{"$f":"J4","v":100},{"$f":"J21","v":5.5319347333333333},{"$f":"K21","v":-4.6297201999999995}]
[-4,{"$f":"'DF raw data'!D12/100","v":-0.7633297},{"$f":"'DF raw data'!E12/100","v":8.0329101999999999},{"$f":"'DF raw data'!I12/100","v":-17.4725},{"$f":"'DF raw data'!J12/100","v":17.493086000000002},{"$f":"'DF raw data'!N12/100","v":3.1367688},{"$f":"'DF raw data'!O12/100","v":10.359766},{"$f":"'DF raw data'!S12/100","v":-11.266304999999999},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-7.7350609999999991},{"$f":"'DF raw data'!Y12/100","v":-3.6373633000000001}]
[-2,{"$f":"'DF raw data'!D13/100","v":8.3640515000000004},{"$f":"'DF raw data'!E13/100","v":-1.5471680000000001},{"$f":"'DF raw data'!I13/100","v":21.071382},{"$f":"'DF raw data'!J13/100","v":13.826025},{"$f":"'DF raw data'!N13/100","v":6.9839398000000008},{"$f":"'DF raw data'!O13/100","v":11.363505999999999},{"$f":"'DF raw data'!S13/100","v":-11.969128000000001},1.63822,{"$f":"'DF raw data'!X13/100","v":-1.3694766},{"$f":"'DF raw data'!Y13/100","v":-15.658769999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":19.325362999999999},{"$f":"'DF raw data'!E14/100","v":-8.3718749999999993},{"$f":"'DF raw data'!I14/100","v":-13.427104},{"$f":"'DF raw data'!J14/100","v":10.234404},{"$f":"'DF raw data'!N14/100","v":4.0326886000000002},{"$f":"'DF raw data'!O14/100","v":6.2073046999999999},{"$f":"'DF raw data'!S14/100","v":-7.1812267999999992},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":8.3286651999999997},{"$f":"'DF raw data'!Y14/100","v":-3.6434570000000002},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":17.642671},{"$f":"'DF raw data'!E15/100","v":10.841806999999999},{"$f":"'DF raw data'!I15/100","v":10.796296},{"$f":"'DF raw data'!J15/100","v":-14.840498},{"$f":"'DF raw data'!N15/100","v":15.339198},{"$f":"'DF raw data'!O15/100","v":22.917831999999997},{"$f":"'DF raw data'!S15/100","v":14.020342000000001},16.2971,{"$f":"'DF raw data'!X15/100","v":9.6366156000000007},{"$f":"'DF raw data'!Y15/100","v":5.4130663999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":1.8436788999999998},{"$f":"'DF raw data'!E16/100","v":32.426475000000003},{"$f":"'DF raw data'!I16/100","v":3.2330386},{"$f":"'DF raw data'!J16/100","v":-7.2216895000000001},{"$f":"'DF raw data'!N16/100","v":10.302873999999999},{"$f":"'DF raw data'!O16/100","v":14.346602000000001},{"$f":"'DF raw data'!S16/100","v":2.4529602000000001},10.6371,{"$f":"'DF raw data'!X16/100","v":-17.482211},{"$f":"'DF raw data'!Y16/100","v":12.660029}]
[6,{"$f":"'DF raw data'!D17/100","v":9.7159618999999999},{"$f":"'DF raw data'!E17/100","v":37.151338000000003},{"$f":"'DF raw data'!I17/100","v":-4.7525231999999997},{"$f":"'DF raw data'!J17/100","v":24.048993999999997},{"$f":"'DF raw data'!N17/100","v":22.884243000000001},{"$f":"'DF raw data'!O17/100","v":18.994678},{"$f":"'DF raw data'!S17/100","v":25.370626999999999},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":6.4724792000000004},{"$f":"'DF raw data'!Y17/100","v":4.2419140999999998}]
[8,{"$f":"'DF raw data'!D18/100","v":-5.8115423999999996},{"$f":"'DF raw data'!E18/100","v":40.362304999999999},{"$f":"'DF raw data'!I18/100","v":-28.794656},{"$f":"'DF raw data'!J18/100","v":5.0694238},{"$f":"'DF raw data'!N18/100","v":-5.0592255000000002},{"$f":"'DF raw data'!O18/100","v":27.735273000000003},{"$f":"'DF raw data'!S18/100","v":22.739337999999996},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":13.597935},{"$f":"'DF raw data'!Y18/100","v":-0.25234381}]
[10,{"$f":"'DF raw data'!D19/100","v":34.758447000000004},{"$f":"'DF raw data'!E19/100","v":46.257313999999994},{"$f":"'DF raw data'!I19/100","v":48.456410999999996},{"$f":"'DF raw data'!J19/100","v":40.547167999999999},{"$f":"'DF raw data'!N19/100","v":37.993148999999995},{"$f":"'DF raw data'!O19/100","v":20.897012},{"$f":"'DF raw data'!S19/100","v":38.512402000000002},13.3657,{"$f":"'DF raw data'!X19/100","v":0.2552026},{"$f":"'DF raw data'!Y19/100","v":-15.57708}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":303},null,{"$f":"'DF raw data'!G4","v":304},null,{"$f":"'DF raw data'!L4","v":305},null,{"$f":"'DF raw data'!Q4","v":306},null,{"$f":"'DF raw data'!V4","v":307}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":15.110695166666666},{"$f":"AVERAGE(C11:C13)","v":0.30758800000000025},6.146857999999999,3.0733103333333331,8.7852754666666666,13.496214233333333,-1.710004266666667,7.7428600000000003,5.5319347333333333,-4.6297201999999995]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16340, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16340, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16340, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16340, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16340, 100 mm gap"]
["Date:","23. mar 2017","10:01:20",null,null,"Date:","23. mar 2017","10:20:12",null,null,"Date:","23. mar 2017","10:38:40",null,null,"Date:","23. mar 2017","10:53:28",null,null,"Date:","23. mar 2017","11:08:48"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",303,null,null,null,"Indice:",304,null,null,null,"Indice:",305,null,null,null,"Indice:",306,null,null,null,"Indice:",307]
["Background Indice:",303,null,null,null,"Background Indice:",304,null,null,null,"Background Indice:",305,null,null,null,"Background Indice:",306,null,null,null,"Background Indice:",307]
["Temperature:",21.65,null,null,null,"Temperature:",21.6,null,null,null,"Temperature:",21.59,null,null,null,"Temperature:",21.64,null,null,null,"Temperature:",21.62]
["Encoder:",390844,null,null,null,"Encoder:",491397,null,null,null,"Encoder:",591770,null,null,null,"Encoder:",792001,null,null,null,"Encoder:",2191994]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-6.8290003000000002E-2,-19.517638999999999,-339.11255,-3210.2676000000001,-1,-2.44028,-12.52792,1281.5769,-2072.1093999999998,-1,-4.4517398000000004,-5.9031700999999996,-1446.3363999999999,-675.73242000000005,-1,-5.8045998000000001,-3.5119699999999998,-5119.2758999999996,307.44824,-1,-0.19878000000000001,12.60059,747.48784999999998,1290.4032999999999]
[-0.80000000999999998,2.8762400000000001,-15.25848,-40.658810000000003,-1661.9141,-0.80000000999999998,5.0292500999999996,-12.528840000000001,45.365600999999998,-974.03417999999999,-0.80000000999999998,1.3040299,-3.0804501000000002,-77.520142000000007,326.96289000000002,-0.80000000999999998,16.958919999999999,-1.68808,-122.74236999999999,287.11523,-0.80000000999999998,7.0257000999999999,2.0891299000000001,-392.28778,1808.5546999999999]
[-0.60000001999999997,5.4912200000000002,-22.148340000000001,1386.9584,-2433.7372999999998,-0.60000001999999997,8.4145097999999994,-6.4864302,1823.8972000000001,626.76464999999996,-0.60000001999999997,2.02915,-7.2805499999999999,-348.78680000000003,-252.25,-0.60000001999999997,-2.43648,-4.28688,675.66741999999999,-71.366211000000007,-0.60000001999999997,4.1708698000000002,-1.4681200000000001,1601.5917999999999,-1.4521500000000001]
[-0.40000001000000002,-4.0674600999999999,-6.6291599000000003,-76.332970000000003,803.29102,-0.40000001000000002,5.3109598,-5.0586801000000001,-1747.25,1749.3086000000001,-0.40000001000000002,0.10771,4.5401201000000002,313.67687999999998,1035.9766,-0.40000001000000002,-6.8024402000000004,6.9815202000000003,-1126.6305,2579.4814000000001,-0.40000001000000002,-4.7411798999999997,-0.23544000000000001,-773.50609999999995,-363.73633000000001]
[-0.2,8.4053602000000005,-6.9767999999999999,836.40515000000005,-154.71680000000001,-0.2,6.6448897999999996,4.8124199000000001,2107.1381999999999,1382.6025,-0.2,1.64649,-0.58117001999999995,698.39398000000006,1136.3506,-0.2,-1.7440899999999999,3.3041100999999999,-1196.9128000000001,1619.3690999999999,-0.2,0.61342001000000002,-4.3091401999999999,-136.94766000000001,-1565.877]
[0,9.7709799000000004,-0.83595001999999996,1932.5363,-837.1875,0,10.281890000000001,-0.63569998999999999,-1342.7103999999999,1023.4404,0,13.842790000000001,5.1065402000000004,403.26886000000002,620.73046999999997,0,3.5578899000000002,7.0382198999999996,-718.12267999999995,694.51367000000005,0,5.52597,3.9647100000000002,832.86652000000004,-364.34570000000002]
[0.2,7.7214799000000003,0.40452999000000001,1764.2671,1084.1806999999999,0.2,-1.0961699,0.68361998000000002,1079.6296,-1484.0498,0.2,3.11341,10.87372,1533.9197999999999,2291.7831999999999,0.2,4.83744,-5.4839200999999997,1402.0342000000001,4101.7421999999997,0.2,4.3608899000000001,1.8602000000000001,963.66156000000001,541.30664000000002]
[0.40000001000000002,3.5197699,8.5720997000000008,184.36788999999999,3242.6475,0.40000001000000002,13.760630000000001,-5.3653301999999998,323.30385999999999,-722.16895,0.40000001000000002,12.694559999999999,12.10873,1030.2873999999999,1434.6602,0.40000001000000002,9.6094503000000007,5.4543900000000001,245.29602,-679.11621000000002,0.40000001000000002,6.2482100000000003,6.6396198000000002,-1748.2211,1266.0029]
[0.60000001999999997,5.6844001000000004,11.212249999999999,971.59618999999998,3715.1338000000001,0.60000001999999997,8.6774301999999999,6.7131701000000001,-475.25232,2404.8993999999998,0.60000001999999997,17.488689000000001,7.4182800999999996,2288.4243000000001,1899.4677999999999,0.60000001999999997,14.684419999999999,5.9612198000000003,2537.0626999999999,2065.9268000000002,0.60000001999999997,1.9536800000000001,-6.1159401000000004,647.24792000000002,424.19141000000002]
[0.80000000999999998,4.7358098000000002,17.431141,-581.15423999999996,4036.2305000000001,0.80000000999999998,0.24697,15.20241,-2879.4656,506.94238000000001,0.80000000999999998,3.7042500999999999,8.3168296999999995,-505.92255,2773.5273000000002,0.80000000999999998,9.0475396999999997,14.99663,2273.9337999999998,1622.5332000000001,0.80000000999999998,2.4740099999999998,-7.6116900000000003,1359.7935,-25.234380999999999]
[1,16.159089999999999,25.608179,3475.8447000000001,4625.7313999999997,1,15.22213,24.848551,4845.6410999999998,4054.7168000000001,1,21.175899999999999,12.63808,3799.3148999999999,2089.7012,1,16.217251000000001,5.4799699999999998,3851.2402000000002,-75.918953000000002,1,-3.4462199,-6.8116101999999996,25.52026,-1557.7080000000001]