{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":644,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":4460446},null,{"$f":"'DF raw data'!G7","v":4440399},null,{"$f":"'DF raw data'!L7","v":4420385},null,{"$f":"'DF raw data'!Q7","v":4380373},null,{"$f":"'DF raw data'!V7","v":4100376},null,null,{"$f":"B4","v":10},{"$f":"B21","v":5.0313433333333331},{"$f":"C21","v":-7.951722966666666},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":7.0141766433333332},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":8.3150463333333331}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":1.4238733300000002},{"$f":"E21","v":8.0623466666666907E-2},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":8.6073167366666663},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":8.3150463333333331}]
[-10,{"$f":"'DF raw data'!B9","v":9.0254697999999998},{"$f":"'DF raw data'!C9","v":14.898910000000001},{"$f":"'DF raw data'!G9","v":-2.0175098999999999},{"$f":"'DF raw data'!H9","v":14.129},{"$f":"'DF raw data'!L9","v":1.77145},{"$f":"'DF raw data'!M9","v":8.9104404000000006},{"$f":"'DF raw data'!Q9","v":-7.0439800999999997},{"$f":"'DF raw data'!R9","v":4.5016198000000003},{"$f":"'DF raw data'!V9","v":-1.2188300000000001},{"$f":"'DF raw data'!W9","v":9.6121902000000006},null,{"$f":"F4","v":20},{"$f":"F21","v":4.4277900333333333},{"$f":"G21","v":0.36332336666666681}]
[-8,{"$f":"'DF raw data'!B10","v":13.71156},{"$f":"'DF raw data'!C10","v":-0.66443998000000004},{"$f":"'DF raw data'!G10","v":8.9245195000000006},{"$f":"'DF raw data'!H10","v":8.9199103999999991},{"$f":"'DF raw data'!L10","v":0.92080998000000003},{"$f":"'DF raw data'!M10","v":-2.7620800000000001},{"$f":"'DF raw data'!Q10","v":3.3050499000000002},{"$f":"'DF raw data'!R10","v":0.96265000000000001},{"$f":"'DF raw data'!V10","v":-9.0067395999999995},{"$f":"'DF raw data'!W10","v":1.4834501},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.98283331},{"$f":"I21","v":-0.32642650000000023},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":7.9510198000000001},{"$f":"'DF raw data'!C11","v":25.194050000000001},{"$f":"'DF raw data'!G11","v":5.9669198999999997},{"$f":"'DF raw data'!H11","v":12.623329999999999},{"$f":"'DF raw data'!L11","v":1.2092700000000001},{"$f":"'DF raw data'!M11","v":8.3177996000000007},{"$f":"'DF raw data'!Q11","v":7.5000800999999999},{"$f":"'DF raw data'!R11","v":1.4298},{"$f":"'DF raw data'!V11","v":-2.4087200000000002},{"$f":"'DF raw data'!W11","v":7.4584498000000004},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.5759734033333337},{"$f":"K21","v":4.6800100000000157E-2},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-2.1257199999999998},{"$f":"'DF raw data'!C12","v":13.003819999999999},{"$f":"'DF raw data'!G12","v":-2.0887300999999998},{"$f":"'DF raw data'!H12","v":11.99447},{"$f":"'DF raw data'!L12","v":1.3855999999999999},{"$f":"'DF raw data'!M12","v":7.4009198999999999},{"$f":"'DF raw data'!Q12","v":1.24244},{"$f":"'DF raw data'!R12","v":1.58575},{"$f":"'DF raw data'!V12","v":-8.4417294999999992},{"$f":"'DF raw data'!W12","v":1.7544200000000001},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-3.2935500000000002},{"$f":"'DF raw data'!C13","v":-17.843188999999999},{"$f":"'DF raw data'!G13","v":0.19535000999999999},{"$f":"'DF raw data'!H13","v":0.10432},{"$f":"'DF raw data'!L13","v":-2.47194},{"$f":"'DF raw data'!M13","v":-2.2794398999999999},{"$f":"'DF raw data'!Q13","v":-4.5613998999999996},{"$f":"'DF raw data'!R13","v":-9.9406099000000001},{"$f":"'DF raw data'!V13","v":-0.45785000999999997},{"$f":"'DF raw data'!W13","v":-3.8789598999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":15.021570000000001},{"$f":"'DF raw data'!C14","v":4.6991500999999998},{"$f":"'DF raw data'!G14","v":0.88419998},{"$f":"'DF raw data'!H14","v":5.4109302000000001},{"$f":"'DF raw data'!L14","v":2.3873300999999998},{"$f":"'DF raw data'!M14","v":7.4356799000000002},{"$f":"'DF raw data'!Q14","v":-1.93808},{"$f":"'DF raw data'!R14","v":9.1021003999999994},{"$f":"'DF raw data'!V14","v":-8.3740702000000002},{"$f":"'DF raw data'!W14","v":5.3782902000000004},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":3.3660100000000002},{"$f":"'DF raw data'!C15","v":-10.711130000000001},{"$f":"'DF raw data'!G15","v":3.1920700000000002},{"$f":"'DF raw data'!H15","v":-5.2733797999999998},{"$f":"'DF raw data'!L15","v":13.367979999999999},{"$f":"'DF raw data'!M15","v":-4.0662699},{"$f":"'DF raw data'!Q15","v":0.55097996999999999},{"$f":"'DF raw data'!R15","v":-0.14077000000000001},{"$f":"'DF raw data'!V15","v":-1.8959999999999999},{"$f":"'DF raw data'!W15","v":-1.35893},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-0.60676998000000004},{"$f":"'DF raw data'!C16","v":1.68005},{"$f":"'DF raw data'!G16","v":-1.14625},{"$f":"'DF raw data'!H16","v":0.88958000999999998},{"$f":"'DF raw data'!L16","v":-1.78891},{"$f":"'DF raw data'!M16","v":-1.7601100000000001},{"$f":"'DF raw data'!Q16","v":2.2085599999999999},{"$f":"'DF raw data'!R16","v":-2.8910401000000001},{"$f":"'DF raw data'!V16","v":-3.7793100000000002},{"$f":"'DF raw data'!W16","v":2.9851201000000001}]
[6,{"$f":"'DF raw data'!B17","v":5.1009402000000001},{"$f":"'DF raw data'!C17","v":-9.5909901000000009},{"$f":"'DF raw data'!G17","v":-0.33925},{"$f":"'DF raw data'!H17","v":-2.9060600000000001},{"$f":"'DF raw data'!L17","v":-3.4977800999999999},{"$f":"'DF raw data'!M17","v":-7.6403398999999999},{"$f":"'DF raw data'!Q17","v":0.41644998999999999},{"$f":"'DF raw data'!R17","v":-1.6794800000000001},{"$f":"'DF raw data'!V17","v":-2.1207799999999999},{"$f":"'DF raw data'!W17","v":4.4977898999999999}]
[8,{"$f":"'DF raw data'!B18","v":-2.1432899999999999},{"$f":"'DF raw data'!C18","v":-4.8645500999999998},{"$f":"'DF raw data'!G18","v":-6.1020497999999996},{"$f":"'DF raw data'!H18","v":-21.350059999999999},{"$f":"'DF raw data'!L18","v":-5.3624802000000003},{"$f":"'DF raw data'!M18","v":-8.4564695000000007},{"$f":"'DF raw data'!Q18","v":-4.1889500999999996},{"$f":"'DF raw data'!R18","v":-7.8342900000000002},{"$f":"'DF raw data'!V18","v":-1.0652900000000001},{"$f":"'DF raw data'!W18","v":5.0336800000000004}]
[10,{"$f":"'DF raw data'!B19","v":-27.396681000000001},{"$f":"'DF raw data'!C19","v":2.1192801000000001},{"$f":"'DF raw data'!G19","v":-7.2850399000000001},{"$f":"'DF raw data'!H19","v":-13.206950000000001},{"$f":"'DF raw data'!L19","v":-7.1146301999999997},{"$f":"'DF raw data'!M19","v":-9.2704897000000006},{"$f":"'DF raw data'!Q19","v":-7.0103201999999998},{"$f":"'DF raw data'!R19","v":-3.4681498999999998},{"$f":"'DF raw data'!V19","v":-15.932370000000001},{"$f":"'DF raw data'!W19","v":5.0955399999999997}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":903},null,{"$f":"'DF raw data'!G4","v":904},null,{"$f":"'DF raw data'!L4","v":905},null,{"$f":"'DF raw data'!Q4","v":906},null,{"$f":"'DF raw data'!V4","v":907}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":5.0313433333333331},{"$f":"AVERAGE(C11:C13)","v":-7.951722966666666},1.4238733300000002,8.0623466666666907E-2,4.4277900333333333,0.36332336666666681,-1.98283331,-0.32642650000000023,-3.5759734033333337,4.6800100000000157E-2]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":4460446},null,{"$f":"'DF raw data'!G7","v":4440399},null,{"$f":"'DF raw data'!L7","v":4420385},null,{"$f":"'DF raw data'!Q7","v":4380373},null,{"$f":"'DF raw data'!V7","v":4100376},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-7.7117245999999993},{"$f":"C21","v":-14.511816133333332},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":14.805109133333334},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":14.747427966666665}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":7.0933845333333343},{"$f":"E21","v":0.23561183333333316},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":14.805109133333334},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":14.747427966666665}]
[-10,{"$f":"'DF raw data'!D9/100","v":56.950063},{"$f":"'DF raw data'!E9/100","v":-1.4873144999999999},{"$f":"'DF raw data'!I9/100","v":24.034951},{"$f":"'DF raw data'!J9/100","v":25.952871000000002},{"$f":"'DF raw data'!N9/100","v":10.801388999999999},{"$f":"'DF raw data'!O9/100","v":3.7110059},{"$f":"'DF raw data'!S9/100","v":-21.339326},{"$f":"'DF raw data'!T9/100","v":19.768350000000002},{"$f":"'DF raw data'!X9/100","v":19.361279},{"$f":"'DF raw data'!Y9/100","v":17.391923999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":4.6782707666666665},{"$f":"G21","v":-3.4156738733333327}]
[-8,{"$f":"'DF raw data'!D10/100","v":30.237734},{"$f":"'DF raw data'!E10/100","v":23.252206999999999},{"$f":"'DF raw data'!I10/100","v":9.1471605999999994},{"$f":"'DF raw data'!J10/100","v":21.920165999999998},{"$f":"'DF raw data'!N10/100","v":8.5871375000000008},{"$f":"'DF raw data'!O10/100","v":10.186152},{"$f":"'DF raw data'!S10/100","v":11.882190000000001},{"$f":"'DF raw data'!T10/100","v":28.150342000000002},{"$f":"'DF raw data'!X10/100","v":-15.441412},{"$f":"'DF raw data'!Y10/100","v":4.0346776999999996},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.2188532666666667},{"$f":"I21","v":-0.23220693333333373}]
[-6,{"$f":"'DF raw data'!D11/100","v":54.644858000000006},{"$f":"'DF raw data'!E11/100","v":29.440830000000002},{"$f":"'DF raw data'!I11/100","v":29.680212000000001},{"$f":"'DF raw data'!J11/100","v":24.551035000000002},{"$f":"'DF raw data'!N11/100","v":9.3280560000000001},{"$f":"'DF raw data'!O11/100","v":12.949580000000001},{"$f":"'DF raw data'!S11/100","v":29.727537000000002},{"$f":"'DF raw data'!T11/100","v":15.483730000000001},{"$f":"'DF raw data'!X11/100","v":5.4701275999999996},{"$f":"'DF raw data'!Y11/100","v":21.968739999999997},null,{"$f":"J4","v":100},{"$f":"J21","v":-2.9945688000000001},{"$f":"K21","v":-4.5786297033333332}]
[-4,{"$f":"'DF raw data'!D12/100","v":-6.3945276},{"$f":"'DF raw data'!E12/100","v":21.127694999999999},{"$f":"'DF raw data'!I12/100","v":13.654824000000001},{"$f":"'DF raw data'!J12/100","v":31.064032999999998},{"$f":"'DF raw data'!N12/100","v":-0.95217239000000009},{"$f":"'DF raw data'!O12/100","v":25.081338000000002},{"$f":"'DF raw data'!S12/100","v":-7.4826514},{"$f":"'DF raw data'!T12/100","v":33.287725000000002},{"$f":"'DF raw data'!X12/100","v":-12.671351},{"$f":"'DF raw data'!Y12/100","v":27.305635000000002}]
[-2,{"$f":"'DF raw data'!D13/100","v":9.6867932000000003},{"$f":"'DF raw data'!E13/100","v":-59.206981999999996},{"$f":"'DF raw data'!I13/100","v":3.1874252000000003},{"$f":"'DF raw data'!J13/100","v":-5.8576757999999991},{"$f":"'DF raw data'!N13/100","v":-2.0534792999999998},{"$f":"'DF raw data'!O13/100","v":-20.682597999999999},{"$f":"'DF raw data'!S13/100","v":-1.5908783},{"$f":"'DF raw data'!T13/100","v":-13.994413999999999},{"$f":"'DF raw data'!X13/100","v":-3.2965706999999997},{"$f":"'DF raw data'!Y13/100","v":-0.43400390999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-27.472707999999997},{"$f":"'DF raw data'!E14/100","v":22.058018000000001},{"$f":"'DF raw data'!I14/100","v":2.6042694000000002},{"$f":"'DF raw data'!J14/100","v":11.793524999999999},{"$f":"'DF raw data'!N14/100","v":4.0178696},{"$f":"'DF raw data'!O14/100","v":10.098594},{"$f":"'DF raw data'!S14/100","v":1.99207},{"$f":"'DF raw data'!T14/100","v":16.299550999999997},{"$f":"'DF raw data'!X14/100","v":-13.190760000000001},{"$f":"'DF raw data'!Y14/100","v":2.2067188},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-5.349259},{"$f":"'DF raw data'!E15/100","v":-6.3864844000000005},{"$f":"'DF raw data'!I15/100","v":15.488459000000001},{"$f":"'DF raw data'!J15/100","v":-5.2290137000000003},{"$f":"'DF raw data'!N15/100","v":12.070422000000001},{"$f":"'DF raw data'!O15/100","v":0.33698238000000003},{"$f":"'DF raw data'!S15/100","v":-4.0577515000000002},{"$f":"'DF raw data'!T15/100","v":-3.0017577999999996},{"$f":"'DF raw data'!X15/100","v":7.5036243000000002},{"$f":"'DF raw data'!Y15/100","v":-15.508604},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":2.0658734000000001},{"$f":"'DF raw data'!E16/100","v":8.6569921999999995},{"$f":"'DF raw data'!I16/100","v":-2.5640393000000001},{"$f":"'DF raw data'!J16/100","v":3.2124706999999999},{"$f":"'DF raw data'!N16/100","v":6.2996581999999997},{"$f":"'DF raw data'!O16/100","v":8.0229687999999992},{"$f":"'DF raw data'!S16/100","v":9.9551183999999999},{"$f":"'DF raw data'!T16/100","v":7.9924609000000002},{"$f":"'DF raw data'!X16/100","v":-18.419021000000001},{"$f":"'DF raw data'!Y16/100","v":17.982147999999999}]
[6,{"$f":"'DF raw data'!D17/100","v":29.891577000000002},{"$f":"'DF raw data'!E17/100","v":-23.270479000000002},{"$f":"'DF raw data'!I17/100","v":-2.5833819999999998},{"$f":"'DF raw data'!J17/100","v":-8.7550585999999999},{"$f":"'DF raw data'!N17/100","v":7.2384698000000007},{"$f":"'DF raw data'!O17/100","v":-3.8583398},{"$f":"'DF raw data'!S17/100","v":11.199290000000001},{"$f":"'DF raw data'!T17/100","v":1.8444921999999999},{"$f":"'DF raw data'!X17/100","v":2.4544635000000001},{"$f":"'DF raw data'!Y17/100","v":5.2328710999999997}]
[8,{"$f":"'DF raw data'!D18/100","v":-3.2052618000000002},{"$f":"'DF raw data'!E18/100","v":11.214209},{"$f":"'DF raw data'!I18/100","v":-3.4260455000000003},{"$f":"'DF raw data'!J18/100","v":-25.602256000000001},{"$f":"'DF raw data'!N18/100","v":-16.095768},{"$f":"'DF raw data'!O18/100","v":-2.0461230000000001},{"$f":"'DF raw data'!S18/100","v":-3.0913601999999996},{"$f":"'DF raw data'!T18/100","v":13.689463},{"$f":"'DF raw data'!X18/100","v":3.7142635999999998},{"$f":"'DF raw data'!Y18/100","v":24.812314000000001}]
[10,{"$f":"'DF raw data'!D19/100","v":-67.914990000000003},{"$f":"'DF raw data'!E19/100","v":-5.0437891000000006},{"$f":"'DF raw data'!I19/100","v":-7.9714984000000007},{"$f":"'DF raw data'!J19/100","v":-34.31165},{"$f":"'DF raw data'!N19/100","v":-5.0643079000000002},{"$f":"'DF raw data'!O19/100","v":-18.734102},{"$f":"'DF raw data'!S19/100","v":-6.9353026999999994},{"$f":"'DF raw data'!T19/100","v":-3.6864941},{"$f":"'DF raw data'!X19/100","v":-28.259524000000003},{"$f":"'DF raw data'!Y19/100","v":-18.804110999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":903},null,{"$f":"'DF raw data'!G4","v":904},null,{"$f":"'DF raw data'!L4","v":905},null,{"$f":"'DF raw data'!Q4","v":906},null,{"$f":"'DF raw data'!V4","v":907}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-7.7117245999999993},{"$f":"AVERAGE(C11:C13)","v":-14.511816133333332},7.0933845333333343,0.23561183333333316,4.6782707666666665,-3.4156738733333327,-1.2188532666666667,-0.23220693333333373,-2.9945688000000001,-4.5786297033333332]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16331, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16331, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16331, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16331, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16331, 100 mm gap"]
["Date:","6. okt 2017","08:27:28",null,null,"Date:","5. okt 2017","16:27:02",null,null,"Date:","5. okt 2017","18:36:33",null,null,"Date:","5. okt 2017","18:49:24",null,null,"Date:","6. okt 2017","08:27:28"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",903,null,null,null,"Indice:",904,null,null,null,"Indice:",905,null,null,null,"Indice:",906,null,null,null,"Indice:",907]
["Background Indice:",903,null,null,null,"Background Indice:",904,null,null,null,"Background Indice:",905,null,null,null,"Background Indice:",906,null,null,null,"Background Indice:",907]
["Temperature:",21.86,null,null,null,"Temperature:",21.77,null,null,null,"Temperature:",21.78,null,null,null,"Temperature:",21.68,null,null,null,"Temperature:",21.69]
["Encoder:",4460446,null,null,null,"Encoder:",4440399,null,null,null,"Encoder:",4420385,null,null,null,"Encoder:",4380373,null,null,null,"Encoder:",4100376]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,9.0254697999999998,14.898910000000001,5695.0063,-148.73145,-1,-2.0175098999999999,14.129,2403.4951000000001,2595.2871,-1,1.77145,8.9104404000000006,1080.1388999999999,371.10059000000001,-1,-7.0439800999999997,4.5016198000000003,-2133.9326000000001,1976.835,-1,-1.2188300000000001,9.6121902000000006,1936.1279,1739.1923999999999]
[-0.80000000999999998,13.71156,-0.66443998000000004,3023.7734,2325.2206999999999,-0.80000000999999998,8.9245195000000006,8.9199103999999991,914.71605999999997,2192.0165999999999,-0.80000000999999998,0.92080998000000003,-2.7620800000000001,858.71375,1018.6152,-0.80000000999999998,3.3050499000000002,0.96265000000000001,1188.2190000000001,2815.0342000000001,-0.80000000999999998,-9.0067395999999995,1.4834501,-1544.1412,403.46776999999997]
[-0.60000001999999997,7.9510198000000001,25.194050000000001,5464.4858000000004,2944.0830000000001,-0.60000001999999997,5.9669198999999997,12.623329999999999,2968.0212000000001,2455.1035000000002,-0.60000001999999997,1.2092700000000001,8.3177996000000007,932.80560000000003,1294.9580000000001,-0.60000001999999997,7.5000800999999999,1.4298,2972.7537000000002,1548.373,-0.60000001999999997,-2.4087200000000002,7.4584498000000004,547.01275999999996,2196.8739999999998]
[-0.40000001000000002,-2.1257199999999998,13.003819999999999,-639.45276000000001,2112.7694999999999,-0.40000001000000002,-2.0887300999999998,11.99447,1365.4824000000001,3106.4032999999999,-0.40000001000000002,1.3855999999999999,7.4009198999999999,-95.217239000000006,2508.1338000000001,-0.40000001000000002,1.24244,1.58575,-748.26513999999997,3328.7725,-0.40000001000000002,-8.4417294999999992,1.7544200000000001,-1267.1351,2730.5635000000002]
[-0.2,-3.2935500000000002,-17.843188999999999,968.67931999999996,-5920.6981999999998,-0.2,0.19535000999999999,0.10432,318.74252000000001,-585.76757999999995,-0.2,-2.47194,-2.2794398999999999,-205.34792999999999,-2068.2597999999998,-0.2,-4.5613998999999996,-9.9406099000000001,-159.08783,-1399.4413999999999,-0.2,-0.45785000999999997,-3.8789598999999999,-329.65706999999998,-43.400390999999999]
[0,15.021570000000001,4.6991500999999998,-2747.2707999999998,2205.8018000000002,0,0.88419998,5.4109302000000001,260.42694,1179.3525,0,2.3873300999999998,7.4356799000000002,401.78696000000002,1009.8594000000001,0,-1.93808,9.1021003999999994,199.20699999999999,1629.9550999999999,0,-8.3740702000000002,5.3782902000000004,-1319.076,220.67187999999999]
[0.2,3.3660100000000002,-10.711130000000001,-534.92589999999996,-638.64844000000005,0.2,3.1920700000000002,-5.2733797999999998,1548.8459,-522.90137000000004,0.2,13.367979999999999,-4.0662699,1207.0422000000001,33.698238000000003,0.2,0.55097996999999999,-0.14077000000000001,-405.77515,-300.17577999999997,0.2,-1.8959999999999999,-1.35893,750.36243000000002,-1550.8604]
[0.40000001000000002,-0.60676998000000004,1.68005,206.58734000000001,865.69921999999997,0.40000001000000002,-1.14625,0.88958000999999998,-256.40393,321.24707000000001,0.40000001000000002,-1.78891,-1.7601100000000001,629.96582000000001,802.29687999999999,0.40000001000000002,2.2085599999999999,-2.8910401000000001,995.51184000000001,799.24608999999998,0.40000001000000002,-3.7793100000000002,2.9851201000000001,-1841.9021,1798.2148]
[0.60000001999999997,5.1009402000000001,-9.5909901000000009,2989.1577000000002,-2327.0479,0.60000001999999997,-0.33925,-2.9060600000000001,-258.33819999999997,-875.50585999999998,0.60000001999999997,-3.4977800999999999,-7.6403398999999999,723.84698000000003,-385.83398,0.60000001999999997,0.41644998999999999,-1.6794800000000001,1119.9290000000001,184.44922,0.60000001999999997,-2.1207799999999999,4.4977898999999999,245.44635,523.28710999999998]
[0.80000000999999998,-2.1432899999999999,-4.8645500999999998,-320.52618000000001,1121.4209000000001,0.80000000999999998,-6.1020497999999996,-21.350059999999999,-342.60455000000002,-2560.2256000000002,0.80000000999999998,-5.3624802000000003,-8.4564695000000007,-1609.5768,-204.6123,0.80000000999999998,-4.1889500999999996,-7.8342900000000002,-309.13601999999997,1368.9463000000001,0.80000000999999998,-1.0652900000000001,5.0336800000000004,371.42635999999999,2481.2314000000001]
[1,-27.396681000000001,2.1192801000000001,-6791.4989999999998,-504.37891000000002,1,-7.2850399000000001,-13.206950000000001,-797.14984000000004,-3431.165,1,-7.1146301999999997,-9.2704897000000006,-506.43079,-1873.4102,1,-7.0103201999999998,-3.4681498999999998,-693.53026999999997,-368.64940999999999,1,-15.932370000000001,5.0955399999999997,-2825.9524000000001,-1880.4111]