{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":143,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386933},null,{"$f":"'DF raw data'!G7","v":487489},null,{"$f":"'DF raw data'!L7","v":587241},null,{"$f":"'DF raw data'!Q7","v":787168},null,{"$f":"'DF raw data'!V7","v":2186989},null,null,{"$f":"B4","v":10},{"$f":"B21","v":8.2352730666666663},{"$f":"C21","v":-13.700210333333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":9.9700964633333324},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":7.3220970666666663}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-0.87310659999999984},{"$f":"E21","v":-8.7614867333333333},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":11.742086566666666},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":8.4756703333333334}]
[-10,{"$f":"'DF raw data'!B9","v":-14.97559},{"$f":"'DF raw data'!C9","v":-4.9862799999999998},{"$f":"'DF raw data'!G9","v":-4.6874199000000001},{"$f":"'DF raw data'!H9","v":-6.6470900000000004},{"$f":"'DF raw data'!L9","v":3.4340002000000001E-2},{"$f":"'DF raw data'!M9","v":-3.4840700999999998},{"$f":"'DF raw data'!Q9","v":-1.17943},{"$f":"'DF raw data'!R9","v":-4.9147501},{"$f":"'DF raw data'!V9","v":-0.82550000999999995},{"$f":"'DF raw data'!W9","v":-0.89825999999999995},null,{"$f":"F4","v":20},{"$f":"F21","v":-0.64507329666666668},{"$f":"G21","v":-6.3781132666666664}]
[-8,{"$f":"'DF raw data'!B10","v":13.39006},{"$f":"'DF raw data'!C10","v":-8.7949800000000007},{"$f":"'DF raw data'!G10","v":-4.8155998999999996},{"$f":"'DF raw data'!H10","v":-8.0960196999999994},{"$f":"'DF raw data'!L10","v":0.24071001},{"$f":"'DF raw data'!M10","v":-8.1861896999999999},{"$f":"'DF raw data'!Q10","v":-2.7597100999999999},{"$f":"'DF raw data'!R10","v":-3.1984300999999999},{"$f":"'DF raw data'!V10","v":-2.4169198999999999},{"$f":"'DF raw data'!W10","v":0.14963999},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.7348233966666669},{"$f":"I21","v":-7.4792232666666676},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-8.6817101999999995},{"$f":"'DF raw data'!C11","v":-17.113430000000001},{"$f":"'DF raw data'!G11","v":-14.882199999999999},{"$f":"'DF raw data'!H11","v":-5.5401601999999999},{"$f":"'DF raw data'!L11","v":-2.7519100000000001},{"$f":"'DF raw data'!M11","v":-5.6615099999999998},{"$f":"'DF raw data'!Q11","v":-1.5382899999999999},{"$f":"'DF raw data'!R11","v":-4.4498701000000001},{"$f":"'DF raw data'!V11","v":-1.96715},{"$f":"'DF raw data'!W11","v":-2.1907698999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.5068134999999998},{"$f":"K21","v":-5.2245400000000002},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-6.7670897999999999},{"$f":"'DF raw data'!C12","v":-14.171810000000001},{"$f":"'DF raw data'!G12","v":1.0783100000000001},{"$f":"'DF raw data'!H12","v":-3.4222701},{"$f":"'DF raw data'!L12","v":-4.3057198999999997},{"$f":"'DF raw data'!M12","v":-12.391690000000001},{"$f":"'DF raw data'!Q12","v":-1.9920001E-2},{"$f":"'DF raw data'!R12","v":-5.8362002000000004},{"$f":"'DF raw data'!V12","v":-8.9434605000000005},{"$f":"'DF raw data'!W12","v":-2.3641700999999999},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":1.9028700999999999},{"$f":"'DF raw data'!C13","v":-10.109439999999999},{"$f":"'DF raw data'!G13","v":-3.4736099},{"$f":"'DF raw data'!H13","v":-11.83028},{"$f":"'DF raw data'!L13","v":-2.1435298999999999},{"$f":"'DF raw data'!M13","v":-6.9230799999999997},{"$f":"'DF raw data'!Q13","v":1.88422},{"$f":"'DF raw data'!R13","v":-13.84843},{"$f":"'DF raw data'!V13","v":1.5425899999999999},{"$f":"'DF raw data'!W13","v":-5.5815200999999997},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":6.7992300999999999},{"$f":"'DF raw data'!C14","v":-18.470381},{"$f":"'DF raw data'!G14","v":-1.12073},{"$f":"'DF raw data'!H14","v":-6.9392899999999997},{"$f":"'DF raw data'!L14","v":-0.495},{"$f":"'DF raw data'!M14","v":-4.6118898000000002},{"$f":"'DF raw data'!Q14","v":-7.4295502000000004},{"$f":"'DF raw data'!R14","v":-2.9727800000000002},{"$f":"'DF raw data'!V14","v":-9.5497903999999991},{"$f":"'DF raw data'!W14","v":-3.1992600000000002},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":16.003719},{"$f":"'DF raw data'!C15","v":-12.520810000000001},{"$f":"'DF raw data'!G15","v":1.9750201000000001},{"$f":"'DF raw data'!H15","v":-7.5148902},{"$f":"'DF raw data'!L15","v":0.70331001000000004},{"$f":"'DF raw data'!M15","v":-7.5993700000000004},{"$f":"'DF raw data'!Q15","v":0.34086000999999999},{"$f":"'DF raw data'!R15","v":-5.6164598000000003},{"$f":"'DF raw data'!V15","v":-2.5132401},{"$f":"'DF raw data'!W15","v":-6.8928399000000002},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":8.8404197999999994},{"$f":"'DF raw data'!C16","v":-9.4395904999999996},{"$f":"'DF raw data'!G16","v":3.6509299},{"$f":"'DF raw data'!H16","v":-7.4595599000000004},{"$f":"'DF raw data'!L16","v":7.4782299999999999},{"$f":"'DF raw data'!M16","v":-6.5159301999999997},{"$f":"'DF raw data'!Q16","v":-1.4336899999999999},{"$f":"'DF raw data'!R16","v":-5.1015201000000001},{"$f":"'DF raw data'!V16","v":-2.8066399},{"$f":"'DF raw data'!W16","v":-2.5860099999999999}]
[6,{"$f":"'DF raw data'!B17","v":2.57057},{"$f":"'DF raw data'!C17","v":-0.10091},{"$f":"'DF raw data'!G17","v":4.5738301000000003},{"$f":"'DF raw data'!H17","v":-2.5337600999999998},{"$f":"'DF raw data'!L17","v":-1.62297},{"$f":"'DF raw data'!M17","v":-3.2243400000000002},{"$f":"'DF raw data'!Q17","v":-3.3176600999999999},{"$f":"'DF raw data'!R17","v":-4.9355897999999998},{"$f":"'DF raw data'!V17","v":-1.4091199999999999},{"$f":"'DF raw data'!W17","v":-1.11093}]
[8,{"$f":"'DF raw data'!B18","v":3.5577198999999999},{"$f":"'DF raw data'!C18","v":0.74836999000000004},{"$f":"'DF raw data'!G18","v":-0.14503000999999999},{"$f":"'DF raw data'!H18","v":-3.3697400000000002},{"$f":"'DF raw data'!L18","v":-2.6793000999999999},{"$f":"'DF raw data'!M18","v":0.95789999000000003},{"$f":"'DF raw data'!Q18","v":-4.5813898999999996},{"$f":"'DF raw data'!R18","v":-1.9738800999999999},{"$f":"'DF raw data'!V18","v":-6.1530001000000001E-2},{"$f":"'DF raw data'!W18","v":-2.4602100999999998}]
[10,{"$f":"'DF raw data'!B19","v":-3.6413701000000001},{"$f":"'DF raw data'!C19","v":-0.87312000999999995},{"$f":"'DF raw data'!G19","v":1.3877701},{"$f":"'DF raw data'!H19","v":0.81182997999999995},{"$f":"'DF raw data'!L19","v":3.0899600999999999},{"$f":"'DF raw data'!M19","v":2.76159},{"$f":"'DF raw data'!Q19","v":-1.9483999999999999},{"$f":"'DF raw data'!R19","v":-1.78843},{"$f":"'DF raw data'!V19","v":-0.42926999999999998},{"$f":"'DF raw data'!W19","v":-5.3123598000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":973},null,{"$f":"'DF raw data'!G4","v":974},null,{"$f":"'DF raw data'!L4","v":975},null,{"$f":"'DF raw data'!Q4","v":976},null,{"$f":"'DF raw data'!V4","v":977}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":8.2352730666666663},{"$f":"AVERAGE(C11:C13)","v":-13.700210333333333},-0.87310659999999984,-8.7614867333333333,-0.64507329666666668,-6.3781132666666664,-1.7348233966666669,-7.4792232666666676,-3.5068134999999998,-5.2245400000000002]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386933},null,{"$f":"'DF raw data'!G7","v":487489},null,{"$f":"'DF raw data'!L7","v":587241},null,{"$f":"'DF raw data'!Q7","v":787168},null,{"$f":"'DF raw data'!V7","v":2186989},null,null,{"$f":"B4","v":10},{"$f":"B21","v":2.5743785333333329},{"$f":"C21","v":-17.428213066666665},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":7.8626441666666675},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":15.503066566666664}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-5.2882656333333342},{"$f":"E21","v":-9.4774087333333323},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":9.4873554000000002},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":15.503066566666664}]
[-10,{"$f":"'DF raw data'!D9/100","v":-29.284094},{"$f":"'DF raw data'!E9/100","v":-17.602080000000001},{"$f":"'DF raw data'!I9/100","v":1.2134180000000001},{"$f":"'DF raw data'!J9/100","v":24.014618999999996},{"$f":"'DF raw data'!N9/100","v":2.1976544000000002},{"$f":"'DF raw data'!O9/100","v":7.6217772999999998},{"$f":"'DF raw data'!S9/100","v":-9.2737780999999995},{"$f":"'DF raw data'!T9/100","v":1.1029395},{"$f":"'DF raw data'!X9/100","v":-5.0144115999999999},{"$f":"'DF raw data'!Y9/100","v":2.7294043000000001},null,{"$f":"F4","v":20},{"$f":"F21","v":-2.0371886333333333},{"$f":"G21","v":-6.4879882333333327}]
[-8,{"$f":"'DF raw data'!D10/100","v":9.3878168000000013},{"$f":"'DF raw data'!E10/100","v":-8.1737401999999992},{"$f":"'DF raw data'!I10/100","v":-14.885714999999999},{"$f":"'DF raw data'!J10/100","v":-3.9900878999999998},{"$f":"'DF raw data'!N10/100","v":-5.9906775000000003},{"$f":"'DF raw data'!O10/100","v":-1.9479004},{"$f":"'DF raw data'!S10/100","v":-8.8378510000000006},{"$f":"'DF raw data'!T10/100","v":1.2272460999999999},{"$f":"'DF raw data'!X10/100","v":-11.574539},{"$f":"'DF raw data'!Y10/100","v":-1.0324316},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.6752783999999998},{"$f":"I21","v":-1.9251465000000001}]
[-6,{"$f":"'DF raw data'!D11/100","v":-29.678872000000002},{"$f":"'DF raw data'!E11/100","v":-12.592695000000001},{"$f":"'DF raw data'!I11/100","v":-42.016869999999997},{"$f":"'DF raw data'!J11/100","v":0.17142579999999999},{"$f":"'DF raw data'!N11/100","v":-9.7115033000000004},{"$f":"'DF raw data'!O11/100","v":-8.6846875000000008},{"$f":"'DF raw data'!S11/100","v":-20.829622000000001},{"$f":"'DF raw data'!T11/100","v":-3.3843164000000003},{"$f":"'DF raw data'!X11/100","v":-9.9722106999999998},{"$f":"'DF raw data'!Y11/100","v":-12.008778999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":-6.9129768666666678},{"$f":"K21","v":-9.1272820666666661}]
[-4,{"$f":"'DF raw data'!D12/100","v":-27.758573999999999},{"$f":"'DF raw data'!E12/100","v":-3.8225878999999998},{"$f":"'DF raw data'!I12/100","v":-30.148125},{"$f":"'DF raw data'!J12/100","v":7.6540819999999998},{"$f":"'DF raw data'!N12/100","v":-28.842593000000001},{"$f":"'DF raw data'!O12/100","v":-14.20665},{"$f":"'DF raw data'!S12/100","v":-12.725841000000001},{"$f":"'DF raw data'!T12/100","v":-2.1315234000000003},{"$f":"'DF raw data'!X12/100","v":-32.245556999999998},{"$f":"'DF raw data'!Y12/100","v":0.47137699}]
[-2,{"$f":"'DF raw data'!D13/100","v":-12.402457},{"$f":"'DF raw data'!E13/100","v":-2.2835741999999999},{"$f":"'DF raw data'!I13/100","v":-13.179056000000001},{"$f":"'DF raw data'!J13/100","v":-7.1606348000000004},{"$f":"'DF raw data'!N13/100","v":-3.6725983000000002},{"$f":"'DF raw data'!O13/100","v":-2.8629784999999996},{"$f":"'DF raw data'!S13/100","v":4.7968871999999996},{"$f":"'DF raw data'!T13/100","v":-9.0481055000000001},{"$f":"'DF raw data'!X13/100","v":1.9816364},{"$f":"'DF raw data'!Y13/100","v":-0.28482420000000003},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":1.5019826000000001},{"$f":"'DF raw data'!E14/100","v":-38.733145},{"$f":"'DF raw data'!I14/100","v":-6.8455145000000002},{"$f":"'DF raw data'!J14/100","v":-13.852440999999999},{"$f":"'DF raw data'!N14/100","v":-3.9030939},{"$f":"'DF raw data'!O14/100","v":-10.113994},{"$f":"'DF raw data'!S14/100","v":-11.224713999999999},{"$f":"'DF raw data'!T14/100","v":3.1753418},{"$f":"'DF raw data'!X14/100","v":-12.600405},{"$f":"'DF raw data'!Y14/100","v":-13.217139},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":18.623609999999999},{"$f":"'DF raw data'!E15/100","v":-11.267919999999998},{"$f":"'DF raw data'!I15/100","v":4.1597735999999994},{"$f":"'DF raw data'!J15/100","v":-7.4191503999999995},{"$f":"'DF raw data'!N15/100","v":1.4641263},{"$f":"'DF raw data'!O15/100","v":-6.4869921999999995},{"$f":"'DF raw data'!S15/100","v":1.4019916000000001},{"$f":"'DF raw data'!T15/100","v":9.73242E-2},{"$f":"'DF raw data'!X15/100","v":-10.120162000000001},{"$f":"'DF raw data'!Y15/100","v":-13.879883},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":14.869523999999998},{"$f":"'DF raw data'!E16/100","v":7.0217187999999995},{"$f":"'DF raw data'!I16/100","v":5.0703722999999998},{"$f":"'DF raw data'!J16/100","v":-3.6238184000000002},{"$f":"'DF raw data'!N16/100","v":14.606824},{"$f":"'DF raw data'!O16/100","v":2.7140233999999999},{"$f":"'DF raw data'!S16/100","v":-16.483605000000001},{"$f":"'DF raw data'!T16/100","v":-0.71675781000000005},{"$f":"'DF raw data'!X16/100","v":-5.5067345999999997},{"$f":"'DF raw data'!Y16/100","v":-0.2155859}]
[6,{"$f":"'DF raw data'!D17/100","v":15.840569},{"$f":"'DF raw data'!E17/100","v":11.554247999999999},{"$f":"'DF raw data'!I17/100","v":-6.6530651999999995},{"$f":"'DF raw data'!J17/100","v":5.8840137000000006},{"$f":"'DF raw data'!N17/100","v":-16.242398999999999},{"$f":"'DF raw data'!O17/100","v":-1.3966015999999999},{"$f":"'DF raw data'!S17/100","v":-11.789902},{"$f":"'DF raw data'!T17/100","v":-11.811063999999998},{"$f":"'DF raw data'!X17/100","v":-8.2986707000000006},{"$f":"'DF raw data'!Y17/100","v":-10.457284999999999}]
[8,{"$f":"'DF raw data'!D18/100","v":-13.143851},{"$f":"'DF raw data'!E18/100","v":18.968223000000002},{"$f":"'DF raw data'!I18/100","v":-12.341078},{"$f":"'DF raw data'!J18/100","v":-4.0101952999999995},{"$f":"'DF raw data'!N18/100","v":-10.211793},{"$f":"'DF raw data'!O18/100","v":4.1353026999999996},{"$f":"'DF raw data'!S18/100","v":-20.782800000000002},{"$f":"'DF raw data'!T18/100","v":10.538740000000001},{"$f":"'DF raw data'!X18/100","v":-5.8093451000000007},{"$f":"'DF raw data'!Y18/100","v":-1.569375}]
[10,{"$f":"'DF raw data'!D19/100","v":-31.647368},{"$f":"'DF raw data'!E19/100","v":2.7432324000000001},{"$f":"'DF raw data'!I19/100","v":-14.129042999999999},{"$f":"'DF raw data'!J19/100","v":8.8965624999999999},{"$f":"'DF raw data'!N19/100","v":-1.0871032999999999},{"$f":"'DF raw data'!O19/100","v":17.781230000000001},{"$f":"'DF raw data'!S19/100","v":-18.233428},{"$f":"'DF raw data'!T19/100","v":4.5048145000000002},{"$f":"'DF raw data'!X19/100","v":-14.18835},{"$f":"'DF raw data'!Y19/100","v":-5.6599316000000002}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":973},null,{"$f":"'DF raw data'!G4","v":974},null,{"$f":"'DF raw data'!L4","v":975},null,{"$f":"'DF raw data'!Q4","v":976},null,{"$f":"'DF raw data'!V4","v":977}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":2.5743785333333329},{"$f":"AVERAGE(C11:C13)","v":-17.428213066666665},-5.2882656333333342,-9.4774087333333323,-2.0371886333333333,-6.4879882333333327,-1.6752783999999998,-1.9251465000000001,-6.9129768666666678,-9.1272820666666661]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16352, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16352, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16352, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16352, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16352, 100 mm gap"]
["Date:","20. dec 2017","17:36:54",null,null,"Date:","20. dec 2017","17:06:41",null,null,"Date:","20. dec 2017","17:20:59",null,null,"Date:","20. dec 2017","17:36:54",null,null,"Date:","20. dec 2017","18:33:12"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",973,null,null,null,"Indice:",974,null,null,null,"Indice:",975,null,null,null,"Indice:",976,null,null,null,"Indice:",977]
["Background Indice:",973,null,null,null,"Background Indice:",974,null,null,null,"Background Indice:",975,null,null,null,"Background Indice:",976,null,null,null,"Background Indice:",977]
["Temperature:",23.74,null,null,null,"Temperature:",23.7,null,null,null,"Temperature:",23.7,null,null,null,"Temperature:",23.74,null,null,null,"Temperature:",23.58]
["Encoder:",386933,null,null,null,"Encoder:",487489,null,null,null,"Encoder:",587241,null,null,null,"Encoder:",787168,null,null,null,"Encoder:",2186989]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-14.97559,-4.9862799999999998,-2928.4094,-1760.2080000000001,-1,-4.6874199000000001,-6.6470900000000004,121.34180000000001,2401.4618999999998,-1,3.4340002000000001E-2,-3.4840700999999998,219.76544000000001,762.17773,-1,-1.17943,-4.9147501,-927.37780999999995,110.29395,-1,-0.82550000999999995,-0.89825999999999995,-501.44116000000002,272.94042999999999]
[-0.80000000999999998,13.39006,-8.7949800000000007,938.78168000000005,-817.37401999999997,-0.80000000999999998,-4.8155998999999996,-8.0960196999999994,-1488.5715,-399.00878999999998,-0.80000000999999998,0.24071001,-8.1861896999999999,-599.06775000000005,-194.79004,-0.80000000999999998,-2.7597100999999999,-3.1984300999999999,-883.78510000000006,122.72461,-0.80000000999999998,-2.4169198999999999,0.14963999,-1157.4539,-103.24316]
[-0.60000001999999997,-8.6817101999999995,-17.113430000000001,-2967.8872000000001,-1259.2695000000001,-0.60000001999999997,-14.882199999999999,-5.5401601999999999,-4201.6869999999999,17.142579999999999,-0.60000001999999997,-2.7519100000000001,-5.6615099999999998,-971.15033000000005,-868.46875,-0.60000001999999997,-1.5382899999999999,-4.4498701000000001,-2082.9621999999999,-338.43164000000002,-0.60000001999999997,-1.96715,-2.1907698999999998,-997.22107000000005,-1200.8779]
[-0.40000001000000002,-6.7670897999999999,-14.171810000000001,-2775.8573999999999,-382.25878999999998,-0.40000001000000002,1.0783100000000001,-3.4222701,-3014.8125,765.40819999999997,-0.40000001000000002,-4.3057198999999997,-12.391690000000001,-2884.2593000000002,-1420.665,-0.40000001000000002,-1.9920001E-2,-5.8362002000000004,-1272.5841,-213.15234000000001,-0.40000001000000002,-8.9434605000000005,-2.3641700999999999,-3224.5556999999999,47.137698999999998]
[-0.2,1.9028700999999999,-10.109439999999999,-1240.2456999999999,-228.35741999999999,-0.2,-3.4736099,-11.83028,-1317.9056,-716.06348000000003,-0.2,-2.1435298999999999,-6.9230799999999997,-367.25983000000002,-286.29784999999998,-0.2,1.88422,-13.84843,479.68871999999999,-904.81055000000003,-0.2,1.5425899999999999,-5.5815200999999997,198.16363999999999,-28.482420000000001]
[0,6.7992300999999999,-18.470381,150.19826,-3873.3145,0,-1.12073,-6.9392899999999997,-684.55145000000005,-1385.2440999999999,0,-0.495,-4.6118898000000002,-390.30939000000001,-1011.3994,0,-7.4295502000000004,-2.9727800000000002,-1122.4713999999999,317.53417999999999,0,-9.5497903999999991,-3.1992600000000002,-1260.0405000000001,-1321.7139]
[0.2,16.003719,-12.520810000000001,1862.3610000000001,-1126.7919999999999,0.2,1.9750201000000001,-7.5148902,415.97735999999998,-741.91503999999998,0.2,0.70331001000000004,-7.5993700000000004,146.41263000000001,-648.69921999999997,0.2,0.34086000999999999,-5.6164598000000003,140.19916000000001,9.7324199999999994,0.2,-2.5132401,-6.8928399000000002,-1012.0162,-1387.9883]
[0.40000001000000002,8.8404197999999994,-9.4395904999999996,1486.9523999999999,702.17187999999999,0.40000001000000002,3.6509299,-7.4595599000000004,507.03723000000002,-362.38184000000001,0.40000001000000002,7.4782299999999999,-6.5159301999999997,1460.6823999999999,271.40233999999998,0.40000001000000002,-1.4336899999999999,-5.1015201000000001,-1648.3605,-71.675781000000001,0.40000001000000002,-2.8066399,-2.5860099999999999,-550.67345999999998,-21.558589999999999]
[0.60000001999999997,2.57057,-0.10091,1584.0569,1155.4248,0.60000001999999997,4.5738301000000003,-2.5337600999999998,-665.30651999999998,588.40137000000004,0.60000001999999997,-1.62297,-3.2243400000000002,-1624.2399,-139.66015999999999,0.60000001999999997,-3.3176600999999999,-4.9355897999999998,-1178.9902,-1181.1063999999999,0.60000001999999997,-1.4091199999999999,-1.11093,-829.86707000000001,-1045.7284999999999]
[0.80000000999999998,3.5577198999999999,0.74836999000000004,-1314.3851,1896.8223,0.80000000999999998,-0.14503000999999999,-3.3697400000000002,-1234.1078,-401.01952999999997,0.80000000999999998,-2.6793000999999999,0.95789999000000003,-1021.1793,413.53026999999997,0.80000000999999998,-4.5813898999999996,-1.9738800999999999,-2078.2800000000002,1053.874,0.80000000999999998,-6.1530001000000001E-2,-2.4602100999999998,-580.93451000000005,-156.9375]
[1,-3.6413701000000001,-0.87312000999999995,-3164.7368000000001,274.32324,1,1.3877701,0.81182997999999995,-1412.9042999999999,889.65625,1,3.0899600999999999,2.76159,-108.71033,1778.123,1,-1.9483999999999999,-1.78843,-1823.3427999999999,450.48145,1,-0.42926999999999998,-5.3123598000000003,-1418.835,-565.99315999999999]