## Slide 1
06/11/2019 Discussion of Magnetic Measurements data:
Jim Welch, Mark Woodley, and Scott Anderson
septum bend strength is OK (capable to beyond 15 GeV)
field harmonics in bending channel are OK … harmonic tolerances specified in “Cu2SXR” PRD are easily met
trim winding has little effect in field-free channel and will probably not be used during operation (septum has its own power supply)
measured residual integrated field (By and Bx) in field-free channel at maximum septum operating current is negligible … any steering effects can be handled with normal steering

## Slide 2
engineering type: 0.625SD38.98
SLAC drawing number: SA-375-150-75
SLAC magnet bar-code: 4584
vendor serial number: 2
when measured: 06-Jun-2019 13:23:47
when analyzed: 11-Jun-2019
MAD assignment: BLRCUS

## Slide 3
Emax
Imax ≈ 83 A

## Slide 4
wirevsx.ru4
@ r= 10 mm
|b1/b0| = 0.005 [0.200] %
|b2/b0| = 0.030 [5.000] %

## Slide 5
name: BLRCUS Trim
# of main coil turns: 45
# of trim coil turns: 20
ratio of turns (Nmain/Ntrim): 2.2500
ratio of slopes (measured): 2.3793
IMMO (amps): [-6,+6]
BMMO (main coil amps): [-2.5217,+2.5217]
Itrim = 0
Imain = 130 A
Imain = 0
wiredat.ru3

## Slide 6
Itrim = 0
Imain = 130 A
Imain = 130 A
wiredat.ru2
trim strength reduced by ≈ 2%
when Imain = 130 A

## Slide 7
standardize, then set Imain to 0
wiredat.r10
standardize, then set Imain to 130 A
wiredat.r11
standardize, then set Imain to 0
standardize, then set Imain to 130 A
wiredat.ru6
wiredat.ru7

## Slide 8
wiredat.ru5
wiredat.ru9
wirevsx.ru8
wirevsx.r12
BSCx
BSCy
ΔByL(4 GeV) ≈ 1.4 G-m
ΔByL(10 GeV) ≈ 2.7 G-m
ΔBxL(4 GeV) ≈ -0.1 G-m
ΔBxL(10 GeV) ≈ -0.2 G-m