## Slide 1: HGVPU final tuning results
I. Vasserman, N. Strelnikov, J. Xu
March 17, 2016

## Slide 2: Outline
HGVPU:
3.4m full length device with progressive conical spring force compensation mechanism.
Revised pole height from 1.2mm to 0.5mm elevation above the surface of magnets.
Re-calibrated spring cages.
New aluminum girder.

## Slide 3: HGVPU in MM1

## Slide 4: Forces on the actuators < +/-30 lbs across the gaps.

## Slide 5: The straightness of the poles and the gap <+/- 10 microns at gap of 7.2mm

## Slide 6: Peak field distribution in Z at the gaps of 7.2mm and 9mm

## Slide 7: Peak field distribution in Z at the gaps of 13mm and 20mm

## Slide 8: Magnetic field tuning results

## Slide 9: Gap dependent horizontal trajectories

## Slide 10: Gap dependent vertical trajectories

## Slide 11: Gap dependent RMS phase errors
Chart: RMS Phase Err
Chart: RMS Phase Err
| 20 | 15 | 13 | 11 | 9 | 8 | 7.2
Phase Err | 1.42 | 2.33 | 2.72 | 3.21 | 3.33 | 3.16 | 3.34

## Slide 12: Gap dependent effective fields and RMS phase errors
Gap (mm) | RMS Phase Err (deg) | Beff (G)
20 | 1.42 | 2019
15 | 2.33 | 3747
13 | 2.72 | 4831
11 | 3.21 | 6280
9 | 3.33 | 8273
8 | 3.16 | 9555
7.2 | 3.34 | 10725
Req | < 4.00 | >10100@7.2

## Slide 13: Gap dependent integrals and multipoles
Gap | J1x | J1y | J2x | J2y | a1 | a2 | a3 | b1 | b2 | b3
20 | 35 | -25 | 8.1 | 4.7 | 68 | 23 | -39 | 26 | 25 | -18
15 | 10 | 13 | 5.5 | 4.1 | 24 | 25 | 39 | 59 | -16 | -50
13 | -2 | 20 | 4.4 | 2.5 | 10 | 12 | -71 | 6 | -4 | -32
11 | -17 | 22 | 3.4 | 0.4 | -15 | 12 | -82 | 6 | -17 | -32
9 | -33 | 22 | 2.4 | -0.4 | -41 | -71 | 12 | -1 | 42 | -151
8 | -37 | 26 | 1.6 | 0.2 | -36 | -89 | 199 | -1 | 48 | -127
7.2 | 15 | 36 | -0.9 | 2.3 | -20 | -129 | 270 | -26 | 83 | -171
Req | +/- 40 | +/- 40 | +/- 15.0 | +/- 15.0 | +/- 100 | +/- 200 | +/- 400 | +/- 100 | +/- 200 | +/- 400
Units: Gap (mm); J1x (G-cm); J1y (G-cm); J2x (G-cm2); J2y (G-cm2)
a1/b1 (G); a2/b2 (G/cm); a3/b3 (G/cm2)

## Slide 14: Conclusion.
Gap straightness at 7.2mm is better than +/- 10 microns.
No observable lower order deflections of the strongbacks at different operational gap settings.
RMS phase errors of the device are better than the requirement of 4 degrees.
Effective field is over 1.07 T at the gap of 7.2mm, that is better than the requirement of 1.01 T.
First and second integrals and multipoles are also better than the requirements.
Reproducibility and environment test results show no changes to the device performance. The difference in RMS phase errors are better than 0.1 degree, that is less than 2 micron changes in taper.

## Slide 15: Magnetic shim map
Left Jaw
from front
(near side)
Right Jaw
from back
(far side)
Left Jaw
from front
(near side)
Right Jaw
from back
(far side)
4mm Side x1
Flat 0.2mm x1
5mm 0.2mm x1
P5/ P4/ P3
4mm Side x1
Flat 0.2mm x1
5mm 0.2mm x1
Phase x1
P57/P56/P55
Phase x1
4mm Side x1
P62
4mm Side x1
P76
4mm Side x1
P76
4mm Side x1
40mm 0.2 x1
40mm 0.2 x1
P62
P78/P77
Phase x1
Phase x1
P124
4mm Side x1
P124
4mm Side x1
P131
30mm 0.2 trj x2
P131
30mm 0.2 trj x2
P146
2mm Side x1
P146
2mm Side x1
P162
2mm Side x1
P162
2mm Side x1
P165
40mm 0.2 x1
P165
2mm Side x1
2mm Side x1
P177
2mm Side x1
P177
2mm Side x1
P215
2mm Side x1
P215
2mm Side x1
P239
4mm Side x1
P239
4mm Side x1
P258/P257/P256/P255/P254/P253/P252/P251/P250
P57/P56/P55
P5/ P4/ P3
P78/P77
P169/P168/P167/P166
4mm Side x4
20mm 0.2 x4
8mm Side x1
P169/P168/P167/P166
40mm 0.2 x1
Phase 0.2 x2
2mm Side x2
P258/P257/P256/P255/P254/P253/P252/P251/P250
15mm 0.2 x1
4mm Side x4
20mm 0.2 x4
8mm Side x1
2mm Side x2
15mm 0.2 x1

## Slide 16: Attachments

## Slide 17: Attachments

## Slide 18: Attachments

## Slide 19: Attachments