Taper test

The taper was changed with a single PV:  USEG:MMF:1:TAPER_DES

The taper is set in mm, both positive and negative. It is limited to +-1mm by limit switches. The pivot point for the taper is the first pole of the undulator, as on the picture. The average gap is bigger than the nominal gap for positive taper and smaller for negative.

Nominal Gap

Bottom jaw

Nominal Gap + Taper

Top jaw

Top jaw

A set of measurements was done to study the effect of changing the taper on the undulator parameters.

The nominal gap was 10mm.

The field integrals were measured with the long coil for positive taper only, as in the table below.

| Taper(µm) | I1X(µTm) | I1Y(µTm) | I2X(µTm2) | I2Y(µTm2) |
| --- | --- | --- | --- | --- |
| 0 (start) | +14 | -8 | +14 | +59 |
| +100 | +13 | -3 | +12 | +84 |
| +300 | +14 | -2 | +16 | +81 |
| 0 (end) | +15 | -4 | +10 | +76 |
| Comment | No change | Small change | No change | Small change |

Summary: The field integrals change is small, just above the measurement accuracy.

The Hall probe measurements were done for 10mm gap also, both for positive and negative tapers. The magnetic axis position and the pitch were corrected for 10mm gap values. Before each Zscan measurement, the Bscany measurements were done to check changes in magnetic axis position and pitch. Results are in the table below.

| Taper(µm) | YMA(m) | Pitch(µrad) | Ph.Err. RMS() | K | K/K |
| --- | --- | --- | --- | --- | --- |
| 0 (start) | 0.000166 | 0.0 | 2.8 | 4.215545 | 0 |
| +300 | 0.000164 | +3.3 | 56 | 4.147550 | -1.610-2 |
| -300 | 0.000200 | -1.6 | 59 | 4.286648 | +1.710-2 |
| 0 (end) | 0.000172 | +11.5 | 2.9 | 4.216078 | 1.310-4 |
| Comment | 34µm change | Small change | 60 change |  |  |

The measurements data at: S:\magdata\LCLS-II\Undulator\SXU_004\DATASET0001\Tuning\Z Scans\Taper_Study\...