## Slide 1: Measurement and Assembly Procedure for DELTA Undulator
Ed Reese
FDR, Dec. 11, 2013

## Slide 2: Focus of Presentation
This presentation will focus on the areas of the assembly process for the Delta Undulator were the measurement of key components and assemblies will ensure that the tight tolerances needed for the device to function correctly are maintained. It will not provide a step by step narrative for the assembly of the device.

## Slide 3: Measurement of all large components will be done on the LEITZ CMM in the MMF
CMM has measuring range of
X axis = 4500 mm
Y axis = 1200 mm
Z axis = 900 mm
Volumetric Measurement Uncertainty = 1.5+l/350
All components fit on CMM

## Slide 4: Machined Strongback E Sections

## Slide 5: Critical Strongback E Section Measurements
± 0.1 µm/mm
Red surface areas need to be machined precisely
Parallelism < ± 25 µm/m

## Slide 6: Machined Carrier

## Slide 7: Critical Dimensions on Carrier
Red surface areas need to be machined precisely
Parallelism < ± 25 µm/m
Beam axis
Angular tolerance = ± 0.1 µm/mm
Angular tolerance
= ± 0.5 µm/mm

## Slide 8: Carrier First Piece Validation
Measure Critical Dimensions on first Carrier Machined.
Magnet Mounting and Adjustment Features are not
Machine at this point.
Evaluate and Verify Location of Magnet Mounting Plane in Relation to Rail Mounting Surfaces.
If Mounting Plane is in correct location no additional machining steps needed.
If not in correct location, reevaluate process, consider measuring each Carrier before final machining to determine corrections needed.

## Slide 9: Test Assemble Strongback with Rails and Carriers
Strongback is test assembled
With Rails and Carriers, but without
Magnets, to verify final assembly
dimensions

## Slide 10: Attach Drive Units to Strongback Carrier Assembly
Drive units are attached to Strongback E sections with Carriers and Rails installed to
test fit and operation.
Drive Unit

## Slide 11: Fit Mounting Feet to joined Strongback Sections
Mounting Feet are attached to joined Strongback sections to calculate machining needed to correctly position Undulator centerline. Values calculated from archive data for Girder in position #33.

## Slide 12: Verify critical dimensions on Calibration Fixture
The critical surfaces on the Calibration Fixture have the same dimensions and tolerances as one half of the Strongback E section. The fixture is inclined
45 deg so that the magnets, when installed will be in a vertical orientation.
Critical Surfaces

## Slide 13: Install Carrier/Rail assemblies on Fixture and verify alignment of reference surfaces for Keyence sensors
Reference Surface
Reference Surface will be used to adjust pitch of fixture on Kuglar Bench

## Slide 14: Install Magnets onto Carriers

## Slide 15: Move Tuning Fixture to CMM and measure Tip circle location of each magnet
The Tip of each magnet is scanned by the CMM. A circle of fixed diameter is fit to the scanned point set giving the center location of the magnet.

## Slide 16: Evaluate mechanical alignment of Magnets installed on Carrier
The center points all magnets are collected and their deviation in the YZ and XZ planes of the nominal centerline are calculated

## Slide 17: Tuning Fixture moved to Kuglar Bench for Magnetic Measurement
After initial CMM measurement the Tuning Fixture with Magnet/Carriers installed is moved to the Kuglar Bench for Magnetic Measurement.
After Magnetic Measurement the magnets may be adjusted with shims in the Y direction or shifted with set screws in the X direction. This process may go through several iterations.
After Magnetic Measurement is complete the Tuning fixture will be returned to the CMM for a final measurement on magnet positions.

## Slide 18: Installation of Carrier assemblies to Strongback

## Slide 19: Installing Lower Carrier assemblies
Rigging one magnet carrier
Translation stage
connected to the strong back
Support
Strong back

## Slide 20: Installing Upper Carrier assemblies

## Slide 21: Joining fully assembled Strongback halves together
Strongback sections
Sliding Fixtures
The two Strongback sections will be mounted onto sliding fixtures. These fixture will be used to control the magnetic forces as the two halves are brought together.