## Slide 1: Moving Wire, Pulsed Wire and Long Coil Measurements on the SSRL Beam Line 5 EPU
Scott Anderson
SLAC Magnetic Measurements

## Slide 2: Moving Wire
100 mm Single CuBe Wire 5.45 meters in length.
Newport ILS stages with 0.1 mm resolution linear scales.
Agilent 3458A Digital voltmeter.
Wire is moved a known distance at each point. Voltage signal is sampled during the movement and then integrated. Voltage offset is corrected.
Voltage offset is important since integrated voltage levels are on the nanovolt-sec level.
Movement is +/-250 mm, 6 times back and forth centered at each point.

## Slide 3: Moving Wire Setup Upstream End

## Slide 4: Moving Wire Setup Downstream End

## Slide 5: Moving Wire Signal

## Slide 6: Moving Wire and Danfysik 1st Integral Measurement Comparisons at 13 mm gap.
Danfysik measurements were made using a 5 mm flip coil and the background subtraction was done using 120 mm (fully open) gap measurements.
Background subtraction was done using 120 mm (fully open) gap measurements.
4 modes are compared CP+, CP-, VP+ and HP modes.

## Slide 7: Bx 1St Integral CP+ Mode

## Slide 8: Bx 1St Integral CP+ Mode at Y +/-2.5, 2, 1 and 0 mm

## Slide 9: By 1St Integral CP+ Mode

## Slide 10: Bx 1St Integral CP- Mode

## Slide 11: Bx 1St Integral CP- Mode at Y +/-2.5, 2, 1 and 0 mm

## Slide 12: By 1St Integral CP- Mode

## Slide 13: VP+ Mode Bx 1st Integral

## Slide 14: Bx 1St Integral VP+ Mode at Y +/-2.5, 2, 1 and 0 mm

## Slide 15: By 1St Integral VP+ Mode

## Slide 16: Bx 1St Integral HP Mode

## Slide 17: Bx 1St Integral HP Mode at Y +/-2.5, 2, 1 and 0 mm

## Slide 18: By 1St Integral HP Mode

## Slide 19: Cross Check of Moving Wire Data
Measurements satisfy Maxwell’s equation dBx/dy = dBy/dx
  VP+ mode: dBx/dy = 0.62, dBy/dx = 0.58, within 7.3%
  VP- mode: dBx/dy = 0.61, dBy/dx = 0.67, within 7.9%
  CP+ mode: dBx/dy = 0.24, dBy/dx = 0.25, within 4%
  CP- mode: dBx/dy = 0.24, dBy/dx = 0.24, within 0.2%
  HP mode: dBx/dy = -0.038 , dBy/dx = -0.031 , within 28%
  Pulsed Wire Measurements were also made as a check.
  First experience using Pulsed Wire with an undulator

## Slide 20: Pulsed Wire Setup
100 mm Single CuBe Wire 5.45 meters in length.
HV pulser. 5 msec square pulse at 600 VDC
Laser/ 25 mm Slit Detector measures sub micron motion of wire.
Signal readout on Oscilloscope.
Voltage-Field Integral calibration made by using calibration magnets.
Large noise is caused by vibration of supports.
Cross check of moving wire measurements

## Slide 21: Pulsed Wire Signal with Calibration Magnet

## Slide 22: Pulsed Wire Noise

## Slide 23: Bx 1St Integral CP- Mode Pulsed Wire

## Slide 24: By 1St Integral CP- Mode Pulsed Wire

## Slide 25: Bx 1St Integral VP+ Mode Pulsed Wire

## Slide 26: By 1St Integral VP+ Mode Pulsed Wire

## Slide 27: Z2 Jaw Control Problem
On February 6th Z2 jaw motor driver found to be not moving and sitting at its zero point.
Motor driver was tripped due to overheating.
After motor was restored subsequent measurements of Bx 1st integrals for CP+, CP-, VP+, VP- changed from previous.
Z2 jaw was set to zero point on purpose and the measurements made in this state matched the measurements made previous to the Z2 motor drive being reset.

## Slide 28: Z2 Jaw at 0 and nominal, CP- mode

## Slide 29: Z2 Jaw at 0 and nominal, VP- mode

## Slide 30: Long Coil Measurements
A long coil was originally used to measure the integrals of the EPU.
The long coil has 350 turns and gives larger signals than a single moving wire.
The By 1st integrals match Danfysik data, but the Bx 1st integral measurements do not, except for the HP mode. The CP and VP mode long coil Bx 1st integrals show large integral values away from x = 0.

## Slide 31: Long Coil Setup

## Slide 32: Long Coil Bx 1St Integral CP- Mode

## Slide 33: Long Coil By 1St Integral CP- Mode

## Slide 34: Long Coil Bx 1St Integral CP+ Mode

## Slide 35: Long Coil By 1St Integral CP+ Mode

## Slide 36: Long Coil Bx 1St Integral HP Mode

## Slide 37: Long Coil By 1St Integral HP Mode

## Slide 38: Long Coil Bx integral Investigation
Long Coil back on measurement bench.
Working to determine source of the Long Coil Bx integral errors for the CP and VP modes by systematically checking each component of the system.
  Swap each component of the long coil system with a moving wire component to isolate the source of the error in the long coil system.

## Slide 39: Discussion