## Slide 1: Magnetic Measurements And Magnet Fiducialization At SLAC
Zachary Wolf
SLAC
August 25, 2015

## Slide 2: Overview
Field harmonics in a multipole magnet
Rotating coil
Field maps
Hall probe, NMR probe, pulsed wire, short rotating coil
Field integrals
Moving wire
Fiducialization
Dipoles
Quadrupoles, Sextupoles, etc.
How to locate a wire
Solenoids
Undulators
How to locate a Hall probe

## Slide 3: Field In A Multipole Magnet
R = ‘reference radius’
Assume ∂z = 0

## Slide 4: Voltage From A Rotating Coil
Main harmonic

## Slide 5: Effect Of Rotation Errors
Coil motion error signals look like harmonics
Solution: integrate voltage to remove speed variation effects (use flux vs angle),
buck out main harmonic, only measure small harmonics
Eg. Bad bearing
Eg. Speed variations

## Slide 6: Bucking Coil
Example from Tanabe, “Iron Dominated Electromagnets”

## Slide 7: Integrator
Trigger with encoder.
Measure Φ(θ).
No coil speed dependence.
Calculate field harmonics
From flux harmonics.
Digital integrator

## Slide 8: Rotating Coil Measurements
Measuring magnetic crosstalk between magnets
Bucking Coil
Encoder
Slip Rings
Motor
S. Anderson

## Slide 9: Example: Harmonics Measurement Of A Quadrupole

## Slide 10: Hall Probes, NMR Probes
NMR
accurate,
large sample size,
requires uniform field
Hall Probe
small sample size
nonlinear, offset drift, gain drift, …
needs lots of calibration
NMR
General Purpose Hall Probe
Custom Hall probe for Delta undulator

## Slide 11: Hall Probe Calibration
Y. Levashov

## Slide 12: Field Maps With Hall/NMR Probes
Dipole magnet.
Use Hall probe at ends.
Use NMR in center.
Apply small correction
to end fields using NMR
in center.

## Slide 13: Undulator Hall Probe Measurements
Delta undulator quadrant
LCLS undulator

## Slide 14: Pulsed Wire
Initial impulse
Travelling wave
Derivative of wire position
signal gives field

## Slide 15: Field Map From A Pulsed Wire
Signal from an undulator.
Calibration magnet puts in
step at the end.
When the wave hits the
end of the wire, it reflects.
S. Anderson

## Slide 16: Field Integrals From A Moving Wire
V+ = dΦ/dt + Vt
V- = -dΦ/dt + Vt
VT = (VT+ - VT-) / 2
= ΔΦ
Take out thermal emf:

## Slide 17: Fiducialization
Suppose we know the beam line coordinates Xb and Yb
in the telescope’s coordinate system.
We know Δx1, Δx2, Δy1, and Δy2 from the fiducialization.
We put the tooling balls at Xb+ Δx1, Xb+ Δx2, Yb+ Δy1, Yb+ Δy2
in order to place the magnet on the beam line.
In this way, we don’t need to know the details of the magnet’s
magnetic field in the tunnel.

## Slide 18: Dipoles Mechanical Fiducialization
Steel dipoles can be fiducialized mechanically.
A CMM can measure the position of the poles
and relate the pole positions to the tooling ball
positions.
In steel magnets, the magnetic field shape is
‘determined’ by the steel.
Typically the CMM can only touch the poles near the ends and the middle
is ignored. Adequate for short magnets. Also, the poles have some roughness.
Roll accuracy estimate: ~25 µm / 5 cm = 5 x 10-4 rad

## Slide 19: Dipoles Mechanical Fiducialization
Capacitive sensors can measure a reference pole
that has been fiducialized, and then measure the
magnet. Tooling balls on the magnet can be related
to tooling balls on the reference pole.
Sensors
Reference Pole
Electronic Level
Position Detector
Measures over entire length of C-magnets. Need to correct for
sensor position and roll.
Roll accuracy estimate: ~10 µm / 10 cm = 1 x 10-4 rad

## Slide 20: Dipoles Magnetic Measurement, Moving Wire
Move wire
vertically,
measure ΔΦ.
Rotate magnet
until ΔΦ=0.
Move wire
horizontally,
measure ΔΦ.
Move magnet
until ΔΦ profile
is centered on
wire.
Similar for finding
vertical center.
Relate magnet position to wire.
Roll
X
Roll accuracy estimate: ~10 µTm / 1 Tm = 1 x 10-5 rad

## Slide 21: Quadrupoles And Higher Multipoles Mechanical Fiducialization
CMM probe touches poles
to measure point locations
on the poles.
An algorithm finds the mechanical
center from the sample points.
The mechanical center is related to
tooling balls on the magnet.
Touch poles at quadrupole ends. The middle is ignored. The poles have a roughness.
Quadrupole center position accuracy estimate: ~50 µm.

## Slide 22: Quadrupoles And Higher Multipoles Moving Wire Fiducialization
Move the wire to the location where ∫By dz = 0 and ∫Bx dz = 0.
Locate the wire relative to the tooling balls.
(May need to take out pitch and yaw for some applications.)
Quadrupole:
Bx = G y
By = G x
Quadrupole center position accuracy estimate: ~10 µTm / 100 T = small.
Accuracy dominated by locating wire and locating tooling balls.

## Slide 23: Pulsed Wire, Vibrating Wire
The wire has a force on it and moves when a current flows through
the wire in the presence of a magnetic field.
*

## Slide 24: Quadrupoles And Higher Multipoles Vibrating Wire

## Slide 25: Vibrating Wire
If the ac current carrying wire goes
through a magnetic field, the force
excites a mechanical vibration of
the wire.
An optical detector gives
a signal that the wire is vibrating.
The vibration stops when the wire
is at the center of the quadrupole.
This method can also detect pitch and
yaw of the quadrupole.

## Slide 26: Vibrating Wire
Equation of motion of the wire
Solution for the wire motion

## Slide 27: Wire Vibration Detector
Optical sensors
detect wire vibration.

## Slide 28: Vibrating Wire Signals
Examples:
Amplitude of signal from wire vibration sensor vs quadrupole x-position
and quadrupole yaw.

## Slide 29: Quadrupoles Rotating Coil

## Slide 30: Quadrupoles And Higher Multipoles Rotating Coil
Locate axis of rotation from shaft ends.
Locate tooling balls relative to axis of rotation.
Harmonic feed down gives magnetic center
location relative to axis of rotation.
From these, calculate tooling ball location
relative to magnetic center.
Can also use short coils.
Move coil until dipole signal
goes to zero.

## Slide 31: How To Locate A Wire
Find the center of the wire by
finding its edges.
Calibrate the scale offset by
flipping the detector in a
kinematic mount.

## Slide 32: Solenoid Fiducialization Pitch and Yaw
A small rotating coil in the center of
the solenoid can be used to align the
field to the axis of rotation.

## Slide 33: Solenoid Fiducialization X and Y
At the end of the solenoid, the rotating coil gives
zero voltage out when the axis of rotation is on
the magnetic center line.

## Slide 34: Undulator Mechanical Alignment
Capacitive sensors can be used to mechanically align a hybrid undulator
to a reference pole. Alternatively, a CMM can be used.

## Slide 35: Undulator Magnetic Alignment
After undulator is tuned:
X trajectory good.
Y trajectory good.
Phase errors good.
Want the beam to go where
the Hall probe is measuring.
But where is the Hall probe
actually measuring?
Fiducialization tolerance
40 microns.
Probe position given to
100 microns in specs.

## Slide 36: Where is the Hall probe measuring?
Attach to undulator ends, measure offsets
Pointed magnets give a magnetic location that can be found repeatably.

## Slide 37: Pointed Magnet Calibration
Flip magnet in fixture. Distance zero field
point moves gives zero field position
relative to tooling balls.
Yurii Levashov

## Slide 38: Undulator Fiducialization
Want D
D1 : Magnetic Measurement
D2 : Calibration
D3 : CMM
D = D1 + D2 + D3

## Slide 39: Undulator Fiducialization

## Slide 40: Check Fiducialization With Simple Techniques
Check fiducialization.
Use telescopes, pointed magnets,
optical sensor to locate Hall probe

## Slide 41: End