Document Title:	SLAC Traveler for S20 XC1EX Corrector Magnet
Author(s):	 	Phoebe Alva Rosa, Doug Storey

…Revision History

| Revision | Date Released | Description of Change |
| --- | --- | --- |
| R0 | July 25, 2025 | Original Release. |

This traveler is intended to cover reception, preparation, and magnetic measurements of the sector 20 corrector magnet XC1EX. This magnet will replace the existing XCOR:LI20:3276 corrector.

This traveler has been adapted from the FACET BC20BE and BC20CE dipole magnet travelers.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MM group:

| Received by (initials): | SDA |
| --- | --- |
| Date received (mm-dd-yyyy): | 09-30-2025 |

Preparation:

A beam direction arrow, with text “beam direction”, is to be applied as shown in Figure 1.

| Beam-direction arrow, barcode and MAD Name stickers in place (initials): | SDA |
| --- | --- |
| MAD name | XC1EX |
| Barcode | 4786 |

Magnetic Measurements:

- Verify that the magnets are complete and undamaged, including wiring connections.

| Incoming inspection OK (initials): | SDA |
| --- | --- |
| Date of arrival to mag. meas.(mmm-dd-yyyy): | 09-31-2025 |

Enter URL of on-line magnetic measurements data (please modify or correct if necessary):

http://www-group.slac.stanford.edu/met/MagMeas/MagData/FACET-II/Corr/4786

- Measure the inductance and resistance of the main magnet coils:

| Inductance of main coil (mH): | 3.13 mH |
| --- | --- |
| Resistance of main coil (Ohms): | 0.3890 Ohm |

- Determine the main-coil connection polarity (with main supply outputting positive current) which produces a “positive” field polarity for dipoles as shown in Figure 1 below.

Figure 1: Dipoles have “positive” polarity when the electron beam is bent back towards the “wall-side” (+X) of the accelerator housing (as opposed to aisle-side).

- Mark the polarity near the main magnet leads with clear “+” and “−” labels on/near the terminal block.

| Main coil polarity set to Positive from Figure 1 is (P): | Positive |
| --- | --- |

- Connect the main magnet terminals in the correct polarity, as established above, to a bipolar power supply with a current range of ±12 A. Measure the temperature of both coils and the core at 12 amps for 7 hours.

| Top Coil Temperature (oC) | 48.6 |
| --- | --- |
| Bottom Coil Temperature (oC) | 46.8 |
| Core Temperature (oC) | 34.4 |
| Ambient Temperature (oC) | 21.3 |

- Standardize the magnet from -12 to 12 to -12 amps 3 times using a linear ramp of 2 A/s, with a 10 second wait time.  Use the same ramp characteristics for trimming the magnet too.

- Maintaining this cycle history, measure the length-integrated vertical dipole field, ∫Bydl, from -12 A to 12 A in 1-A steps, including zero (25 ‘up’ measurements).  Please record (below) the magnetic field strength (in kG-m) at ± 12 A. Then, still maintaining the cycle history, measure ∫Bydl back down from 12 A to -12 A in 1-A steps, including zero (25 ‘down’ measurements).

| Field Strength at +12 A | 0.06268 kG-m |
| --- | --- |
| Field Strength at -12 A | 0.06223 kG-m |
| Filename & run number of ∫Bydl up & down data: | Wiredat.ru1 |

- With stretched wire, measure the vertical length-integrated field component over a horizontal span of ±10 mm in 2 mm intervals at 12 A.

| Filename & run # of ∫Bydl vs. x data at 12 A: | Wirevsx.ru2 |
| --- | --- |

- Upon completion of tests, email URL of on-line data to Doug Storey (dstorey@slac.stanford.edu) for determination if the magnet is accepted, with Phoebe Alva Rosa (aphoebe@slac.stanford.edu) in cc. Upon acceptance of magnet, analysis data will be placed in the on-line data folder.

| Magnet accepted: | Yes |
| --- | --- |
| Assigned beamline location (MAD-deck name): | XC1EX |
| Magnet marked with assigned MAD-deck name (initials): | SDA |