## Slide 1
engineering type: 1.182D6.82T
SLAC drawing number: SA-380-300-00
SLAC magnet bar-code: F2-001
vendor serial number: 1
when measured: 27-Feb-2018 14:08:57
when analyzed: 05-Jun-2018
MAD assignment: BX10751

## Slide 2
|b1/b0| ≈ 0.0456 % > 0.02 %
|b2/b0| ≈ 0.1886 % < 3.6 %
R = 20 mm
hardat.ru2

## Slide 3
name: BX10751 trim
# of main coil turns: 36
# of trim coil turns: 34
ratio of turns (Nmain/Ntrim): 1.0588
ratio of slopes (measured): 1.0964
IMMO (amps): [-6,6]
BMMO (main coil amps): [-5.4726, 5.4726]
Itrim = 0
Imain = 0