## Slide 1
engineering type: 0.315D13.08-C
SLAC drawing number: Stangenes 731-30822-10 r2
SLAC magnet bar-code: 4595
vendor serial number: 30822-4
when measured: 06-Nov-2025 16:10:28
when analyzed: 02-Dec-2025
MAD assignment: BCX203280
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0 (Itrim = 0):
after bipolar standardize = -15.50 +/- 0.21 G-m
after degauss = 0.113 +/- 0.016 G-m
|ΔB/B0|max = 0.48% @ X = +9 mm < 0.5%
Imain = 13 A (wirevsx.ru9)

## Slide 2
Harmonic tolerance: |ΔB/B0| < 0.5% for |X| ≤ 9 mm
wirevsx.ru9
wirevsx.r10
BCX203280 (Imain=13 A, Itrim=0)
BCX203280 (Imain=14.5 A, Itrim=0)

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
name: BCX203280T
# of main coil turns: 668
# of trim coil turns: 164
ratio of turns (Nmain/Ntrim): 4.0732
ratio of slopes (measured): 4.1658
IMMO (trim coil amps): [-6,6]
BMMO (main coil amps): [-1.4402,+1.4402]
BCX203280 main coil: BL vs I
BCX203280 trim coil: BL vs I
BCX203280 trim

## Slide 4
Imain = 0 (after degauss)
wiredat.ru7
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
BCX203280 main coil: BL vs I
BCX203280 trim coil: BL vs I
BCX203280 trim

## Slide 5
ramp type = COSINE
ramp rate = 2 A/s
settle time = 3 s
remnant BL = +0.1130 +/- 0.0163 G-m
degauss time = 268 s
remnant BL: Degauss Run Summary 4595.xlsx
BCX203280 degauss

## Slide 6
BCX203280: BL vs Imain
BCX203280: dBL/dImain vs Imain

## Slide 7
Imain=0 ; 87.5 G-m/Atrim
IVBBTRM = [0,4.1911]
Imain= 0 ; 88.4 G-m/Atrim
IVBBTRM = [0,4.1467]
Imain= +13 A ; 62.3 G-m/Atrim
IVBBTRM = [0,4.2988]
BCX203280: dBL vs Itrim
BCX203280: Imain=0 (after degauss)
BCX203280: Imain=0 (after bipolar standardize)
BCX203280: Imain=+13A (after bipolar standardize)