## Slide 1
Note: the LCLS Control System requires all dipole magnets powered in series to have positive IVB polynomials (i.e. BDES[kG-m] > 0 gives IDES[A] > 0). To that end, this dipole’s IVB (which is properly negative) has had its sign reversed both in Oracle and in EPICS.

## Slide 2
engineering type: 0.315D13.08-C
SLAC drawing number: Stangenes 731-30822-10 r2
SLAC magnet bar-code: 4597
vendor serial number: 30822-1
when measured: 21-Nov-2025 11:08:56
when reanalyzed: 03-Feb-2026
MAD assignment: BCX203284
Itrim = 0 (after bipolar standardize)
|ΔB/B0| ≤ 0.46% for |X| ≤ 9 mm
Imain = 13 A (wirevsx.ru8)
BCX203284: BL vs X
measured bipolar
operate unipolar
BTRM(BDES) = -0.07117 Amain
BTRM(IDES) = -0.2962 Atrim
for BL = 0 kG-m

## Slide 3
Harmonic tolerance: |ΔB/B0| < 0.5% for |X| ≤ 9 mm
wiredat.ru8
wiredat.ru9
BCX203284 (Imain=13 A, Itrim=0)
BCX203284 (Imain=14.5 A, Itrim=0)

## Slide 4
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
name: BCX203284T
# of main coil turns: 668
# of trim coil turns: 164
ratio of turns (Nmain/Ntrim): 4.0732
ratio of slopes (measured): 4.1622
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.4415,+1.4415]
fit to DOWN data
fit to UP data