## Slide 1
engineering type: 0.315D13.08-C
SLAC drawing number: Stangenes 731-30822-10 r2
SLAC magnet bar-code: 4598
vendor serial number: 30822-3
when measured: 03-Dec-2025 13:05:09
when analyzed: 09-Dec-2025
MAD assignment: BCX203286
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -14.28 +/- 0.23 G-m
after degauss = 0.0266 +/- 0.0273 G-m
|ΔB/B0| = 0.38% @ X = +9 mm < 0.5%
Imain = 13 A (wirevsx.ru8)
BCX203286: set on the way UP
BCX203286: BL vs X

## Slide 2
Harmonic tolerance: |ΔB/B0| < 0.5% @ |X| ≤ 9 mm
BCX203286 (Imain=13 A, Itrim=0)
BCX203286 (Imain=14.5 A, Itrim=0)

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
name: BCX203286T
# of main coil turns: 668
# of trim coil turns: 164
ratio of turns (Nmain/Ntrim): 4.0732
ratio of slopes (measured): 4.1668
IMMO (amps): [-6,6]
BMMO (main coil amps): [-1.4399,+1.4399]
BCX203286 main coil: BL vs I
BCX203286 trim coil: BL vs I
BCX203286 trim

## Slide 4
Imain = 0 (after degauss)
wiredat.ru7
|Imain| < 10 A, Itrim = 0
(after bipolar standardize)
BCX203286 main coil: BL vs I
BCX203286 trim coil: BL vs I
BCX203286 trim

## Slide 5
ramp type = COSINE
ramp rate = 2 A/s
settle time = 3 s
remnant BL = +0.0466 +/- 0.0273 G-m
degauss time = 268 s
remnant BL: Degauss Run Summary 4598.xlsx
BCX203286 degauss

## Slide 6
bhvszdat.r11
BCX203286