
[image: image1]

SLAC Traveler for HE-SXR Undulator #29A804-009 HE 
in MMF
Dataset 001
This traveler is intended to cover preparation, magnetic measurements, and fiducialization of the SXR undulator segments at MMF.   

Preparation:
· Remove all protective covers. Install wooden platform between the undulator and granite. 
· Remove plastic cover. Clean up the granite.

· Attach temperature sensors.

· Attach ground wire for capacitive sensors.

· Check linear potentiometers. Install brackets if necessary.

· Label modules from 1 to 3.

· Install coil supports.

	Vendor serial number from undulator label:
	

	Technician (initials):
	

	Date/Time:
	


· Call controls group to connect and check the control system. Check completed.

· Call alignment crew to align the undulator to the bench. Alignment done.
· Cycle gap, ~5000 gap changes (between 7.2mm and 33 mm)

 Mechanical measurements:
· Open gap to 10mm (tuning gap). Check gap straightness with the touch probe, correct if necessary by replacing magnet keepers’ shims.

· Correct encoders’ offsets to make jaws parallel and level.

· Set half-gap encoders’ offsets at 10mm gap to read 5.0000mm.

· Check compensation spring stack gaps (refer to nominal values).

· Measure every pole at 10mm gap, correct pole positions to ±20µm, if necessary.

· At 10mm gap set the midplane encoders’ offsets, so the encoders read 5.0000mm.

· Remove touch probe. Install the Hall probe, align. 

· Find the third pole z-position.                                  Z3 =
	Engineer(initials)/date and time:
	


Tuning.
Follow the check list below, consult with the tuning test plan (LCLS-TN-22-09). 
Tuning process checklist:

	Engineer (initials):
	

	Tuning started, Date (mm-dd-yyyy):
	

	· Undulator magnetically aligned at tuning Gap = 10mm (adjust encoder offsets if necessary), results recorded into the table below.

	· MA shift vs Gap measured with Hall probe for gaps = 7.2,8.5,10,13,17,25, and 33 mm

	· Fields measured for central tuning x = ± 2mm at tuning gap=10mm

	· Upstream PM measured

	· Kugler bench OFF, Long coil set, PDI and XPS ON

	· Field integrals measured for tuning trajectories

	· Coil and stages in park position, PDI and XPS OFF, the coil removed for shimming.

	· Shim table prepared. Gap opened to 180mm.

	· Undulator trajectory shimming completed

	· Kugler bench ON and re-initialized

	· Trajectories and phase errors checked. (If additional shimming needed, repeat Hall probe measurements, coil measurements and tuning).

	· Trajectories and phase errors tuned to tolerances.

	· Fields measured for tuning ends.

	· Kugler bench OFF. Long coil set, PDI and XPS ON

	· Field integrals measured for tuning first and second field integrals.

	· Shim table prepared. Gap opened to 180mm. The coil removed for shimming. 


	· Field integrals tuning done.

	· Long coil set back, PDI and XPS ON.

	· Undulator field integrals checked. If additional shimming needed, repeat coil measurements and tuning.

	· Field integrals shimming completed.

	· Field integrals measured for finding commissioning gap. 

	· Coil and stages set in park position, PDI and XPS OFF, the coil removed. 

	· Kugler bench ON and re-initialized.

	· MA position vs Y checked.

	· Trajectories and phase errors checked and within the tolerances. 

	· Hall probe measurements to find commissioning gap done. (K=6.08)

	· Phase matching done.

	· Coil installed; optical measurements of coil position done

	· Field integrals vs gap checked.

· 

	· Parameters file checked before final

	· Final scan done, coil

	· Coil and stages set in park position. PDI and XPS OFF, the coil removed.

	· Gap clearance checked. Gap =7.350mm


	· Kugler bench ON and re-initialized.

	· Yaw and Pitch checked

	· Reference magnet measured.

	· MA position vs Y checked. Gap = 10mm
· 

	· Half-gap encoders checked; Gap=10mm

	· Commissioning gap checked.

	· Parameters file checked before final

	· Final scan done, Hall probe.

	· PMs measured for fiducialization, file with results created.

	· Fiducialization done.

	· Touch probe measurements at 10mm gap done.

	· Touch probe measurements with taper +100µm, +300µm, -100µm done.

	· Kugler bench OFF

	· Capacitive sensor measurements, poles #10, 15, 101, 106 done.

	· Tuning Evaluation traveler created.

	· Controller Configuration saved.


The following information is to be noted by an engineer:

	Average X (m):
	

	Average Y (m):
	

	Final Pitch (µrad):
	

	Final Yaw (µrad):
	

	Magnetic axis position X (m):
	

	Magnetic axis position Y (m):
	

	Coil X-position (m)
	

	Coil Y-position (m)
	

	Commissioning Gap (mm)
	

	Reference magnet measurement (T)
	

	US Half-Gap Encoder
	

	DS Half-Gap Encoder
	


Fiducialization:
The following information is to be noted by an engineer:

	Engineer (Initials):
	

	Upstream RM X offset (m):
	

	Upstream RM Y offset (m):
	

	Downstream RM X offset (m):
	

	Downstream RM Y offset (m):
	


Have an alignment crew to measure offsets between RM and the undulator fiducials. 
	Surveyor (initials):
	

	Date (mm/dd/yyyy):
	


Prepare the undulator for the move into the storage area:

· Power OFF the controller. Disconnect the network cable.

· Remove linear potentiometer brackets, if installed.

· Remove long coil brackets.

· Remove temperature sensors.

· Remove ground wire for capacitive sensors.

· Install all protective covers. Remove wooden platform between the undulator and granite.

· The undulator is ready for move into the storage area.

	Technician (initials) / date:
	


URL for data:

	V:\MET\MagServe\MagData\LCLS-II-HE\Undulator\HE-SXU_009\DATASET0001\Final Results\


	Undulator tuning completed (signed):
	


[image: image1][image: image2.png]


