## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2319
vendor serial number: 004
when measured: 28-Jan-2022 09:58:35
when analyzed: 17-Feb-2022
MAD assignment: BCXCBX11
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -27.15 +/- 0.02 G-m
after degauss = -0.53 +/- 0.47 G-m
pole-tip field @ magnet center B = -0.3 +/- 0.082 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.038 [0.500] %
|b2/b0| = 0.096 [1.000] %
p0 = 1.3930e-02
p1 = 3.7853e-03
p2 = -9.5847e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.027 [0.500] %
|b2/b0| = 0.062 [1.000] %
p0 = 8.9437e-03
p1 = 2.7125e-03
p2 = -6.2485e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX11T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6974
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.2986,+0.2986]

## Slide 4
After degauss, BL=0 at ITRIM = +0.1151 A
wiredat.ru6