## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2320
vendor serial number: 005
when measured: 03-Feb-2022 13:35:32
when analyzed: 19-Feb-2022
MAD assignment: BCXCBX12
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = +27.21 +/- 0.06 G-m
after degauss = +1.41 +/- 1.28 G-m
pole-tip field @ magnet center B = -0.1 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.009 [0.500] %
|b2/b0| = 0.058 [1.000] %
p0 = -1.1626e-02
p1 = -9.0286e-04
p2 = -5.8461e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.029 [0.500] %
|b2/b0| = 0.088 [1.000] %
p0 = 4.6953e-03
p1 = -2.8826e-03
p2 = -8.8238e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX12T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.8440
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.2922,+0.2922]

## Slide 4
After degauss, BL=0 at ITRIM = 0.0934 A
wiredat.ru5