## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2321
vendor serial number: 006
when measured: 08-Feb-2022 11:39:37
when analyzed: 21-Feb-2022
MAD assignment: BCXCBX13
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = +28.43 +/- 0.11 G-m
after degauss = +1.25 +/- 0.30 G-m
pole-tip field @ magnet center B = +0.45 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.001 [0.500] %
|b2/b0| = 0.075 [1.000] %
p0 = 9.2017e-03
p1 = -1.1761e-04
p2 = -7.5377e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.013 [0.500] %
|b2/b0| = 0.078 [1.000] %
p0 = 1.9325e-05
p1 = 1.3076e-03
p2 = -7.8484e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX13T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6287
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3017,+0.3017]

## Slide 4
After degauss, BL=0 at ITRIM = +0.0939 A
wiredat.ru5