## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2321
vendor serial number: 007
when measured: 15-Feb-2022 13:56:59
when analyzed: 21-Feb-2022
MAD assignment: BCXCBX14
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -27.89 +/- 0.11 G-m
after degauss = -1.42 +/- 1.19 G-m
pole-tip field @ magnet center B = -0.45 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.007 [0.500] %
|b2/b0| = 0.062 [1.000] %
p0 = 4.9706e-03
p1 = -6.8025e-04
p2 = -6.1646e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.019 [0.500] %
|b2/b0| = 0.042 [1.000] %
p0 = 1.2282e-02
p1 = 1.8614e-03
p2 = -4.2012e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru4
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX14T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6461
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3009,+0.3009]

## Slide 4
After degauss, BL=0 at ITRIM = -0.0786 A
wiredat.ru5