## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2323
vendor serial number: 008
when measured: 26-Apr-2022 10:50:27
when analyzed: 28-Apr-2022
MAD assignment: BCXCBX21
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = -26.91 +/- 0.05 G-m
after degauss = -0.02 +/- 0.09 G-m …
pole-tip field @ magnet center B = 0.15 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.004 [0.500] %
|b2/b0| = 0.066 [1.000] %
p0 = 8.7299e-03
p1 = 4.3420e-04
p2 = -6.6018e-04
wirevsx.ru1
wirevsx.ru2
BCXCBX21
p0 = -2.6413e-02
p1 = -6.3758e-04
p2 = -5.4867e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.006 [0.500] %
|b2/b0| = 0.055 [1.000] %

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX21T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.6592
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3003,+0.3003]

## Slide 4
Imain = 0 (after degauss)
wiredat.ru5
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
BCXCBX21