## Slide 1
engineering type: 0.276D14-C
SLAC drawing number: Everson Tesla
SLAC magnet bar-code: 2324
vendor serial number: 009
when measured: 16-Feb-2022 15:43:59
when analyzed: 28-Apr-2022
MAD assignment: BCXCBX22
Itrim = 0 (after bipolar standardize)
BL @ Imain = 0:
after bipolar standardize = +28.16 +/- 0.09 G-m
after degauss = +0.08 +/- 0.07 G-m …
pole-tip field @ magnet center B = 0.0 G

## Slide 2
Use harmonic tolerances from SXRSS bends
Harmonics @ r= 10 mm:
|b1/b0| = 0.010 [0.500] %
|b2/b0| = 0.075 [1.000] %
p0 = -4.9260e-03
p1 = -9.5556e-04
p2 = -7.4980e-04
Harmonics @ r= 10 mm:
|b1/b0| = 0.005 [0.500] %
|b2/b0| = 0.066 [1.000] %
p0 = -5.0019e-03
p1 = 5.1407e-04
p2 = -6.6135e-04
wirevsx.ru1
wirevsx.ru2

## Slide 3
Imain = 0 (after bipolar standardize)
wiredat.ru3
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)
name: BCXCBX22T
# of main coil turns: 234
# of trim coil turns: 69/2
ratio of turns (Nmain/Ntrim): 6.7826
ratio of slopes (measured): 6.5910
IMMO (amps): [-2,2]
BMMO (main coil amps): [-0.3034,+0.3034]

## Slide 4
Imain = 0 (after degauss)
wiredat.ru5
|Imain| < 9 A, Itrim = 0
(after bipolar standardize)