## Slide 1
engineering type: 0.433D14-C (1.575D14-C)
SLAC drawing number: ASM-20171113-6552 Red_Gap CHG_11mm_9_5_19.pdf
SLAC magnet bar-code: 4550
vendor serial number: 1
when measured: 28-Oct-2019 15:32:48
when analyzed: 07-Nov-2019
MAD assignment: BCXXL1
BL @ Imain = 0 (after stdz): -19.41 G-m

## Slide 2
quadrupole component removed
@ r= 10 mm
|b1/b0| = 0.041 [0.500] %
|b2/b0| = 0.198 [1.000] %
Imain | b1/b0 % | b2/b0 % | b3/b0 %
12.99 | 0.0789 | 0.0596 | 0.0781
17.50 | 0.0750 | 0.0603 | 0.0325
tolerance | 0.5 | 1.0 | -
Harmonic content vs Imain @ r = 1 cm
hardat.ru8

## Slide 3
after standardize
Imain = 0
slope @ Itrim = 0 : +4.1839 G-m/A
after degauss
Imain = 0
slope @ Itrim = 0 : +3.8174 G-m/A
wiredat.ru5
wiredat.ru3

## Slide 4
name: BCXXL1_Trim
# of main coil turns: 378
# of trim coil turns: 10
ratio of turns (Nmain/Ntrim): 37.8
ratio of slopes (measured): 40.012
IMMO (amps): [-6,6]
BMMO (main coil amps): [-0. 1495,+0. 1495]
Imain = 0 (after STDZ)
wiredat.ru3
Itrim = 0

## Slide 5
ramp type = COSINE
ramp rate = 2 A/s
settle time = 5 s
remnant BL = +0.7832+/- 0.0310 G-m
degauss time = 420 s
remnant BL: wiredat.ru4

## Slide 6
bhvszdat.ru6