## Slide 1
engineering type: 0.433D14-C (1.575D14-C)
SLAC drawing number: ASM-20171113-6552 Red_Gap CHG_11mm_9_5_19.pdf
SLAC magnet bar-code: 4553
vendor serial number: 4
when measured: 21-Oct-2019 09:36:14
when analyzed: 24-Oct-2019
MAD assignment: BCXXL4
BL @ Imain = 0 (after stdz): -19.85 G-m

## Slide 2
after standardize
Imain = 0
slope @ Itrim = 0 : +4.1704 G-m/A
after degauss
Imain = 0
slope @ Itrim = 0 : +3.8444 G-m/A
wiredat.ru5
wiredat.ru3

## Slide 3
name: BCXXL4_Trim
# of main coil turns: 378
# of trim coil turns: 10
ratio of turns (Nmain/Ntrim): 37.8
ratio of slopes (measured): 40.018
IMMO (amps): [-6,6]
BMMO (main coil amps): [-0. 1499,+0. 1499]
Imain = 0 (after STDZ)
wiredat.ru3
Itrim = 0

## Slide 4
ramp type = COSINE
ramp rate = 2 A/s
settle time = 5 s
remnant BL = +0.9103+/- 0.0288 G-m
degauss time = 420 s
remnant BL: wiredat.ru4