## Slide 1
engineering type: 1.0D38.37
SLAC drawing number: SA-344-100-01
SLAC magnet bar-code: 4554
vendor serial number: 15
when measured: 31-Oct-2018 15:39:57
when analyzed: 05-Nov-2018
MAD assignment: BXSP1S

## Slide 3
wirevsx.ru1
@ r= 20 mm
|b1/b0| = 0.014 [0.200] %
|b2/b0| = 0.446 [100.000] %

## Slide 4
name: BXSP1S Trim
# of main coil turns: 36
# of trim coil turns: 17
ratio of turns (Nmain/Ntrim): 2.1176
ratio of slopes (measured): 2.1888
IMMO (amps): [-6,+6]
BMMO (main coil amps): [-2.7411,+2.7411]
Itrim = 0
Imain < 150 A
Imain = 0