## Test Plan for ‘On The Curve’ Measurements using a LCLS-II 1D22.625 Dipole Magnet

(June 19, 2025)

This measurement plan is to test if staying on the standardize curve (i.e. ‘On The Curve’) is a stable method of setting the magnet to a lower current from a higher current, without doing a full standardize cycle.  The magnet to be used for the test is a PEPII Bypass B2 dipole, 1D22.625 (SA-344-100-37) which has been refurbished.

Receiving:

The following information is to be noted upon receipt of the magnets by the SLAC MM group:

| Received by (initials): | SDA |
| --- | --- |
| Date placed on test stand (dd-mmm-yyyy): | 7/15/2025 |
| SLAC barcode number: | 4565 |
| Vendor serial number from magnet label: | N/A |

- Connect the magnet terminals, in the correct polarity as established above, to a unipolar power supply with maximum current I  125 A.

- Connect magnet to LCW supply.  Flow rate should be ~0.94 gpm at 100 delta PSI.   Run the magnet at 125 A for ~1 hour to warm it up (record, delta P, flow rate, and magnet coil and steel temperature).

| LCW delta P (psi) | 93.6 psi |
| --- | --- |
| LCW flow rate (gpm) | 0.92 gpm |
| LCW delta T (°C) | 1.3 °C |
| Ambient temperature (°C): | 23.4 °C |
| Final magnet core temperature (°C): | 25.9 °C |
| Final magnet coil temperature (°C): | 29.1 °C |

- The magnet standardization is from zero to 125 A and back to zero, through three full cycles, finally ending at zero, with a flat-top pause time (at both 0 and 125 A) of 10 seconds.  Use a three linear ramp rate of 5 A/sec

- Current values to be used for the test are as follows:

| Current Setting Number | Percent of Maximum Current of 125 Amps | Current |
| --- | --- | --- |
| (#) | (%) | (Amps) |
| 1 | 70 | 87.5 |
| 2 | 60 | 75 |
| 3 | 80 | 100 |
| 4 | 30 | 37.5 |

## Normal Standardize Tests

- Standardize magnet as in step 3. Set the current to 87.5 (i.e. 70% of maximum current or current setting 1) amps and then measure the Bydl field. Set the magnet to 75 amps, standardize magnet as in step 3, and then set the current to 75 amps.  Measure the Bydl field.

- Repeat the step 5 measurements a total of 10 times.

| Filenamemof Bydl at 87.5 and 75 amps, normal standardize : | Wiredat.ru4, wireplt.ru4 |
| --- | --- |

- Standardize magnet as in step 3.  Set the current to 100 amps and then measure the Bydl field. Set the magnet to 37.5 amps,  standardize magnet as in step 3, and then set the current to 37.5 amps.  Measure the Bydl field.

- Repeat the step 7 measurements a total of 10 times.

| Filenamemof Bydl at 100 and 37.5 amps, normal standardize : | Wiredat.ru5, wireplt.ru5 |
| --- | --- |

## On the Curve Tests

- Standardize magnet as in step 3.

- Set the current to 87.5 amps and then measure the Bydl field.  Ramp the magnet to 125 amps, then to 0 amps then set the magnet to 75 amps.  Measure the Bydl field.

- Repeat the step 10 measurements a total of 10 times.

| Filename of Bydl at 87.5 and 75 amps, on the Curve : | Wiredat.ru6, wireplt.ru6 |
| --- | --- |

- Standardize magnet as in step 3.

- Set the current to 100 amps and then measure the Bydl field.  Ramp the magnet to 125 amps, then to 0 amps then set the magnet to 37.5 amps.  Measure the Bydl field.

- Repeat the step 13 measurements a total of 10 times.

| Filename of Bydl at 100 and 37.5 amps, on the Curve : |  |
| --- | --- |

- Upon completion of tests, send data link to Ben Ripman, F.J. Decker and Tim Maxwell.