{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":103,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":395529},null,{"$f":"'DF raw data'!G7","v":496028},null,{"$f":"'DF raw data'!L7","v":596356},null,{"$f":"'DF raw data'!Q7","v":796619},null,{"$f":"'DF raw data'!V7","v":2196657},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-0.18865667000000005},{"$f":"C21","v":-4.2862899666666667},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":13.377210003333333},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":2.0352501666666667}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":5.5240100000000005},{"$f":"E21","v":-3.026616666666667},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":13.377210003333333},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":6.4867300666666665}]
[-10,{"$f":"'DF raw data'!B9","v":27.934570000000001},{"$f":"'DF raw data'!C9","v":1.3626},{"$f":"'DF raw data'!G9","v":14.221209999999999},{"$f":"'DF raw data'!H9","v":2.1308999000000002},{"$f":"'DF raw data'!L9","v":15.92634},{"$f":"'DF raw data'!M9","v":7.4160199000000002},{"$f":"'DF raw data'!Q9","v":7.8275299},{"$f":"'DF raw data'!R9","v":0.55756998000000002},{"$f":"'DF raw data'!V9","v":-10.55452},{"$f":"'DF raw data'!W9","v":-12.74288},null,{"$f":"F4","v":20},{"$f":"F21","v":3.9938967000000001},{"$f":"G21","v":-4.3878334333333333}]
[-8,{"$f":"'DF raw data'!B10","v":-1.6990000000000002E-2},{"$f":"'DF raw data'!C10","v":7.2243298999999999},{"$f":"'DF raw data'!G10","v":-1.6467499999999999},{"$f":"'DF raw data'!H10","v":4.9633798999999996},{"$f":"'DF raw data'!L10","v":4.8802599999999998},{"$f":"'DF raw data'!M10","v":-2.2915800000000002},{"$f":"'DF raw data'!Q10","v":0.86504000000000003},{"$f":"'DF raw data'!R10","v":-16.539949},{"$f":"'DF raw data'!V10","v":1.9489700000000001},{"$f":"'DF raw data'!W10","v":-1.5458799999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":13.188553333333333},{"$f":"I21","v":-2.3525832666666666},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":1.6431100000000001},{"$f":"'DF raw data'!C11","v":5.5592699000000003},{"$f":"'DF raw data'!G11","v":4.7610998000000002},{"$f":"'DF raw data'!H11","v":1.58945},{"$f":"'DF raw data'!L11","v":-4.3952498000000002},{"$f":"'DF raw data'!M11","v":2.3433001},{"$f":"'DF raw data'!Q11","v":10.037559999999999},{"$f":"'DF raw data'!R11","v":-6.1542602000000004},{"$f":"'DF raw data'!V11","v":-13.202819999999999},{"$f":"'DF raw data'!W11","v":-2.0933199},null,{"$f":"J4","v":100},{"$f":"J21","v":7.4048266333333332},{"$f":"K21","v":-8.8393133333333331},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-5.9981298000000001},{"$f":"'DF raw data'!C12","v":1.81488},{"$f":"'DF raw data'!G12","v":5.5353398},{"$f":"'DF raw data'!H12","v":0.54697001000000001},{"$f":"'DF raw data'!L12","v":7.9755701999999999},{"$f":"'DF raw data'!M12","v":-2.0940398999999998},{"$f":"'DF raw data'!Q12","v":7.1664900999999999},{"$f":"'DF raw data'!R12","v":-2.8179999000000001E-2},{"$f":"'DF raw data'!V12","v":17.301860999999999},{"$f":"'DF raw data'!W12","v":11.329420000000001},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":1.3716999999999999},{"$f":"'DF raw data'!C13","v":-1.76065},{"$f":"'DF raw data'!G13","v":7.2009401000000004},{"$f":"'DF raw data'!H13","v":-2.3036300999999999},{"$f":"'DF raw data'!L13","v":7.8340502000000001},{"$f":"'DF raw data'!M13","v":-1.01772},{"$f":"'DF raw data'!Q13","v":11.29064},{"$f":"'DF raw data'!R13","v":-8.1140699000000005},{"$f":"'DF raw data'!V13","v":11.10721},{"$f":"'DF raw data'!W13","v":-6.3233499999999996},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-2.89412},{"$f":"'DF raw data'!C14","v":-4.8203000999999999},{"$f":"'DF raw data'!G14","v":2.2576599000000002},{"$f":"'DF raw data'!H14","v":1.1411099},{"$f":"'DF raw data'!L14","v":-6.7171000999999997},{"$f":"'DF raw data'!M14","v":-2.7479998999999999},{"$f":"'DF raw data'!Q14","v":12.53135},{"$f":"'DF raw data'!R14","v":10.754960000000001},{"$f":"'DF raw data'!V14","v":-4.4783201000000004},{"$f":"'DF raw data'!W14","v":-17.812550000000002},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":0.95644998999999997},{"$f":"'DF raw data'!C15","v":-6.2779198000000003},{"$f":"'DF raw data'!G15","v":7.1134300000000001},{"$f":"'DF raw data'!H15","v":-7.9173298000000001},{"$f":"'DF raw data'!L15","v":10.864739999999999},{"$f":"'DF raw data'!M15","v":-9.3977804000000003},{"$f":"'DF raw data'!Q15","v":15.74367},{"$f":"'DF raw data'!R15","v":-9.6986398999999999},{"$f":"'DF raw data'!V15","v":15.58559},{"$f":"'DF raw data'!W15","v":-2.3820399999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":3.6779999999999999},{"$f":"'DF raw data'!C16","v":-3.4892299000000002},{"$f":"'DF raw data'!G16","v":4.4216699999999998},{"$f":"'DF raw data'!H16","v":-10.7037},{"$f":"'DF raw data'!L16","v":6.9540100000000002},{"$f":"'DF raw data'!M16","v":-6.4906997999999998},{"$f":"'DF raw data'!Q16","v":17.021311000000001},{"$f":"'DF raw data'!R16","v":-3.6398999999999999},{"$f":"'DF raw data'!V16","v":-10.49193},{"$f":"'DF raw data'!W16","v":-4.6835298999999999}]
[6,{"$f":"'DF raw data'!B17","v":11.51876},{"$f":"'DF raw data'!C17","v":-3.6308701000000001},{"$f":"'DF raw data'!G17","v":15.76632},{"$f":"'DF raw data'!H17","v":-3.0899000000000001},{"$f":"'DF raw data'!L17","v":16.39649},{"$f":"'DF raw data'!M17","v":-0.29780999000000002},{"$f":"'DF raw data'!Q17","v":18.184290000000001},{"$f":"'DF raw data'!R17","v":-8.4832096000000003},{"$f":"'DF raw data'!V17","v":6.9077802000000004},{"$f":"'DF raw data'!W17","v":-6.8184699999999996}]
[8,{"$f":"'DF raw data'!B18","v":17.527820999999999},{"$f":"'DF raw data'!C18","v":-11.95119},{"$f":"'DF raw data'!G18","v":25.7013},{"$f":"'DF raw data'!H18","v":-9.3127297999999996},{"$f":"'DF raw data'!L18","v":18.706109999999999},{"$f":"'DF raw data'!M18","v":-7.3182302000000004},{"$f":"'DF raw data'!Q18","v":17.959471000000001},{"$f":"'DF raw data'!R18","v":-8.5843095999999992},{"$f":"'DF raw data'!V18","v":11.612310000000001},{"$f":"'DF raw data'!W18","v":-9.8637505000000001}]
[10,{"$f":"'DF raw data'!B19","v":28.09647},{"$f":"'DF raw data'!C19","v":-7.1880999000000001},{"$f":"'DF raw data'!G19","v":29.601410000000001},{"$f":"'DF raw data'!H19","v":-5.2975402000000003},{"$f":"'DF raw data'!L19","v":26.870398999999999},{"$f":"'DF raw data'!M19","v":2.88062},{"$f":"'DF raw data'!Q19","v":27.828130999999999},{"$f":"'DF raw data'!R19","v":-5.7245498000000001},{"$f":"'DF raw data'!V19","v":10.83212},{"$f":"'DF raw data'!W19","v":-7.5278100999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":367},null,{"$f":"'DF raw data'!G4","v":368},null,{"$f":"'DF raw data'!L4","v":369},null,{"$f":"'DF raw data'!Q4","v":370},null,{"$f":"'DF raw data'!V4","v":371}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-0.18865667000000005},{"$f":"AVERAGE(C11:C13)","v":-4.2862899666666667},5.5240100000000005,-3.026616666666667,3.9938967000000001,-4.3878334333333333,13.188553333333333,-2.3525832666666666,7.4048266333333332,-8.8393133333333331]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":395529},null,{"$f":"'DF raw data'!G7","v":496028},null,{"$f":"'DF raw data'!L7","v":596356},null,{"$f":"'DF raw data'!Q7","v":796619},null,{"$f":"'DF raw data'!V7","v":2196657},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-13.580022666666666},{"$f":"C21","v":-9.4549738333333337},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":15.277954933333334},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":17.197833833333334}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-4.0055094666666671},{"$f":"E21","v":1.6345082666666666},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":19.776204033333332},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":23.716811966666668}]
[-10,{"$f":"'DF raw data'!D9/100","v":23.581738000000001},{"$f":"'DF raw data'!E9/100","v":9.3390625000000007},{"$f":"'DF raw data'!I9/100","v":-3.1784707999999999},{"$f":"'DF raw data'!J9/100","v":6.9946680000000008},{"$f":"'DF raw data'!N9/100","v":-2.0047954999999997},{"$f":"'DF raw data'!O9/100","v":14.090693000000002},{"$f":"'DF raw data'!S9/100","v":-13.812666},{"$f":"'DF raw data'!T9/100","v":6.7921973000000007},{"$f":"'DF raw data'!X9/100","v":-47.760815000000001},{"$f":"'DF raw data'!Y9/100","v":-17.032627000000002},null,{"$f":"F4","v":20},{"$f":"F21","v":-5.4245681999999995},{"$f":"G21","v":-3.0425457333333328}]
[-8,{"$f":"'DF raw data'!D10/100","v":36.375399999999999},{"$f":"'DF raw data'!E10/100","v":25.278964999999999},{"$f":"'DF raw data'!I10/100","v":-3.9704138000000002},{"$f":"'DF raw data'!J10/100","v":10.198407999999999},{"$f":"'DF raw data'!N10/100","v":-13.880153999999999},{"$f":"'DF raw data'!O10/100","v":5.4219336},{"$f":"'DF raw data'!S10/100","v":12.762562000000001},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":5.4018933000000002},{"$f":"'DF raw data'!Y10/100","v":-30.718027000000003},null,{"$f":"H4","v":30},{"$f":"H21","v":1.6979322666666663},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":-6.3284149000000003},{"$f":"'DF raw data'!E11/100","v":-0.65848633000000012},{"$f":"'DF raw data'!I11/100","v":-10.607598000000001},{"$f":"'DF raw data'!J11/100","v":-11.388672},{"$f":"'DF raw data'!N11/100","v":-27.324016},{"$f":"'DF raw data'!O11/100","v":-13.269229000000001},{"$f":"'DF raw data'!S11/100","v":19.275380999999999},2.20269,{"$f":"'DF raw data'!X11/100","v":-27.911624000000003},{"$f":"'DF raw data'!Y11/100","v":-24.683173999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":6.1961813666666652},{"$f":"K21","v":-15.97395196666667}]
[-4,{"$f":"'DF raw data'!D12/100","v":-40.753315000000001},{"$f":"'DF raw data'!E12/100","v":6.3682128999999996},{"$f":"'DF raw data'!I12/100","v":-10.346611000000001},{"$f":"'DF raw data'!J12/100","v":8.1890526999999995},{"$f":"'DF raw data'!N12/100","v":-3.1017730999999999},{"$f":"'DF raw data'!O12/100","v":15.700986},{"$f":"'DF raw data'!S12/100","v":9.4626526000000002},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-4.9237554999999995},{"$f":"'DF raw data'!Y12/100","v":5.6611229999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":-13.997678000000001},{"$f":"'DF raw data'!E13/100","v":-12.938153999999999},{"$f":"'DF raw data'!I13/100","v":1.2905797999999999},{"$f":"'DF raw data'!J13/100","v":6.6494237999999992},{"$f":"'DF raw data'!N13/100","v":-5.5282001000000003},{"$f":"'DF raw data'!O13/100","v":15.374530999999999},{"$f":"'DF raw data'!S13/100","v":-7.7624054000000005},1.63822,{"$f":"'DF raw data'!X13/100","v":36.855390999999997},{"$f":"'DF raw data'!Y13/100","v":1.1173242000000001},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-17.410879999999999},{"$f":"'DF raw data'!E14/100","v":-6.2086816000000002},{"$f":"'DF raw data'!I14/100","v":-18.836449000000002},{"$f":"'DF raw data'!J14/100","v":13.808271},{"$f":"'DF raw data'!N14/100","v":-14.751232999999999},{"$f":"'DF raw data'!O14/100","v":-9.0538671999999991},{"$f":"'DF raw data'!S14/100","v":14.463526999999999},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":-12.205459000000001},{"$f":"'DF raw data'!Y14/100","v":-58.282061000000006},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-9.3315099999999997},{"$f":"'DF raw data'!E15/100","v":-9.2180859000000002},{"$f":"'DF raw data'!I15/100","v":5.5293407999999999},{"$f":"'DF raw data'!J15/100","v":-15.554169999999999},{"$f":"'DF raw data'!N15/100","v":4.0057285},{"$f":"'DF raw data'!O15/100","v":-15.448300999999999},{"$f":"'DF raw data'!S15/100","v":-1.6073248},16.2971,{"$f":"'DF raw data'!X15/100","v":-6.0613878999999997},{"$f":"'DF raw data'!Y15/100","v":9.2428808999999994},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-6.5662720000000006},{"$f":"'DF raw data'!E16/100","v":10.114746},{"$f":"'DF raw data'!I16/100","v":-4.2277715999999996},{"$f":"'DF raw data'!J16/100","v":-6.7178319999999996},{"$f":"'DF raw data'!N16/100","v":10.892498},{"$f":"'DF raw data'!O16/100","v":3.1858887},{"$f":"'DF raw data'!S16/100","v":-14.071959999999999},10.6371,{"$f":"'DF raw data'!X16/100","v":-71.393584000000004},{"$f":"'DF raw data'!Y16/100","v":3.4692676000000002}]
[6,{"$f":"'DF raw data'!D17/100","v":15.242627000000001},{"$f":"'DF raw data'!E17/100","v":2.0148047},{"$f":"'DF raw data'!I17/100","v":-0.85529792999999998},{"$f":"'DF raw data'!J17/100","v":2.5303613},{"$f":"'DF raw data'!N17/100","v":20.115988999999999},{"$f":"'DF raw data'!O17/100","v":3.0025293},{"$f":"'DF raw data'!S17/100","v":8.649726600000001},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":-3.6105444000000002},{"$f":"'DF raw data'!Y17/100","v":-12.841094}]
[8,{"$f":"'DF raw data'!D18/100","v":30.004058000000001},{"$f":"'DF raw data'!E18/100","v":7.6208496000000006},{"$f":"'DF raw data'!I18/100","v":32.375767000000003},{"$f":"'DF raw data'!J18/100","v":-0.22289060999999999},{"$f":"'DF raw data'!N18/100","v":33.940810999999997},{"$f":"'DF raw data'!O18/100","v":-0.25494140999999998},{"$f":"'DF raw data'!S18/100","v":0.72053336999999995},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":-8.3062292000000006},{"$f":"'DF raw data'!Y18/100","v":-10.801416}]
[10,{"$f":"'DF raw data'!D19/100","v":35.616320999999999},{"$f":"'DF raw data'!E19/100","v":-11.233145},{"$f":"'DF raw data'!I19/100","v":23.891649999999998},{"$f":"'DF raw data'!J19/100","v":-20.900039000000003},{"$f":"'DF raw data'!N19/100","v":57.275537000000007},{"$f":"'DF raw data'!O19/100","v":19.240840000000002},{"$f":"'DF raw data'!S19/100","v":26.958696},13.3657,{"$f":"'DF raw data'!X19/100","v":3.2912549000000002},{"$f":"'DF raw data'!Y19/100","v":-22.159355000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":367},null,{"$f":"'DF raw data'!G4","v":368},null,{"$f":"'DF raw data'!L4","v":369},null,{"$f":"'DF raw data'!Q4","v":370},null,{"$f":"'DF raw data'!V4","v":371}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-13.580022666666666},{"$f":"AVERAGE(C11:C13)","v":-9.4549738333333337},-4.0055094666666671,1.6345082666666666,-5.4245681999999995,-3.0425457333333328,1.6979322666666663,7.7428600000000003,6.1961813666666652,-15.97395196666667]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16345, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16345, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16345, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16345, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16345, 100 mm gap"]
["Date:","5. apr 2017","08:50:23",null,null,"Date:","5. apr 2017","08:50:46",null,null,"Date:","5. apr 2017","08:50:53",null,null,"Date:","5. apr 2017","08:51:03",null,null,"Date:","5. apr 2017","08:51:09"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",367,null,null,null,"Indice:",368,null,null,null,"Indice:",369,null,null,null,"Indice:",370,null,null,null,"Indice:",371]
["Background Indice:",367,null,null,null,"Background Indice:",368,null,null,null,"Background Indice:",369,null,null,null,"Background Indice:",370,null,null,null,"Background Indice:",371]
["Temperature:",21.18,null,null,null,"Temperature:",21.26,null,null,null,"Temperature:",21.31,null,null,null,"Temperature:",21.3,null,null,null,"Temperature:",21.33]
["Encoder:",395529,null,null,null,"Encoder:",496028,null,null,null,"Encoder:",596356,null,null,null,"Encoder:",796619,null,null,null,"Encoder:",2196657]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,27.934570000000001,1.3626,2358.1738,933.90625,-1,14.221209999999999,2.1308999000000002,-317.84708000000001,699.46680000000003,-1,15.92634,7.4160199000000002,-200.47954999999999,1409.0693000000001,-1,7.8275299,0.55756998000000002,-1381.2665999999999,679.21973000000003,-1,-10.55452,-12.74288,-4776.0815000000002,-1703.2627]
[-0.80000000999999998,-1.6990000000000002E-2,7.2243298999999999,3637.54,2527.8964999999998,-0.80000000999999998,-1.6467499999999999,4.9633798999999996,-397.04138,1019.8407999999999,-0.80000000999999998,4.8802599999999998,-2.2915800000000002,-1388.0154,542.19335999999998,-0.80000000999999998,0.86504000000000003,-16.539949,1276.2562,-6226.7227000000003,-0.80000000999999998,1.9489700000000001,-1.5458799999999999,540.18933000000004,-3071.8027000000002]
[-0.60000001999999997,1.6431100000000001,5.5592699000000003,-632.84149000000002,-65.848633000000007,-0.60000001999999997,4.7610998000000002,1.58945,-1060.7598,-1138.8671999999999,-0.60000001999999997,-4.3952498000000002,2.3433001,-2732.4016000000001,-1326.9229,-0.60000001999999997,10.037559999999999,-6.1542602000000004,1927.5381,-2795.7334000000001,-0.60000001999999997,-13.202819999999999,-2.0933199,-2791.1624000000002,-2468.3173999999999]
[-0.40000001000000002,-5.9981298000000001,1.81488,-4075.3314999999998,636.82128999999998,-0.40000001000000002,5.5353398,0.54697001000000001,-1034.6611,818.90526999999997,-0.40000001000000002,7.9755701999999999,-2.0940398999999998,-310.17730999999998,1570.0986,-0.40000001000000002,7.1664900999999999,-2.8179999000000001E-2,946.26526000000001,423.94238000000001,-0.40000001000000002,17.301860999999999,11.329420000000001,-492.37554999999998,566.1123]
[-0.2,1.3716999999999999,-1.76065,-1399.7678000000001,-1293.8154,-0.2,7.2009401000000004,-2.3036300999999999,129.05797999999999,664.94237999999996,-0.2,7.8340502000000001,-1.01772,-552.82001000000002,1537.4530999999999,-0.2,11.29064,-8.1140699000000005,-776.24054000000001,68.912109000000001,-0.2,11.10721,-6.3233499999999996,3685.5391,111.73242]
[0,-2.89412,-4.8203000999999999,-1741.088,-620.86815999999999,0,2.2576599000000002,1.1411099,-1883.6449,1380.8271,0,-6.7171000999999997,-2.7479998999999999,-1475.1233,-905.38671999999997,0,12.53135,10.754960000000001,1446.3526999999999,3345.4209000000001,0,-4.4783201000000004,-17.812550000000002,-1220.5459000000001,-5828.2061000000003]
[0.2,0.95644998999999997,-6.2779198000000003,-933.15099999999995,-921.80858999999998,0.2,7.1134300000000001,-7.9173298000000001,552.93407999999999,-1555.4169999999999,0.2,10.864739999999999,-9.3977804000000003,400.57285000000002,-1544.8300999999999,0.2,15.74367,-9.6986398999999999,-160.73248000000001,-2264.0234,0.2,15.58559,-2.3820399999999999,-606.13878999999997,924.28809000000001]
[0.40000001000000002,3.6779999999999999,-3.4892299000000002,-656.62720000000002,1011.4746,0.40000001000000002,4.4216699999999998,-10.7037,-422.77715999999998,-671.78319999999997,0.40000001000000002,6.9540100000000002,-6.4906997999999998,1089.2498000000001,318.58886999999999,0.40000001000000002,17.021311000000001,-3.6398999999999999,-1407.1959999999999,1147.7891,0.40000001000000002,-10.49193,-4.6835298999999999,-7139.3584000000001,346.92676]
[0.60000001999999997,11.51876,-3.6308701000000001,1524.2627,201.48047,0.60000001999999997,15.76632,-3.0899000000000001,-85.529792999999998,253.03613000000001,0.60000001999999997,16.39649,-0.29780999000000002,2011.5989,300.25292999999999,0.60000001999999997,18.184290000000001,-8.4832096000000003,864.97266000000002,-1220.3163999999999,0.60000001999999997,6.9077802000000004,-6.8184699999999996,-361.05444,-1284.1094000000001]
[0.80000000999999998,17.527820999999999,-11.95119,3000.4058,762.08496000000002,0.80000000999999998,25.7013,-9.3127297999999996,3237.5767000000001,-22.289061,0.80000000999999998,18.706109999999999,-7.3182302000000004,3394.0810999999999,-25.494140999999999,0.80000000999999998,17.959471000000001,-8.5843095999999992,72.053336999999999,-657.92969000000005,0.80000000999999998,11.612310000000001,-9.8637505000000001,-830.62292000000002,-1080.1415999999999]
[1,28.09647,-7.1880999000000001,3561.6320999999998,-1123.3145,1,29.601410000000001,-5.2975402000000003,2389.165,-2090.0039000000002,1,26.870398999999999,2.88062,5727.5537000000004,1924.0840000000001,1,27.828130999999999,-5.7245498000000001,2695.8696,-2078.9061999999999,1,10.83212,-7.5278100999999999,329.12549000000001,-2215.9355]