{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":123,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":389145},null,{"$f":"'DF raw data'!G7","v":489812},null,{"$f":"'DF raw data'!L7","v":590205},null,{"$f":"'DF raw data'!Q7","v":790599},null,{"$f":"'DF raw data'!V7","v":2190769},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-5.6973567999999988},{"$f":"C21","v":-9.8069699000000004},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":8.6844134666666655},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":5.0759532666666676}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-2.3405000666666664},{"$f":"E21","v":-8.1739868666666666},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":9.4990134999999984},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":12.812293200000001}]
[-10,{"$f":"'DF raw data'!B9","v":38.136490000000002},{"$f":"'DF raw data'!C9","v":-11.553890000000001},{"$f":"'DF raw data'!G9","v":36.049332},{"$f":"'DF raw data'!H9","v":-20.466159999999999},{"$f":"'DF raw data'!L9","v":26.26079},{"$f":"'DF raw data'!M9","v":-12.46298},{"$f":"'DF raw data'!Q9","v":18.947120999999999},{"$f":"'DF raw data'!R9","v":-12.511850000000001},{"$f":"'DF raw data'!V9","v":2.4825001000000002},{"$f":"'DF raw data'!W9","v":-1.1608000000000001},null,{"$f":"F4","v":20},{"$f":"F21","v":1.9679699600000002},{"$f":"G21","v":-4.7310166333333328}]
[-8,{"$f":"'DF raw data'!B10","v":21.251760000000001},{"$f":"'DF raw data'!C10","v":-16.771021000000001},{"$f":"'DF raw data'!G10","v":21.662201},{"$f":"'DF raw data'!H10","v":-7.2132101000000004},{"$f":"'DF raw data'!L10","v":14.10835},{"$f":"'DF raw data'!M10","v":-15.850989999999999},{"$f":"'DF raw data'!Q10","v":9.2624701999999992},{"$f":"'DF raw data'!R10","v":-20.011399999999998},{"$f":"'DF raw data'!V10","v":-3.7810700000000002},{"$f":"'DF raw data'!W10","v":-10.47893},null,{"$f":"H4","v":30},{"$f":"H21","v":2.9870566666666671},{"$f":"I21","v":-6.6522599000000007},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":16.645479000000002},{"$f":"'DF raw data'!C11","v":-13.652900000000001},{"$f":"'DF raw data'!G11","v":15.715870000000001},{"$f":"'DF raw data'!H11","v":-11.119910000000001},{"$f":"'DF raw data'!L11","v":16.902370000000001},{"$f":"'DF raw data'!M11","v":-11.25952},{"$f":"'DF raw data'!Q11","v":5.2541599000000003},{"$f":"'DF raw data'!R11","v":-4.5560597999999999},{"$f":"'DF raw data'!V11","v":2.0509200000000001},{"$f":"'DF raw data'!W11","v":4.7032499000000003},null,{"$f":"J4","v":100},{"$f":"J21","v":3.8016567000000001},{"$f":"K21","v":3.0053233000000001},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":11.145339999999999},{"$f":"'DF raw data'!C12","v":-10.70673},{"$f":"'DF raw data'!G12","v":6.78477},{"$f":"'DF raw data'!H12","v":-11.377370000000001},{"$f":"'DF raw data'!L12","v":12.590809999999999},{"$f":"'DF raw data'!M12","v":-9.0486001999999992},{"$f":"'DF raw data'!Q12","v":13.54077},{"$f":"'DF raw data'!R12","v":-7.6360897999999997},{"$f":"'DF raw data'!V12","v":0.60870999000000003},{"$f":"'DF raw data'!W12","v":-7.1699299999999999},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":8.7998896000000002},{"$f":"'DF raw data'!C13","v":-8.3759499000000002},{"$f":"'DF raw data'!G13","v":1.5125099},{"$f":"'DF raw data'!H13","v":-13.28717},{"$f":"'DF raw data'!L13","v":6.5915599},{"$f":"'DF raw data'!M13","v":-6.7195100999999999},{"$f":"'DF raw data'!Q13","v":14.93952},{"$f":"'DF raw data'!R13","v":-8.1305799000000007},{"$f":"'DF raw data'!V13","v":2.7971200999999999},{"$f":"'DF raw data'!W13","v":4.3512997999999996},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-12.75343},{"$f":"'DF raw data'!C14","v":-9.3341398000000009},{"$f":"'DF raw data'!G14","v":-5.7297400999999999},{"$f":"'DF raw data'!H14","v":-8.0884704999999997},{"$f":"'DF raw data'!L14","v":0.96719997999999996},{"$f":"'DF raw data'!M14","v":-3.8348898999999999},{"$f":"'DF raw data'!Q14","v":4.3037200000000002},{"$f":"'DF raw data'!R14","v":-3.9731399999999999},{"$f":"'DF raw data'!V14","v":11.042350000000001},{"$f":"'DF raw data'!W14","v":3.9193201000000002},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-13.138529999999999},{"$f":"'DF raw data'!C15","v":-11.71082},{"$f":"'DF raw data'!G15","v":-2.8042699999999998},{"$f":"'DF raw data'!H15","v":-3.1463201000000001},{"$f":"'DF raw data'!L15","v":-1.6548499999999999},{"$f":"'DF raw data'!M15","v":-3.6386498999999999},{"$f":"'DF raw data'!Q15","v":-10.282069999999999},{"$f":"'DF raw data'!R15","v":-7.8530597999999996},{"$f":"'DF raw data'!V15","v":-2.4344999999999999},{"$f":"'DF raw data'!W15","v":0.74534999999999996},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-20.748850000000001},{"$f":"'DF raw data'!C16","v":-12.878439999999999},{"$f":"'DF raw data'!G16","v":-22.049021},{"$f":"'DF raw data'!H16","v":-2.0235801000000002},{"$f":"'DF raw data'!L16","v":-9.0631103999999993},{"$f":"'DF raw data'!M16","v":2.02847},{"$f":"'DF raw data'!Q16","v":-2.8101001000000001},{"$f":"'DF raw data'!R16","v":-5.0809898000000002},{"$f":"'DF raw data'!V16","v":-2.7625899},{"$f":"'DF raw data'!W16","v":-1.3705499999999999}]
[6,{"$f":"'DF raw data'!B17","v":-21.76652},{"$f":"'DF raw data'!C17","v":-18.122181000000001},{"$f":"'DF raw data'!G17","v":-11.49606},{"$f":"'DF raw data'!H17","v":2.3765900000000002},{"$f":"'DF raw data'!L17","v":-9.8072003999999993},{"$f":"'DF raw data'!M17","v":0.50248998},{"$f":"'DF raw data'!Q17","v":-4.0703000999999999},{"$f":"'DF raw data'!R17","v":0.15329999999999999},{"$f":"'DF raw data'!V17","v":1.9900599999999999},{"$f":"'DF raw data'!W17","v":-11.75478}]
[8,{"$f":"'DF raw data'!B18","v":-15.23033},{"$f":"'DF raw data'!C18","v":-2.4559000000000002},{"$f":"'DF raw data'!G18","v":-15.34005},{"$f":"'DF raw data'!H18","v":1.1479599},{"$f":"'DF raw data'!L18","v":-5.4159698000000001},{"$f":"'DF raw data'!M18","v":0.81849002999999998},{"$f":"'DF raw data'!Q18","v":12.354480000000001},{"$f":"'DF raw data'!R18","v":-10.598459999999999},{"$f":"'DF raw data'!V18","v":-2.9976101000000002},{"$f":"'DF raw data'!W18","v":-5.4370599000000004}]
[10,{"$f":"'DF raw data'!B19","v":-9.5017204},{"$f":"'DF raw data'!C19","v":10.08527},{"$f":"'DF raw data'!G19","v":-3.5739901000000001},{"$f":"'DF raw data'!H19","v":19.900708999999999},{"$f":"'DF raw data'!L19","v":1.0569900000000001},{"$f":"'DF raw data'!M19","v":10.473039999999999},{"$f":"'DF raw data'!Q19","v":6.3382101000000004},{"$f":"'DF raw data'!R19","v":8.8903704000000001},{"$f":"'DF raw data'!V19","v":5.1620302000000002},{"$f":"'DF raw data'!W19","v":1.02}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":480},null,{"$f":"'DF raw data'!G4","v":481},null,{"$f":"'DF raw data'!L4","v":482},null,{"$f":"'DF raw data'!Q4","v":483},null,{"$f":"'DF raw data'!V4","v":484}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-5.6973567999999988},{"$f":"AVERAGE(C11:C13)","v":-9.8069699000000004},-2.3405000666666664,-8.1739868666666666,1.9679699600000002,-4.7310166333333328,2.9870566666666671,-6.6522599000000007,3.8016567000000001,3.0053233000000001]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":389145},null,{"$f":"'DF raw data'!G7","v":489812},null,{"$f":"'DF raw data'!L7","v":590205},null,{"$f":"'DF raw data'!Q7","v":790599},null,{"$f":"'DF raw data'!V7","v":2190769},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-22.865584266666662},{"$f":"C21","v":2.1451595000000001},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":17.100522499999997},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":8.8911327999999994}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-11.845721966666666},{"$f":"E21","v":2.2380859999999996},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":17.878113863333329},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":13.623997433333333}]
[-10,{"$f":"'DF raw data'!D9/100","v":40.803339999999999},{"$f":"'DF raw data'!E9/100","v":-3.3331348000000003},{"$f":"'DF raw data'!I9/100","v":27.663393999999997},{"$f":"'DF raw data'!J9/100","v":-28.052645999999999},{"$f":"'DF raw data'!N9/100","v":17.350708999999998},{"$f":"'DF raw data'!O9/100","v":-6.6782129000000001},{"$f":"'DF raw data'!S9/100","v":11.831405},{"$f":"'DF raw data'!T9/100","v":-7.3174315999999999},{"$f":"'DF raw data'!X9/100","v":-5.8140917999999999},{"$f":"'DF raw data'!Y9/100","v":9.4176269999999995},null,{"$f":"F4","v":20},{"$f":"F21","v":-8.4333872999999997},{"$f":"G21","v":11.036292299999999}]
[-8,{"$f":"'DF raw data'!D10/100","v":26.111406000000002},{"$f":"'DF raw data'!E10/100","v":-10.699043},{"$f":"'DF raw data'!I10/100","v":20.128523000000001},{"$f":"'DF raw data'!J10/100","v":23.535273000000004},{"$f":"'DF raw data'!N10/100","v":14.637238999999999},{"$f":"'DF raw data'!O10/100","v":7.2703223000000001},{"$f":"'DF raw data'!S10/100","v":7.1514367999999999},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":-15.438213000000001},{"$f":"'DF raw data'!Y10/100","v":-2.5676269999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":-5.7650617666666664},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":27.335334},{"$f":"'DF raw data'!E11/100","v":-0.85535156000000001},{"$f":"'DF raw data'!I11/100","v":23.309699999999999},{"$f":"'DF raw data'!J11/100","v":-15.733789},{"$f":"'DF raw data'!N11/100","v":26.031299000000001},{"$f":"'DF raw data'!O11/100","v":-6.4382714999999999},{"$f":"'DF raw data'!S11/100","v":8.3143042000000005},2.20269,{"$f":"'DF raw data'!X11/100","v":13.204395},{"$f":"'DF raw data'!Y11/100","v":-35.827354},null,{"$f":"J4","v":100},{"$f":"J21","v":-4.987470403333333},{"$f":"K21","v":15.769156933333333}]
[-4,{"$f":"'DF raw data'!D12/100","v":9.2613763000000002},{"$f":"'DF raw data'!E12/100","v":15.310234000000001},{"$f":"'DF raw data'!I12/100","v":8.8721753000000003},{"$f":"'DF raw data'!J12/100","v":8.9760059000000005},{"$f":"'DF raw data'!N12/100","v":1.1020235},{"$f":"'DF raw data'!O12/100","v":24.680635000000002},{"$f":"'DF raw data'!S12/100","v":14.794031},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-3.8315131000000004},{"$f":"'DF raw data'!Y12/100","v":8.2549901999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":-7.7858367999999993},{"$f":"'DF raw data'!E13/100","v":2.9260448999999999},{"$f":"'DF raw data'!I13/100","v":-1.7060339},{"$f":"'DF raw data'!J13/100","v":-17.680576000000002},{"$f":"'DF raw data'!N13/100","v":-1.6687531000000002},{"$f":"'DF raw data'!O13/100","v":13.176367000000001},{"$f":"'DF raw data'!S13/100","v":13.44975},1.63822,{"$f":"'DF raw data'!X13/100","v":-0.14635500000000001},{"$f":"'DF raw data'!Y13/100","v":28.726543000000003},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-32.705114999999999},{"$f":"'DF raw data'!E14/100","v":3.6134570000000004},{"$f":"'DF raw data'!I14/100","v":-15.615785000000001},{"$f":"'DF raw data'!J14/100","v":12.013945000000001},{"$f":"'DF raw data'!N14/100","v":-6.8737877999999997},{"$f":"'DF raw data'!O14/100","v":9.4510059000000002},{"$f":"'DF raw data'!S14/100","v":-3.5136243},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":0.28992978999999997},{"$f":"'DF raw data'!Y14/100","v":22.616543},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-28.105801},{"$f":"'DF raw data'!E15/100","v":-0.1040234},{"$f":"'DF raw data'!I15/100","v":-18.215346999999998},{"$f":"'DF raw data'!J15/100","v":12.380889},{"$f":"'DF raw data'!N15/100","v":-16.757621},{"$f":"'DF raw data'!O15/100","v":10.481503999999999},{"$f":"'DF raw data'!S15/100","v":-27.231311000000002},16.2971,{"$f":"'DF raw data'!X15/100","v":-15.105986},{"$f":"'DF raw data'!Y15/100","v":-4.0356151999999996},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-45.209849000000006},{"$f":"'DF raw data'!E16/100","v":25.675858999999999},{"$f":"'DF raw data'!I16/100","v":-35.538041999999997},{"$f":"'DF raw data'!J16/100","v":30.994004},{"$f":"'DF raw data'!N16/100","v":-19.335916000000001},{"$f":"'DF raw data'!O16/100","v":25.658389000000003},{"$f":"'DF raw data'!S16/100","v":-14.135341},10.6371,{"$f":"'DF raw data'!X16/100","v":-12.195026},{"$f":"'DF raw data'!Y16/100","v":11.840556999999999}]
[6,{"$f":"'DF raw data'!D17/100","v":-43.725893999999997},{"$f":"'DF raw data'!E17/100","v":-3.5506445000000002},{"$f":"'DF raw data'!I17/100","v":-34.530891000000004},{"$f":"'DF raw data'!J17/100","v":20.894766000000001},{"$f":"'DF raw data'!N17/100","v":-1.6126092999999999},{"$f":"'DF raw data'!O17/100","v":33.086885000000002},{"$f":"'DF raw data'!S17/100","v":-9.4221698000000007},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":1.7194336000000001},{"$f":"'DF raw data'!Y17/100","v":2.3199707000000003}]
[8,{"$f":"'DF raw data'!D18/100","v":-10.835099},{"$f":"'DF raw data'!E18/100","v":41.418369000000006},{"$f":"'DF raw data'!I18/100","v":-27.698647000000001},{"$f":"'DF raw data'!J18/100","v":30.981309},{"$f":"'DF raw data'!N18/100","v":-4.7461754999999997},{"$f":"'DF raw data'!O18/100","v":43.178114999999998},{"$f":"'DF raw data'!S18/100","v":27.060919999999999},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":-3.0616003000000003},{"$f":"'DF raw data'!Y18/100","v":5.7229882999999999}]
[10,{"$f":"'DF raw data'!D19/100","v":-26.533166999999999},{"$f":"'DF raw data'!E19/100","v":17.424199000000002},{"$f":"'DF raw data'!I19/100","v":-13.507440999999998},{"$f":"'DF raw data'!J19/100","v":29.377559000000002},{"$f":"'DF raw data'!N19/100","v":-9.1591655999999997},{"$f":"'DF raw data'!O19/100","v":9.1992577999999998},{"$f":"'DF raw data'!S19/100","v":6.7483447000000005},13.3657,{"$f":"'DF raw data'!X19/100","v":8.6148999000000011},{"$f":"'DF raw data'!Y19/100","v":-30.643047000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":480},null,{"$f":"'DF raw data'!G4","v":481},null,{"$f":"'DF raw data'!L4","v":482},null,{"$f":"'DF raw data'!Q4","v":483},null,{"$f":"'DF raw data'!V4","v":484}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-22.865584266666662},{"$f":"AVERAGE(C11:C13)","v":2.1451595000000001},-11.845721966666666,2.2380859999999996,-8.4333872999999997,11.036292299999999,-5.7650617666666664,7.7428600000000003,-4.987470403333333,15.769156933333333]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16349, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16349, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16349, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16349, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16349, 100 mm gap"]
["Date:","10. maj 2017","11:15:57",null,null,"Date:","10. maj 2017","11:30:38",null,null,"Date:","10. maj 2017","11:44:57",null,null,"Date:","10. maj 2017","12:00:16",null,null,"Date:","10. maj 2017","12:47:40"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",480,null,null,null,"Indice:",481,null,null,null,"Indice:",482,null,null,null,"Indice:",483,null,null,null,"Indice:",484]
["Background Indice:",480,null,null,null,"Background Indice:",481,null,null,null,"Background Indice:",482,null,null,null,"Background Indice:",483,null,null,null,"Background Indice:",484]
["Temperature:",21.38,null,null,null,"Temperature:",21.42,null,null,null,"Temperature:",21.41,null,null,null,"Temperature:",21.47,null,null,null,"Temperature:",21.48]
["Encoder:",389145,null,null,null,"Encoder:",489812,null,null,null,"Encoder:",590205,null,null,null,"Encoder:",790599,null,null,null,"Encoder:",2190769]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,38.136490000000002,-11.553890000000001,4080.3339999999998,-333.31348000000003,-1,36.049332,-20.466159999999999,2766.3393999999998,-2805.2646,-1,26.26079,-12.46298,1735.0708999999999,-667.82128999999998,-1,18.947120999999999,-12.511850000000001,1183.1405,-731.74315999999999,-1,2.4825001000000002,-1.1608000000000001,-581.40917999999999,941.7627]
[-0.80000000999999998,21.251760000000001,-16.771021000000001,2611.1406000000002,-1069.9042999999999,-0.80000000999999998,21.662201,-7.2132101000000004,2012.8523,2353.5273000000002,-0.80000000999999998,14.10835,-15.850989999999999,1463.7239,727.03223000000003,-0.80000000999999998,9.2624701999999992,-20.011399999999998,715.14368000000002,-863.48046999999997,-0.80000000999999998,-3.7810700000000002,-10.47893,-1543.8213000000001,-256.7627]
[-0.60000001999999997,16.645479000000002,-13.652900000000001,2733.5333999999998,-85.535156000000001,-0.60000001999999997,15.715870000000001,-11.119910000000001,2330.9699999999998,-1573.3788999999999,-0.60000001999999997,16.902370000000001,-11.25952,2603.1298999999999,-643.82714999999996,-0.60000001999999997,5.2541599000000003,-4.5560597999999999,831.43042000000003,-141.20214999999999,-0.60000001999999997,2.0509200000000001,4.7032499000000003,1320.4395,-3582.7354]
[-0.40000001000000002,11.145339999999999,-10.70673,926.13762999999994,1531.0234,-0.40000001000000002,6.78477,-11.377370000000001,887.21753000000001,897.60059000000001,-0.40000001000000002,12.590809999999999,-9.0486001999999992,110.20235,2468.0635000000002,-0.40000001000000002,13.54077,-7.6360897999999997,1479.4031,1390.3623,-0.40000001000000002,0.60870999000000003,-7.1699299999999999,-383.15131000000002,825.49901999999997]
[-0.2,8.7998896000000002,-8.3759499000000002,-778.58367999999996,292.60449,-0.2,1.5125099,-13.28717,-170.60338999999999,-1768.0576000000001,-0.2,6.5915599,-6.7195100999999999,-166.87531000000001,1317.6367,-0.2,14.93952,-8.1305799000000007,1344.9749999999999,145.87207000000001,-0.2,2.7971200999999999,4.3512997999999996,-14.6355,2872.6543000000001]
[0,-12.75343,-9.3341398000000009,-3270.5115000000001,361.34570000000002,0,-5.7297400999999999,-8.0884704999999997,-1561.5785000000001,1201.3945000000001,0,0.96719997999999996,-3.8348898999999999,-687.37878000000001,945.10059000000001,0,4.3037200000000002,-3.9731399999999999,-351.36243000000002,997.58887000000004,0,11.042350000000001,3.9193201000000002,28.992978999999998,2261.6543000000001]
[0.2,-13.138529999999999,-11.71082,-2810.5801000000001,-10.402340000000001,0.2,-2.8042699999999998,-3.1463201000000001,-1821.5346999999999,1238.0889,0.2,-1.6548499999999999,-3.6386498999999999,-1675.7620999999999,1048.1504,0.2,-10.282069999999999,-7.8530597999999996,-2723.1311000000001,545.28026999999997,0.2,-2.4344999999999999,0.74534999999999996,-1510.5986,-403.56151999999997]
[0.40000001000000002,-20.748850000000001,-12.878439999999999,-4520.9849000000004,2567.5859,0.40000001000000002,-22.049021,-2.0235801000000002,-3553.8042,3099.4004,0.40000001000000002,-9.0631103999999993,2.02847,-1933.5916,2565.8389000000002,0.40000001000000002,-2.8101001000000001,-5.0809898000000002,-1413.5341000000001,888.73535000000004,0.40000001000000002,-2.7625899,-1.3705499999999999,-1219.5026,1184.0556999999999]
[0.60000001999999997,-21.76652,-18.122181000000001,-4372.5893999999998,-355.06445000000002,0.60000001999999997,-11.49606,2.3765900000000002,-3453.0891000000001,2089.4766,0.60000001999999997,-9.8072003999999993,0.50248998,-161.26093,3308.6885000000002,0.60000001999999997,-4.0703000999999999,0.15329999999999999,-942.21698000000004,1790.4238,0.60000001999999997,1.9900599999999999,-11.75478,171.94336000000001,231.99707000000001]
[0.80000000999999998,-15.23033,-2.4559000000000002,-1083.5099,4141.8369000000002,0.80000000999999998,-15.34005,1.1479599,-2769.8647000000001,3098.1309000000001,0.80000000999999998,-5.4159698000000001,0.81849002999999998,-474.61754999999999,4317.8114999999998,0.80000000999999998,12.354480000000001,-10.598459999999999,2706.0920000000001,827.33690999999999,0.80000000999999998,-2.9976101000000002,-5.4370599000000004,-306.16003000000001,572.29882999999995]
[1,-9.5017204,10.08527,-2653.3166999999999,1742.4199000000001,1,-3.5739901000000001,19.900708999999999,-1350.7440999999999,2937.7559000000001,1,1.0569900000000001,10.473039999999999,-915.91656,919.92578000000003,1,6.3382101000000004,8.8903704000000001,674.83447000000001,230.22754,1,5.1620302000000002,1.02,861.48999000000003,-3064.3047000000001]