{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":133,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":382682},null,{"$f":"'DF raw data'!G7","v":482876},null,{"$f":"'DF raw data'!L7","v":583465},null,{"$f":"'DF raw data'!Q7","v":783949},null,{"$f":"'DF raw data'!V7","v":2184014},null,null,{"$f":"B4","v":10},{"$f":"B21","v":2.3579300340000002},{"$f":"C21","v":-8.1316431666666684},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":13.667300366666666},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":8.6966399366666689}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":0.57635996666666645},{"$f":"E21","v":-12.199973300000002},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":13.667300366666666},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":12.617933400000002}]
[-10,{"$f":"'DF raw data'!B9","v":18.451540000000001},{"$f":"'DF raw data'!C9","v":-10.433199999999999},{"$f":"'DF raw data'!G9","v":25.400289999999998},{"$f":"'DF raw data'!H9","v":-12.693479999999999},{"$f":"'DF raw data'!L9","v":15.39044},{"$f":"'DF raw data'!M9","v":-17.006121},{"$f":"'DF raw data'!Q9","v":22.591749},{"$f":"'DF raw data'!R9","v":-14.703950000000001},{"$f":"'DF raw data'!V9","v":6.4842300000000002},{"$f":"'DF raw data'!W9","v":-19.650829000000002},null,{"$f":"F4","v":20},{"$f":"F21","v":11.768129966666669},{"$f":"G21","v":-7.3540233333333331}]
[-8,{"$f":"'DF raw data'!B10","v":16.907689999999999},{"$f":"'DF raw data'!C10","v":-0.48449998999999999},{"$f":"'DF raw data'!G10","v":19.072689},{"$f":"'DF raw data'!H10","v":-9.5814304000000003},{"$f":"'DF raw data'!L10","v":14.76084},{"$f":"'DF raw data'!M10","v":3.9112201},{"$f":"'DF raw data'!Q10","v":26.377680000000002},{"$f":"'DF raw data'!R10","v":-10.676},{"$f":"'DF raw data'!V10","v":12.29613},{"$f":"'DF raw data'!W10","v":-0.66466999000000004},null,{"$f":"H4","v":30},{"$f":"H21","v":14.243660333333333},{"$f":"I21","v":-3.5033333633333332},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":8.9691401000000006},{"$f":"'DF raw data'!C11","v":-11.27115},{"$f":"'DF raw data'!G11","v":10.25536},{"$f":"'DF raw data'!H11","v":-12.340199999999999},{"$f":"'DF raw data'!L11","v":13.69211},{"$f":"'DF raw data'!M11","v":-10.67784},{"$f":"'DF raw data'!Q11","v":13.147970000000001},{"$f":"'DF raw data'!R11","v":-14.30626},{"$f":"'DF raw data'!V11","v":-3.2323401},{"$f":"'DF raw data'!W11","v":0.51235001999999996},null,{"$f":"J4","v":100},{"$f":"J21","v":3.5937165000000006},{"$f":"K21","v":0.41796009999999978},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":7.2742000000000004},{"$f":"'DF raw data'!C12","v":-16.483910000000002},{"$f":"'DF raw data'!G12","v":11.26111},{"$f":"'DF raw data'!H12","v":-15.65504},{"$f":"'DF raw data'!L12","v":4.9953197999999999},{"$f":"'DF raw data'!M12","v":-11.5055},{"$f":"'DF raw data'!Q12","v":12.08461},{"$f":"'DF raw data'!R12","v":-14.01374},{"$f":"'DF raw data'!V12","v":7.1275101000000003},{"$f":"'DF raw data'!W12","v":-6.4521898999999996},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":5.2047701000000002},{"$f":"'DF raw data'!C13","v":-4.4234099000000002},{"$f":"'DF raw data'!G13","v":-1.668E-2},{"$f":"'DF raw data'!H13","v":-9.6147699000000006},{"$f":"'DF raw data'!L13","v":15.358309999999999},{"$f":"'DF raw data'!M13","v":-4.1945601000000003},{"$f":"'DF raw data'!Q13","v":14.622999999999999},{"$f":"'DF raw data'!R13","v":-3.4010400999999999},{"$f":"'DF raw data'!V13","v":6.8576398000000003},{"$f":"'DF raw data'!W13","v":5.1146301999999997},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":7.5630002000000002E-2},{"$f":"'DF raw data'!C14","v":-9.2191296000000005},{"$f":"'DF raw data'!G14","v":-5.5836201000000001},{"$f":"'DF raw data'!H14","v":-15.34859},{"$f":"'DF raw data'!L14","v":6.0932998999999999},{"$f":"'DF raw data'!M14","v":-8.0899400999999997},{"$f":"'DF raw data'!Q14","v":12.07489},{"$f":"'DF raw data'!R14","v":-0.36280999000000003},{"$f":"'DF raw data'!V14","v":-8.4405202999999993},{"$f":"'DF raw data'!W14","v":-4.5669899000000003},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":1.79339},{"$f":"'DF raw data'!C15","v":-10.75239},{"$f":"'DF raw data'!G15","v":7.3293799999999996},{"$f":"'DF raw data'!H15","v":-11.636559999999999},{"$f":"'DF raw data'!L15","v":13.852779999999999},{"$f":"'DF raw data'!M15","v":-9.7775698000000002},{"$f":"'DF raw data'!Q15","v":16.033090999999999},{"$f":"'DF raw data'!R15","v":-6.7461500000000001},{"$f":"'DF raw data'!V15","v":12.36403},{"$f":"'DF raw data'!W15","v":0.70623999999999998},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-5.6843300000000001},{"$f":"'DF raw data'!C16","v":-17.183879999999998},{"$f":"'DF raw data'!G16","v":-7.2435799000000003},{"$f":"'DF raw data'!H16","v":-23.495819000000001},{"$f":"'DF raw data'!L16","v":5.0407500000000001},{"$f":"'DF raw data'!M16","v":-13.49338},{"$f":"'DF raw data'!Q16","v":13.042669999999999},{"$f":"'DF raw data'!R16","v":1.0651799},{"$f":"'DF raw data'!V16","v":-9.1778697999999999},{"$f":"'DF raw data'!W16","v":11.475490000000001}]
[6,{"$f":"'DF raw data'!B17","v":5.3419398999999999},{"$f":"'DF raw data'!C17","v":-12.49995},{"$f":"'DF raw data'!G17","v":6.8567400000000003},{"$f":"'DF raw data'!H17","v":-19.685590999999999},{"$f":"'DF raw data'!L17","v":9.2915200999999996},{"$f":"'DF raw data'!M17","v":-9.2987403999999998},{"$f":"'DF raw data'!Q17","v":16.524090000000001},{"$f":"'DF raw data'!R17","v":-5.1263699999999996},{"$f":"'DF raw data'!V17","v":6.2684698000000001},{"$f":"'DF raw data'!W17","v":-8.2820101000000008}]
[8,{"$f":"'DF raw data'!B18","v":4.8988098999999998},{"$f":"'DF raw data'!C18","v":-24.249179999999999},{"$f":"'DF raw data'!G18","v":12.17324},{"$f":"'DF raw data'!H18","v":-22.15757},{"$f":"'DF raw data'!L18","v":11.30822},{"$f":"'DF raw data'!M18","v":-8.3954801999999997},{"$f":"'DF raw data'!Q18","v":17.220469999999999},{"$f":"'DF raw data'!R18","v":-1.8977900000000001},{"$f":"'DF raw data'!V18","v":8.8408002999999997},{"$f":"'DF raw data'!W18","v":5.8677602000000002}]
[10,{"$f":"'DF raw data'!B19","v":4.8117599000000002},{"$f":"'DF raw data'!C19","v":-22.311969999999999},{"$f":"'DF raw data'!G19","v":3.5978799000000001},{"$f":"'DF raw data'!H19","v":-13.40743},{"$f":"'DF raw data'!L19","v":11.32062},{"$f":"'DF raw data'!M19","v":-15.46575},{"$f":"'DF raw data'!Q19","v":16.407948999999999},{"$f":"'DF raw data'!R19","v":4.4893599000000002},{"$f":"'DF raw data'!V19","v":-0.25525001000000003},{"$f":"'DF raw data'!W19","v":4.7284898999999996}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":501},null,{"$f":"'DF raw data'!G4","v":502},null,{"$f":"'DF raw data'!L4","v":503},null,{"$f":"'DF raw data'!Q4","v":504},null,{"$f":"'DF raw data'!V4","v":505}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":2.3579300340000002},{"$f":"AVERAGE(C11:C13)","v":-8.1316431666666684},0.57635996666666645,-12.199973300000002,11.768129966666669,-7.3540233333333331,14.243660333333333,-3.5033333633333332,3.5937165000000006,0.41796009999999978]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":382682},null,{"$f":"'DF raw data'!G7","v":482876},null,{"$f":"'DF raw data'!L7","v":583465},null,{"$f":"'DF raw data'!Q7","v":783949},null,{"$f":"'DF raw data'!V7","v":2184014},null,null,{"$f":"B4","v":10},{"$f":"B21","v":8.3647226000000003},{"$f":"C21","v":-0.97556966666666634},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":18.364642400000001},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":12.637814196666667}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":11.122390300000001},{"$f":"E21","v":-4.8949541966666663},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":24.087313133333335},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":21.806233430000002}]
[-10,{"$f":"'DF raw data'!D9/100","v":54.687025999999996},{"$f":"'DF raw data'!E9/100","v":-15.11041},{"$f":"'DF raw data'!I9/100","v":53.004185},{"$f":"'DF raw data'!J9/100","v":-22.168008},{"$f":"'DF raw data'!N9/100","v":32.521903999999999},{"$f":"'DF raw data'!O9/100","v":-24.921444999999999},{"$f":"'DF raw data'!S9/100","v":60.831493999999999},{"$f":"'DF raw data'!T9/100","v":-20.452313999999998},{"$f":"'DF raw data'!X9/100","v":29.417926999999999},{"$f":"'DF raw data'!Y9/100","v":-46.538104999999995},null,{"$f":"F4","v":20},{"$f":"F21","v":16.238385400000002},{"$f":"G21","v":-1.1113932666666664}]
[-8,{"$f":"'DF raw data'!D10/100","v":36.526720999999995},{"$f":"'DF raw data'!E10/100","v":-8.5077539000000009},{"$f":"'DF raw data'!I10/100","v":40.203899},{"$f":"'DF raw data'!J10/100","v":19.832822},{"$f":"'DF raw data'!N10/100","v":25.932402000000003},{"$f":"'DF raw data'!O10/100","v":1.3421875000000001},{"$f":"'DF raw data'!S10/100","v":59.892079999999993},3.3834399999999998,{"$f":"'DF raw data'!X10/100","v":18.412213999999999},{"$f":"'DF raw data'!Y10/100","v":3.6630077999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":26.729365000000001},{"$f":"I21","v":7.7428600000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":19.218538000000002},{"$f":"'DF raw data'!E11/100","v":0.78},{"$f":"'DF raw data'!I11/100","v":25.355395999999999},{"$f":"'DF raw data'!J11/100","v":-16.951953},{"$f":"'DF raw data'!N11/100","v":36.510918000000004},{"$f":"'DF raw data'!O11/100","v":-15.659854000000001},{"$f":"'DF raw data'!S11/100","v":24.222932},2.20269,{"$f":"'DF raw data'!X11/100","v":-1.9038425000000001},{"$f":"'DF raw data'!Y11/100","v":15.262528999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":2.6420518666666668},{"$f":"K21","v":16.911279233333335}]
[-4,{"$f":"'DF raw data'!D12/100","v":2.6888873000000002},{"$f":"'DF raw data'!E12/100","v":-29.196347999999997},{"$f":"'DF raw data'!I12/100","v":59.500225},{"$f":"'DF raw data'!J12/100","v":-1.6464453000000001},{"$f":"'DF raw data'!N12/100","v":18.902532000000001},{"$f":"'DF raw data'!O12/100","v":-22.932559000000001},{"$f":"'DF raw data'!S12/100","v":30.81147},9.0253399999999999,{"$f":"'DF raw data'!X12/100","v":-26.046858},{"$f":"'DF raw data'!Y12/100","v":-34.600664000000002}]
[-2,{"$f":"'DF raw data'!D13/100","v":10.781898000000002},{"$f":"'DF raw data'!E13/100","v":7.4865918000000002},{"$f":"'DF raw data'!I13/100","v":9.9988977000000006},{"$f":"'DF raw data'!J13/100","v":10.252139},{"$f":"'DF raw data'!N13/100","v":25.696143000000003},{"$f":"'DF raw data'!O13/100","v":11.623447000000001},{"$f":"'DF raw data'!S13/100","v":28.482932000000002},1.63822,{"$f":"'DF raw data'!X13/100","v":-2.2891614000000002},{"$f":"'DF raw data'!Y13/100","v":65.624199000000004},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":12.521787},{"$f":"'DF raw data'!E14/100","v":-7.8023339999999992},{"$f":"'DF raw data'!I14/100","v":4.5250922000000005},{"$f":"'DF raw data'!J14/100","v":-25.487568},{"$f":"'DF raw data'!N14/100","v":9.8690301999999992},{"$f":"'DF raw data'!O14/100","v":-12.443408},{"$f":"'DF raw data'!S14/100","v":25.915344000000001},5.2932600000000001,{"$f":"'DF raw data'!X14/100","v":7.9938580000000004},{"$f":"'DF raw data'!Y14/100","v":-22.676083999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":1.7904828000000002},{"$f":"'DF raw data'!E15/100","v":-2.6109667999999999},{"$f":"'DF raw data'!I15/100","v":18.843181000000001},{"$f":"'DF raw data'!J15/100","v":0.55056640999999995},{"$f":"'DF raw data'!N15/100","v":13.149982999999999},{"$f":"'DF raw data'!O15/100","v":-2.5142188000000001},{"$f":"'DF raw data'!S15/100","v":25.789819000000001},16.2971,{"$f":"'DF raw data'!X15/100","v":2.2214590000000003},{"$f":"'DF raw data'!Y15/100","v":7.7857227},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-16.665365999999999},{"$f":"'DF raw data'!E16/100","v":2.2283789000000001},{"$f":"'DF raw data'!I16/100","v":3.3450774999999999},{"$f":"'DF raw data'!J16/100","v":-14.721552999999998},{"$f":"'DF raw data'!N16/100","v":-12.120130999999999},{"$f":"'DF raw data'!O16/100","v":-19.492314},{"$f":"'DF raw data'!S16/100","v":20.295770999999998},10.6371,{"$f":"'DF raw data'!X16/100","v":-14.539505999999999},{"$f":"'DF raw data'!Y16/100","v":29.032939000000002}]
[6,{"$f":"'DF raw data'!D17/100","v":5.1599877999999997},{"$f":"'DF raw data'!E17/100","v":10.076064000000001},{"$f":"'DF raw data'!I17/100","v":11.131139000000001},{"$f":"'DF raw data'!J17/100","v":-6.4063378999999996},{"$f":"'DF raw data'!N17/100","v":3.6462E-3},{"$f":"'DF raw data'!O17/100","v":7.9600194999999996},{"$f":"'DF raw data'!S17/100","v":19.600389},10.636900000000001,{"$f":"'DF raw data'!X17/100","v":-1.5540432999999998},{"$f":"'DF raw data'!Y17/100","v":-32.88082}]
[8,{"$f":"'DF raw data'!D18/100","v":8.7281329000000003},{"$f":"'DF raw data'!E18/100","v":-30.578583999999999},{"$f":"'DF raw data'!I18/100","v":21.566782},{"$f":"'DF raw data'!J18/100","v":-41.425497999999997},{"$f":"'DF raw data'!N18/100","v":34.352021000000001},{"$f":"'DF raw data'!O18/100","v":13.686396},{"$f":"'DF raw data'!S18/100","v":35.744453},11.485300000000001,{"$f":"'DF raw data'!X18/100","v":17.935680999999999},{"$f":"'DF raw data'!Y18/100","v":24.273438000000002}]
[10,{"$f":"'DF raw data'!D19/100","v":-16.907177999999998},{"$f":"'DF raw data'!E19/100","v":-23.122939000000002},{"$f":"'DF raw data'!I19/100","v":1.9135373000000002},{"$f":"'DF raw data'!J19/100","v":-26.739756000000003},{"$f":"'DF raw data'!N19/100","v":0.85526787000000004},{"$f":"'DF raw data'!O19/100","v":-37.786104000000002},{"$f":"'DF raw data'!S19/100","v":21.302202000000001},13.3657,{"$f":"'DF raw data'!X19/100","v":13.034462000000001},{"$f":"'DF raw data'!Y19/100","v":17.381035000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":501},null,{"$f":"'DF raw data'!G4","v":502},null,{"$f":"'DF raw data'!L4","v":503},null,{"$f":"'DF raw data'!Q4","v":504},null,{"$f":"'DF raw data'!V4","v":505}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":8.3647226000000003},{"$f":"AVERAGE(C11:C13)","v":-0.97556966666666634},11.122390300000001,-4.8949541966666663,16.238385400000002,-1.1113932666666664,26.729365000000001,7.7428600000000003,2.6420518666666668,16.911279233333335]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16353, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16353, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16353, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16353, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16353, 100 mm gap"]
["Date:","30. maj 2017","10:55:05",null,null,"Date:","30. maj 2017","11:14:57",null,null,"Date:","30. maj 2017","11:37:53",null,null,"Date:","30. maj 2017","11:56:45",null,null,"Date:","30. maj 2017","12:13:07"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",501,null,null,null,"Indice:",502,null,null,null,"Indice:",503,null,null,null,"Indice:",504,null,null,null,"Indice:",505]
["Background Indice:",501,null,null,null,"Background Indice:",502,null,null,null,"Background Indice:",503,null,null,null,"Background Indice:",504,null,null,null,"Background Indice:",505]
["Temperature:",21.52,null,null,null,"Temperature:",21.54,null,null,null,"Temperature:",21.57,null,null,null,"Temperature:",21.61,null,null,null,"Temperature:",21.65]
["Encoder:",382682,null,null,null,"Encoder:",482876,null,null,null,"Encoder:",583465,null,null,null,"Encoder:",783949,null,null,null,"Encoder:",2184014]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,18.451540000000001,-10.433199999999999,5468.7025999999996,-1511.0409999999999,-1,25.400289999999998,-12.693479999999999,5300.4184999999998,-2216.8008,-1,15.39044,-17.006121,3252.1904,-2492.1444999999999,-1,22.591749,-14.703950000000001,6083.1494000000002,-2045.2313999999999,-1,6.4842300000000002,-19.650829000000002,2941.7927,-4653.8104999999996]
[-0.80000000999999998,16.907689999999999,-0.48449998999999999,3652.6720999999998,-850.77539000000002,-0.80000000999999998,19.072689,-9.5814304000000003,4020.3899000000001,1983.2822000000001,-0.80000000999999998,14.76084,3.9112201,2593.2402000000002,134.21875,-0.80000000999999998,26.377680000000002,-10.676,5989.2079999999996,-2375.0742,-0.80000000999999998,12.29613,-0.66466999000000004,1841.2213999999999,366.30077999999997]
[-0.60000001999999997,8.9691401000000006,-11.27115,1921.8538000000001,78,-0.60000001999999997,10.25536,-12.340199999999999,2535.5396000000001,-1695.1953000000001,-0.60000001999999997,13.69211,-10.67784,3651.0918000000001,-1565.9854,-0.60000001999999997,13.147970000000001,-14.30626,2422.2932000000001,-2300.0654,-0.60000001999999997,-3.2323401,0.51235001999999996,-190.38425000000001,1526.2529]
[-0.40000001000000002,7.2742000000000004,-16.483910000000002,268.88873000000001,-2919.6347999999998,-0.40000001000000002,11.26111,-15.65504,5950.0225,-164.64453,-0.40000001000000002,4.9953197999999999,-11.5055,1890.2532000000001,-2293.2559000000001,-0.40000001000000002,12.08461,-14.01374,3081.1469999999999,-2863.9101999999998,-0.40000001000000002,7.1275101000000003,-6.4521898999999996,-2604.6858000000002,-3460.0664000000002]
[-0.2,5.2047701000000002,-4.4234099000000002,1078.1898000000001,748.65917999999999,-0.2,-1.668E-2,-9.6147699000000006,999.88977,1025.2139,-0.2,15.358309999999999,-4.1945601000000003,2569.6143000000002,1162.3447000000001,-0.2,14.622999999999999,-3.4010400999999999,2848.2932000000001,85.250977000000006,-0.2,6.8576398000000003,5.1146301999999997,-228.91614000000001,6562.4198999999999]
[0,7.5630002000000002E-2,-9.2191296000000005,1252.1786999999999,-780.23339999999996,0,-5.5836201000000001,-15.34859,452.50922000000003,-2548.7568000000001,0,6.0932998999999999,-8.0899400999999997,986.90301999999997,-1244.3407999999999,0,12.07489,-0.36280999000000003,2591.5344,1120.5459000000001,0,-8.4405202999999993,-4.5669899000000003,799.38580000000002,-2267.6084000000001]
[0.2,1.79339,-10.75239,179.04828000000001,-261.09667999999999,0.2,7.3293799999999996,-11.636559999999999,1884.3181,55.056640999999999,0.2,13.852779999999999,-9.7775698000000002,1314.9983,-251.42187999999999,0.2,16.033090999999999,-6.7461500000000001,2578.9819000000002,940.97460999999998,0.2,12.36403,0.70623999999999998,222.14590000000001,778.57227]
[0.40000001000000002,-5.6843300000000001,-17.183879999999998,-1666.5365999999999,222.83788999999999,0.40000001000000002,-7.2435799000000003,-23.495819000000001,334.50774999999999,-1472.1552999999999,0.40000001000000002,5.0407500000000001,-13.49338,-1212.0130999999999,-1949.2313999999999,0.40000001000000002,13.042669999999999,1.0651799,2029.5771,861.2373,0.40000001000000002,-9.1778697999999999,11.475490000000001,-1453.9505999999999,2903.2939000000001]
[0.60000001999999997,5.3419398999999999,-12.49995,515.99878000000001,1007.6064,0.60000001999999997,6.8567400000000003,-19.685590999999999,1113.1139000000001,-640.63378999999998,0.60000001999999997,9.2915200999999996,-9.2987403999999998,0.36462,796.00194999999997,0.60000001999999997,16.524090000000001,-5.1263699999999996,1960.0389,221.76855,0.60000001999999997,6.2684698000000001,-8.2820101000000008,-155.40432999999999,-3288.0819999999999]
[0.80000000999999998,4.8988098999999998,-24.249179999999999,872.81329000000005,-3057.8584000000001,0.80000000999999998,12.17324,-22.15757,2156.6781999999998,-4142.5497999999998,0.80000000999999998,11.30822,-8.3954801999999997,3435.2021,1368.6396,0.80000000999999998,17.220469999999999,-1.8977900000000001,3574.4452999999999,886.17187999999999,0.80000000999999998,8.8408002999999997,5.8677602000000002,1793.5681,2427.3438000000001]
[1,4.8117599000000002,-22.311969999999999,-1690.7177999999999,-2312.2939000000001,1,3.5978799000000001,-13.40743,191.35373000000001,-2673.9756000000002,1,11.32062,-15.46575,85.526786999999999,-3778.6104,1,16.407948999999999,4.4893599000000002,2130.2202000000002,1427.9482,1,-0.25525001000000003,4.7284898999999996,1303.4462000000001,1738.1034999999999]