{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":508,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":1214791},null,{"$f":"'DF raw data'!G7","v":1194742},null,{"$f":"'DF raw data'!L7","v":1174727},null,{"$f":"'DF raw data'!Q7","v":1134721},null,{"$f":"'DF raw data'!V7","v":854727},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-1.18874663},{"$f":"C21","v":3.7917833066666673},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":6.2242668366666667},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":5.8111299933333314}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-6.3037334999999999},{"$f":"E21","v":7.0930266000000008},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":6.2242668366666667},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":13.404776026666665}]
[-10,{"$f":"'DF raw data'!B9","v":-19.348908999999999},{"$f":"'DF raw data'!C9","v":-3.9668801},{"$f":"'DF raw data'!G9","v":-19.014869999999998},{"$f":"'DF raw data'!H9","v":-4.8222299},{"$f":"'DF raw data'!L9","v":-17.425540999999999},{"$f":"'DF raw data'!M9","v":-0.88441002000000002},{"$f":"'DF raw data'!Q9","v":-13.3263},{"$f":"'DF raw data'!R9","v":13.47301},{"$f":"'DF raw data'!V9","v":-4.0952902},{"$f":"'DF raw data'!W9","v":19.697109000000001},null,{"$f":"F4","v":20},{"$f":"F21","v":-6.9110733999999994},{"$f":"G21","v":7.1040800333333332}]
[-8,{"$f":"'DF raw data'!B10","v":-13.243790000000001},{"$f":"'DF raw data'!C10","v":-3.8089499},{"$f":"'DF raw data'!G10","v":-12.76558},{"$f":"'DF raw data'!H10","v":-1.4857800000000001},{"$f":"'DF raw data'!L10","v":-8.3128405000000001},{"$f":"'DF raw data'!M10","v":3.02847},{"$f":"'DF raw data'!Q10","v":-10.16863},{"$f":"'DF raw data'!R10","v":7.1960702000000003},{"$f":"'DF raw data'!V10","v":1.7418499999999999},{"$f":"'DF raw data'!W10","v":18.297070000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-7.4130134666666665},{"$f":"I21","v":9.6029132999999991},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-7.2159500000000003},{"$f":"'DF raw data'!C11","v":-5.9002499999999998},{"$f":"'DF raw data'!G11","v":-10.044560000000001},{"$f":"'DF raw data'!H11","v":-4.2056899000000003},{"$f":"'DF raw data'!L11","v":-6.1832500000000001},{"$f":"'DF raw data'!M11","v":2.6387898999999999},{"$f":"'DF raw data'!Q11","v":-15.169},{"$f":"'DF raw data'!R11","v":5.6915202000000003},{"$f":"'DF raw data'!V11","v":-5.8883099999999997},{"$f":"'DF raw data'!W11","v":14.08803},null,{"$f":"J4","v":100},{"$f":"J21","v":-6.2166099666666668},{"$f":"K21","v":17.196559333333333},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-5.0403599999999997},{"$f":"'DF raw data'!C12","v":-0.72977000000000003},{"$f":"'DF raw data'!G12","v":-10.04214},{"$f":"'DF raw data'!H12","v":2.2477800999999999},{"$f":"'DF raw data'!L12","v":-12.186400000000001},{"$f":"'DF raw data'!M12","v":7.8503698999999996},{"$f":"'DF raw data'!Q12","v":-10.048500000000001},{"$f":"'DF raw data'!R12","v":7.9262600000000001},{"$f":"'DF raw data'!V12","v":-10.00858},{"$f":"'DF raw data'!W12","v":10.95504},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-2.4456799},{"$f":"'DF raw data'!C13","v":0.77819002000000004},{"$f":"'DF raw data'!G13","v":-6.9912400000000003},{"$f":"'DF raw data'!H13","v":3.3221400000000001},{"$f":"'DF raw data'!L13","v":-5.3887900999999996},{"$f":"'DF raw data'!M13","v":3.6391100999999999},{"$f":"'DF raw data'!Q13","v":-5.56318},{"$f":"'DF raw data'!R13","v":12.042299999999999},{"$f":"'DF raw data'!V13","v":-4.6845198000000003},{"$f":"'DF raw data'!W13","v":20.830919000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-0.80410999000000005},{"$f":"'DF raw data'!C14","v":3.6380501000000001},{"$f":"'DF raw data'!G14","v":-9.2934704000000004},{"$f":"'DF raw data'!H14","v":11.08878},{"$f":"'DF raw data'!L14","v":-12.925850000000001},{"$f":"'DF raw data'!M14","v":8.0811100000000007},{"$f":"'DF raw data'!Q14","v":-8.6083201999999996},{"$f":"'DF raw data'!R14","v":4.9469098999999996},{"$f":"'DF raw data'!V14","v":-2.8854301000000002},{"$f":"'DF raw data'!W14","v":16.323919},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-0.31645000000000001},{"$f":"'DF raw data'!C15","v":6.9591098000000002},{"$f":"'DF raw data'!G15","v":-2.6264900999999998},{"$f":"'DF raw data'!H15","v":6.8681597999999999},{"$f":"'DF raw data'!L15","v":-2.4185800999999998},{"$f":"'DF raw data'!M15","v":9.5920199999999998},{"$f":"'DF raw data'!Q15","v":-8.0675401999999998},{"$f":"'DF raw data'!R15","v":11.81953},{"$f":"'DF raw data'!V15","v":-11.079879999999999},{"$f":"'DF raw data'!W15","v":14.434839999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":11.52868},{"$f":"'DF raw data'!C16","v":13.051830000000001},{"$f":"'DF raw data'!G16","v":4.2371898000000003},{"$f":"'DF raw data'!H16","v":15.41897},{"$f":"'DF raw data'!L16","v":5.0889502000000002},{"$f":"'DF raw data'!M16","v":15.12476},{"$f":"'DF raw data'!Q16","v":-0.98031997999999998},{"$f":"'DF raw data'!R16","v":12.992229999999999},{"$f":"'DF raw data'!V16","v":-4.2008200000000002},{"$f":"'DF raw data'!W16","v":18.940450999999999}]
[6,{"$f":"'DF raw data'!B17","v":7.8908801000000004},{"$f":"'DF raw data'!C17","v":21.853909999999999},{"$f":"'DF raw data'!G17","v":17.368539999999999},{"$f":"'DF raw data'!H17","v":15.22932},{"$f":"'DF raw data'!L17","v":4.5868802000000004},{"$f":"'DF raw data'!M17","v":9.4611701999999998},{"$f":"'DF raw data'!Q17","v":6.3200497999999996},{"$f":"'DF raw data'!R17","v":8.2961798000000009},{"$f":"'DF raw data'!V17","v":-1.5975299999999999},{"$f":"'DF raw data'!W17","v":13.53393}]
[8,{"$f":"'DF raw data'!B18","v":6.4142298999999996},{"$f":"'DF raw data'!C18","v":19.969549000000001},{"$f":"'DF raw data'!G18","v":9.2529097},{"$f":"'DF raw data'!H18","v":11.539849999999999},{"$f":"'DF raw data'!L18","v":3.99614},{"$f":"'DF raw data'!M18","v":16.207411},{"$f":"'DF raw data'!Q18","v":3.8977799000000002},{"$f":"'DF raw data'!R18","v":15.91663},{"$f":"'DF raw data'!V18","v":-6.8545097999999998},{"$f":"'DF raw data'!W18","v":25.739571000000002}]
[10,{"$f":"'DF raw data'!B19","v":12.815770000000001},{"$f":"'DF raw data'!C19","v":29.062850999999998},{"$f":"'DF raw data'!G19","v":8.9471597999999997},{"$f":"'DF raw data'!H19","v":18.380320000000001},{"$f":"'DF raw data'!L19","v":8.5399103000000007},{"$f":"'DF raw data'!M19","v":13.579639999999999},{"$f":"'DF raw data'!Q19","v":4.8618598000000004},{"$f":"'DF raw data'!R19","v":12.99456},{"$f":"'DF raw data'!V19","v":4.0019999000000004},{"$f":"'DF raw data'!W19","v":15.44131}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":647},null,{"$f":"'DF raw data'!G4","v":648},null,{"$f":"'DF raw data'!L4","v":649},null,{"$f":"'DF raw data'!Q4","v":650},null,{"$f":"'DF raw data'!V4","v":651}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-1.18874663},{"$f":"AVERAGE(C11:C13)","v":3.7917833066666673},-6.3037334999999999,7.0930266000000008,-6.9110733999999994,7.1040800333333332,-7.4130134666666665,9.6029132999999991,-6.2166099666666668,17.196559333333333]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":1214791},null,{"$f":"'DF raw data'!G7","v":1194742},null,{"$f":"'DF raw data'!L7","v":1174727},null,{"$f":"'DF raw data'!Q7","v":1134721},null,{"$f":"'DF raw data'!V7","v":854727},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-1.7370038000000001},{"$f":"C21","v":7.7030861000000002},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":8.7498604000000011},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":24.810611566666665}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-4.6155322666666674},{"$f":"E21","v":16.616458366666667},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":8.7498604000000011},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":25.330536899999998}]
[-10,{"$f":"'DF raw data'!D9/100","v":-20.945111999999998},{"$f":"'DF raw data'!E9/100","v":11.435917999999999},{"$f":"'DF raw data'!I9/100","v":-19.725255000000001},{"$f":"'DF raw data'!J9/100","v":-4.6293749999999996},{"$f":"'DF raw data'!N9/100","v":-25.862666000000001},{"$f":"'DF raw data'!O9/100","v":23.688457},{"$f":"'DF raw data'!S9/100","v":-3.9413684},{"$f":"'DF raw data'!T9/100","v":46.206630999999994},{"$f":"'DF raw data'!X9/100","v":0.56261538999999994},{"$f":"'DF raw data'!Y9/100","v":13.795938000000001},null,{"$f":"F4","v":20},{"$f":"F21","v":-10.486864200000001},{"$f":"G21","v":22.135205000000003}]
[-8,{"$f":"'DF raw data'!D10/100","v":-10.282292},{"$f":"'DF raw data'!E10/100","v":-4.1195117000000003},{"$f":"'DF raw data'!I10/100","v":-17.796880999999999},{"$f":"'DF raw data'!J10/100","v":8.3114062000000004},{"$f":"'DF raw data'!N10/100","v":-8.1404809999999994},{"$f":"'DF raw data'!O10/100","v":12.934756},{"$f":"'DF raw data'!S10/100","v":-9.2086950999999999},{"$f":"'DF raw data'!T10/100","v":21.000839999999997},{"$f":"'DF raw data'!X10/100","v":10.270373999999999},{"$f":"'DF raw data'!Y10/100","v":35.186934000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-10.441551666666667},{"$f":"I21","v":32.513697666666665}]
[-6,{"$f":"'DF raw data'!D11/100","v":-9.5073137999999986},{"$f":"'DF raw data'!E11/100","v":4.8487695000000004},{"$f":"'DF raw data'!I11/100","v":-11.890087000000001},{"$f":"'DF raw data'!J11/100","v":-6.2273828},{"$f":"'DF raw data'!N11/100","v":-3.4729445999999999},{"$f":"'DF raw data'!O11/100","v":14.669141},{"$f":"'DF raw data'!S11/100","v":-16.429852},{"$f":"'DF raw data'!T11/100","v":24.578935999999999},{"$f":"'DF raw data'!X11/100","v":-11.985973},{"$f":"'DF raw data'!Y11/100","v":31.525214999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":-10.145742123333333},{"$f":"K21","v":33.033622999999999}]
[-4,{"$f":"'DF raw data'!D12/100","v":-3.4533557000000004},{"$f":"'DF raw data'!E12/100","v":13.519844000000001},{"$f":"'DF raw data'!I12/100","v":-13.725315999999999},{"$f":"'DF raw data'!J12/100","v":15.561923999999999},{"$f":"'DF raw data'!N12/100","v":-14.127312999999999},{"$f":"'DF raw data'!O12/100","v":32.314999999999998},{"$f":"'DF raw data'!S12/100","v":-8.8925330000000002},{"$f":"'DF raw data'!T12/100","v":25.397597999999999},{"$f":"'DF raw data'!X12/100","v":-12.49452},{"$f":"'DF raw data'!Y12/100","v":36.463018000000005}]
[-2,{"$f":"'DF raw data'!D13/100","v":-4.9122440000000003},{"$f":"'DF raw data'!E13/100","v":8.2673047000000004},{"$f":"'DF raw data'!I13/100","v":-11.243492000000002},{"$f":"'DF raw data'!J13/100","v":29.799551000000001},{"$f":"'DF raw data'!N13/100","v":-12.148614999999999},{"$f":"'DF raw data'!O13/100","v":18.128281000000001},{"$f":"'DF raw data'!S13/100","v":-8.2403943000000002},{"$f":"'DF raw data'!T13/100","v":31.72832},{"$f":"'DF raw data'!X13/100","v":-0.9067141700000001},{"$f":"'DF raw data'!Y13/100","v":40.540947000000003},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":2.6455721999999997},{"$f":"'DF raw data'!E14/100","v":-4.3505663999999999},{"$f":"'DF raw data'!I14/100","v":-2.7193304},{"$f":"'DF raw data'!J14/100","v":17.841953},{"$f":"'DF raw data'!N14/100","v":-13.605752000000001},{"$f":"'DF raw data'!O14/100","v":14.0275},{"$f":"'DF raw data'!S14/100","v":-13.268498999999998},{"$f":"'DF raw data'!T14/100","v":31.899082},{"$f":"'DF raw data'!X14/100","v":-4.0688712000000002},{"$f":"'DF raw data'!Y14/100","v":28.577949},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-2.9443396000000002},{"$f":"'DF raw data'!E15/100","v":19.192519999999998},{"$f":"'DF raw data'!I15/100","v":0.1162256},{"$f":"'DF raw data'!J15/100","v":2.2078711000000002},{"$f":"'DF raw data'!N15/100","v":-5.7062255999999998},{"$f":"'DF raw data'!O15/100","v":34.249834},{"$f":"'DF raw data'!S15/100","v":-9.8157617000000013},{"$f":"'DF raw data'!T15/100","v":33.913691},{"$f":"'DF raw data'!X15/100","v":-25.461641},{"$f":"'DF raw data'!Y15/100","v":29.981972999999996},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":16.538585999999999},{"$f":"'DF raw data'!E16/100","v":26.531464999999997},{"$f":"'DF raw data'!I16/100","v":1.7776471},{"$f":"'DF raw data'!J16/100","v":26.192705},{"$f":"'DF raw data'!N16/100","v":8.0584673999999996},{"$f":"'DF raw data'!O16/100","v":56.256347999999996},{"$f":"'DF raw data'!S16/100","v":-2.5028725999999999},{"$f":"'DF raw data'!T16/100","v":27.507783},{"$f":"'DF raw data'!X16/100","v":-9.6158020000000004},{"$f":"'DF raw data'!Y16/100","v":41.175137000000007}]
[6,{"$f":"'DF raw data'!D17/100","v":4.8132634999999997},{"$f":"'DF raw data'!E17/100","v":36.70834},{"$f":"'DF raw data'!I17/100","v":34.738323000000001},{"$f":"'DF raw data'!J17/100","v":33.206288999999998},{"$f":"'DF raw data'!N17/100","v":5.0932269000000003},{"$f":"'DF raw data'!O17/100","v":23.955243999999997},{"$f":"'DF raw data'!S17/100","v":31.300320000000003},{"$f":"'DF raw data'!T17/100","v":17.735849999999999},{"$f":"'DF raw data'!X17/100","v":-0.85639008000000005},{"$f":"'DF raw data'!Y17/100","v":29.89293}]
[8,{"$f":"'DF raw data'!D18/100","v":3.5098153999999999},{"$f":"'DF raw data'!E18/100","v":46.867070000000005},{"$f":"'DF raw data'!I18/100","v":21.030493},{"$f":"'DF raw data'!J18/100","v":25.153085999999998},{"$f":"'DF raw data'!N18/100","v":7.7936559999999995},{"$f":"'DF raw data'!O18/100","v":40.080137000000001},{"$f":"'DF raw data'!S18/100","v":3.9644553},{"$f":"'DF raw data'!T18/100","v":32.309375000000003},{"$f":"'DF raw data'!X18/100","v":-20.339911999999998},{"$f":"'DF raw data'!Y18/100","v":55.824209000000003}]
[10,{"$f":"'DF raw data'!D19/100","v":19.907014999999998},{"$f":"'DF raw data'!E19/100","v":55.159237999999995},{"$f":"'DF raw data'!I19/100","v":42.012070000000001},{"$f":"'DF raw data'!J19/100","v":39.296923999999997},{"$f":"'DF raw data'!N19/100","v":19.047393},{"$f":"'DF raw data'!O19/100","v":13.161425999999999},{"$f":"'DF raw data'!S19/100","v":4.1037616000000003},{"$f":"'DF raw data'!T19/100","v":30.179561},{"$f":"'DF raw data'!X19/100","v":2.0871021000000001},{"$f":"'DF raw data'!Y19/100","v":42.883983999999998}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":647},null,{"$f":"'DF raw data'!G4","v":648},null,{"$f":"'DF raw data'!L4","v":649},null,{"$f":"'DF raw data'!Q4","v":650},null,{"$f":"'DF raw data'!V4","v":651}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-1.7370038000000001},{"$f":"AVERAGE(C11:C13)","v":7.7030861000000002},-4.6155322666666674,16.616458366666667,-10.486864200000001,22.135205000000003,-10.441551666666667,32.513697666666665,-10.145742123333333,33.033622999999999]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16308, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16308, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16308, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16308, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16308, 100 mm gap"]
["Date:","9. aug 2017","15:08:15",null,null,"Date:","9. aug 2017","15:22:26",null,null,"Date:","9. aug 2017","16:02:58",null,null,"Date:","9. aug 2017","16:03:05",null,null,"Date:","9. aug 2017","16:03:08"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",647,null,null,null,"Indice:",648,null,null,null,"Indice:",649,null,null,null,"Indice:",650,null,null,null,"Indice:",651]
["Background Indice:",647,null,null,null,"Background Indice:",648,null,null,null,"Background Indice:",649,null,null,null,"Background Indice:",650,null,null,null,"Background Indice:",651]
["Temperature:",22.01,null,null,null,"Temperature:",21.96,null,null,null,"Temperature:",21.92,null,null,null,"Temperature:",21.85,null,null,null,"Temperature:",21.86]
["Encoder:",1214791,null,null,null,"Encoder:",1194742,null,null,null,"Encoder:",1174727,null,null,null,"Encoder:",1134721,null,null,null,"Encoder:",854727]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-19.348908999999999,-3.9668801,-2094.5111999999999,1143.5917999999999,-1,-19.014869999999998,-4.8222299,-1972.5255,-462.9375,-1,-17.425540999999999,-0.88441002000000002,-2586.2665999999999,2368.8456999999999,-1,-13.3263,13.47301,-394.13684000000001,4620.6630999999998,-1,-4.0952902,19.697109000000001,56.261538999999999,1379.5938000000001]
[-0.80000000999999998,-13.243790000000001,-3.8089499,-1028.2292,-411.95116999999999,-0.80000000999999998,-12.76558,-1.4857800000000001,-1779.6881000000001,831.14062000000001,-0.80000000999999998,-8.3128405000000001,3.02847,-814.04809999999998,1293.4756,-0.80000000999999998,-10.16863,7.1960702000000003,-920.86950999999999,2100.0839999999998,-0.80000000999999998,1.7418499999999999,18.297070000000001,1027.0373999999999,3518.6934000000001]
[-0.60000001999999997,-7.2159500000000003,-5.9002499999999998,-950.73137999999994,484.87695000000002,-0.60000001999999997,-10.044560000000001,-4.2056899000000003,-1189.0087000000001,-622.73828000000003,-0.60000001999999997,-6.1832500000000001,2.6387898999999999,-347.29446000000002,1466.9141,-0.60000001999999997,-15.169,5.6915202000000003,-1642.9852000000001,2457.8935999999999,-0.60000001999999997,-5.8883099999999997,14.08803,-1198.5972999999999,3152.5214999999998]
[-0.40000001000000002,-5.0403599999999997,-0.72977000000000003,-345.33557000000002,1351.9844000000001,-0.40000001000000002,-10.04214,2.2477800999999999,-1372.5316,1556.1923999999999,-0.40000001000000002,-12.186400000000001,7.8503698999999996,-1412.7312999999999,3231.5,-0.40000001000000002,-10.048500000000001,7.9262600000000001,-889.25329999999997,2539.7597999999998,-0.40000001000000002,-10.00858,10.95504,-1249.452,3646.3018000000002]
[-0.2,-2.4456799,0.77819002000000004,-491.2244,826.73046999999997,-0.2,-6.9912400000000003,3.3221400000000001,-1124.3492000000001,2979.9551000000001,-0.2,-5.3887900999999996,3.6391100999999999,-1214.8615,1812.8280999999999,-0.2,-5.56318,12.042299999999999,-824.03943000000004,3172.8319999999999,-0.2,-4.6845198000000003,20.830919000000002,-90.671417000000005,4054.0947000000001]
[0,-0.80410999000000005,3.6380501000000001,264.55721999999997,-435.05664000000002,0,-9.2934704000000004,11.08878,-271.93304000000001,1784.1953000000001,0,-12.925850000000001,8.0811100000000007,-1360.5752,1402.75,0,-8.6083201999999996,4.9469098999999996,-1326.8498999999999,3189.9081999999999,0,-2.8854301000000002,16.323919,-406.88711999999998,2857.7948999999999]
[0.2,-0.31645000000000001,6.9591098000000002,-294.43396000000001,1919.252,0.2,-2.6264900999999998,6.8681597999999999,11.62256,220.78711000000001,0.2,-2.4185800999999998,9.5920199999999998,-570.62256000000002,3424.9834000000001,0.2,-8.0675401999999998,11.81953,-981.57617000000005,3391.3690999999999,0.2,-11.079879999999999,14.434839999999999,-2546.1641,2998.1972999999998]
[0.40000001000000002,11.52868,13.051830000000001,1653.8586,2653.1464999999998,0.40000001000000002,4.2371898000000003,15.41897,177.76471000000001,2619.2705000000001,0.40000001000000002,5.0889502000000002,15.12476,805.84673999999995,5625.6347999999998,0.40000001000000002,-0.98031997999999998,12.992229999999999,-250.28726,2750.7782999999999,0.40000001000000002,-4.2008200000000002,18.940450999999999,-961.58019999999999,4117.5137000000004]
[0.60000001999999997,7.8908801000000004,21.853909999999999,481.32634999999999,3670.8339999999998,0.60000001999999997,17.368539999999999,15.22932,3473.8323,3320.6289000000002,0.60000001999999997,4.5868802000000004,9.4611701999999998,509.32269000000002,2395.5243999999998,0.60000001999999997,6.3200497999999996,8.2961798000000009,3130.0320000000002,1773.585,0.60000001999999997,-1.5975299999999999,13.53393,-85.639008000000004,2989.2930000000001]
[0.80000000999999998,6.4142298999999996,19.969549000000001,350.98154,4686.7070000000003,0.80000000999999998,9.2529097,11.539849999999999,2103.0493000000001,2515.3085999999998,0.80000000999999998,3.99614,16.207411,779.36559999999997,4008.0137,0.80000000999999998,3.8977799000000002,15.91663,396.44553000000002,3230.9375,0.80000000999999998,-6.8545097999999998,25.739571000000002,-2033.9911999999999,5582.4209000000001]
[1,12.815770000000001,29.062850999999998,1990.7014999999999,5515.9237999999996,1,8.9471597999999997,18.380320000000001,4201.2070000000003,3929.6923999999999,1,8.5399103000000007,13.579639999999999,1904.7393,1316.1425999999999,1,4.8618598000000004,12.99456,410.37616000000003,3017.9560999999999,1,4.0019999000000004,15.44131,208.71020999999999,4288.3984]