{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":518,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17223830},null,{"$f":"'DF raw data'!G7","v":17203783},null,{"$f":"'DF raw data'!L7","v":17183757},null,{"$f":"'DF raw data'!Q7","v":17143759},null,{"$f":"'DF raw data'!V7","v":16863824},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-12.921723333333333},{"$f":"C21","v":12.617086333333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":4.8753470666666647},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":11.5885397}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-8.1202602000000006},{"$f":"E21","v":10.733776633333333},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":5.2683133333333343},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":11.5885397}]
[-10,{"$f":"'DF raw data'!B9","v":-8.9007100999999995},{"$f":"'DF raw data'!C9","v":3.0968100999999999},{"$f":"'DF raw data'!G9","v":-5.0999799000000001},{"$f":"'DF raw data'!H9","v":12.19251},{"$f":"'DF raw data'!L9","v":-3.8068298999999999},{"$f":"'DF raw data'!M9","v":4.0122198999999998},{"$f":"'DF raw data'!Q9","v":-4.4894099000000001},{"$f":"'DF raw data'!R9","v":3.8707299000000002},{"$f":"'DF raw data'!V9","v":-13.81026},{"$f":"'DF raw data'!W9","v":7.3996500999999997},null,{"$f":"F4","v":20},{"$f":"F21","v":-12.771786666666666},{"$f":"G21","v":9.6717166999999993}]
[-8,{"$f":"'DF raw data'!B10","v":-8.4511994999999995},{"$f":"'DF raw data'!C10","v":-1.5645601},{"$f":"'DF raw data'!G10","v":-10.004580000000001},{"$f":"'DF raw data'!H10","v":4.4597502000000002},{"$f":"'DF raw data'!L10","v":-13.654730000000001},{"$f":"'DF raw data'!M10","v":2.6844698999999999},{"$f":"'DF raw data'!Q10","v":-13.77477},{"$f":"'DF raw data'!R10","v":3.9579},{"$f":"'DF raw data'!V10","v":-2.8476300000000001},{"$f":"'DF raw data'!W10","v":2.3055899000000002},null,{"$f":"H4","v":30},{"$f":"H21","v":-8.0463762666666678},{"$f":"I21","v":1.0285466333333335},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-8.6374396999999998},{"$f":"'DF raw data'!C11","v":8.4064502999999995},{"$f":"'DF raw data'!G11","v":-11.151960000000001},{"$f":"'DF raw data'!H11","v":10.965479999999999},{"$f":"'DF raw data'!L11","v":-8.4804200999999999},{"$f":"'DF raw data'!M11","v":17.376349999999999},{"$f":"'DF raw data'!Q11","v":-5.3592199999999997},{"$f":"'DF raw data'!R11","v":7.6689800999999997},{"$f":"'DF raw data'!V11","v":-2.68153},{"$f":"'DF raw data'!W11","v":2.7684801000000001},null,{"$f":"J4","v":100},{"$f":"J21","v":-7.6534099999999983},{"$f":"K21","v":4.5214333333333334},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-14.062569999999999},{"$f":"'DF raw data'!C12","v":15.13763},{"$f":"'DF raw data'!G12","v":-10.23232},{"$f":"'DF raw data'!H12","v":7.9688600999999997},{"$f":"'DF raw data'!L12","v":-15.23143},{"$f":"'DF raw data'!M12","v":6.0974497999999997},{"$f":"'DF raw data'!Q12","v":-10.68206},{"$f":"'DF raw data'!R12","v":8.5982398999999994},{"$f":"'DF raw data'!V12","v":-13.484080000000001},{"$f":"'DF raw data'!W12","v":0.28973000999999998},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-10.93055},{"$f":"'DF raw data'!C13","v":16.295718999999998},{"$f":"'DF raw data'!G13","v":-7.3672199000000003},{"$f":"'DF raw data'!H13","v":12.145770000000001},{"$f":"'DF raw data'!L13","v":-11.61788},{"$f":"'DF raw data'!M13","v":11.3614},{"$f":"'DF raw data'!Q13","v":-6.6286997999999997},{"$f":"'DF raw data'!R13","v":2.9102199},{"$f":"'DF raw data'!V13","v":-11.301729999999999},{"$f":"'DF raw data'!W13","v":10.075670000000001},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-13.9009},{"$f":"'DF raw data'!C14","v":15.41611},{"$f":"'DF raw data'!G14","v":-8.0986404000000007},{"$f":"'DF raw data'!H14","v":8.4029998999999993},{"$f":"'DF raw data'!L14","v":-15.460839999999999},{"$f":"'DF raw data'!M14","v":11.19684},{"$f":"'DF raw data'!Q14","v":2.93573},{"$f":"'DF raw data'!R14","v":-1.6899},{"$f":"'DF raw data'!V14","v":-3.57951},{"$f":"'DF raw data'!W14","v":2.3546200000000002},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-13.933719999999999},{"$f":"'DF raw data'!C15","v":6.1394299999999999},{"$f":"'DF raw data'!G15","v":-8.8949203000000008},{"$f":"'DF raw data'!H15","v":11.652559999999999},{"$f":"'DF raw data'!L15","v":-11.23664},{"$f":"'DF raw data'!M15","v":6.4569101},{"$f":"'DF raw data'!Q15","v":-20.446159000000002},{"$f":"'DF raw data'!R15","v":1.8653200000000001},{"$f":"'DF raw data'!V15","v":-8.0789899999999992},{"$f":"'DF raw data'!W15","v":1.13401},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-16.361111000000001},{"$f":"'DF raw data'!C16","v":21.691151000000001},{"$f":"'DF raw data'!G16","v":-14.22251},{"$f":"'DF raw data'!H16","v":14.89589},{"$f":"'DF raw data'!L16","v":-13.83168},{"$f":"'DF raw data'!M16","v":13.78973},{"$f":"'DF raw data'!Q16","v":-19.136040000000001},{"$f":"'DF raw data'!R16","v":7.1746802000000001},{"$f":"'DF raw data'!V16","v":-10.768969999999999},{"$f":"'DF raw data'!W16","v":9.2693501000000005}]
[6,{"$f":"'DF raw data'!B17","v":-12.05936},{"$f":"'DF raw data'!C17","v":11.04419},{"$f":"'DF raw data'!G17","v":-8.8364601},{"$f":"'DF raw data'!H17","v":13.29058},{"$f":"'DF raw data'!L17","v":-5.3401398999999996},{"$f":"'DF raw data'!M17","v":15.27843},{"$f":"'DF raw data'!Q17","v":-8.3539695999999992},{"$f":"'DF raw data'!R17","v":8.6690102000000007},{"$f":"'DF raw data'!V17","v":-2.1698998999999999},{"$f":"'DF raw data'!W17","v":8.8828297000000003}]
[8,{"$f":"'DF raw data'!B18","v":-21.515920999999999},{"$f":"'DF raw data'!C18","v":11.39953},{"$f":"'DF raw data'!G18","v":-13.0307},{"$f":"'DF raw data'!H18","v":13.09844},{"$f":"'DF raw data'!L18","v":-18.226579999999998},{"$f":"'DF raw data'!M18","v":13.00825},{"$f":"'DF raw data'!Q18","v":-17.439329000000001},{"$f":"'DF raw data'!R18","v":-9.2672997000000006},{"$f":"'DF raw data'!V18","v":-2.6844001},{"$f":"'DF raw data'!W18","v":8.0508699000000004}]
[10,{"$f":"'DF raw data'!B19","v":-19.391559999999998},{"$f":"'DF raw data'!C19","v":13.680120000000001},{"$f":"'DF raw data'!G19","v":-13.87011},{"$f":"'DF raw data'!H19","v":9.2010802999999992},{"$f":"'DF raw data'!L19","v":-17.339569000000001},{"$f":"'DF raw data'!M19","v":7.4057497999999997},{"$f":"'DF raw data'!Q19","v":-2.2480500000000001},{"$f":"'DF raw data'!R19","v":11.425319999999999},{"$f":"'DF raw data'!V19","v":-0.79680996999999998},{"$f":"'DF raw data'!W19","v":7.0063300000000002}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":698},null,{"$f":"'DF raw data'!G4","v":699},null,{"$f":"'DF raw data'!L4","v":700},null,{"$f":"'DF raw data'!Q4","v":701},null,{"$f":"'DF raw data'!V4","v":702}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-12.921723333333333},{"$f":"AVERAGE(C11:C13)","v":12.617086333333333},-8.1202602000000006,10.733776633333333,-12.771786666666666,9.6717166999999993,-8.0463762666666678,1.0285466333333335,-7.6534099999999983,4.5214333333333334]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17223830},null,{"$f":"'DF raw data'!G7","v":17203783},null,{"$f":"'DF raw data'!L7","v":17183757},null,{"$f":"'DF raw data'!Q7","v":17143759},null,{"$f":"'DF raw data'!V7","v":16863824},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-22.957288000000002},{"$f":"C21","v":21.218538366666667},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":13.664738300000002},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":43.077184133333333}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-19.612789933333332},{"$f":"E21","v":14.853363},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":14.384240366666669},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":43.077184133333333}]
[-10,{"$f":"'DF raw data'!D9/100","v":-11.479203999999999},{"$f":"'DF raw data'!E9/100","v":7.3335254000000001},{"$f":"'DF raw data'!I9/100","v":-2.6173428000000003},{"$f":"'DF raw data'!J9/100","v":7.5850097999999999},{"$f":"'DF raw data'!N9/100","v":-4.4881875999999998},{"$f":"'DF raw data'!O9/100","v":7.9554785000000008},{"$f":"'DF raw data'!S9/100","v":-18.952971000000002},{"$f":"'DF raw data'!T9/100","v":18.536992000000001},{"$f":"'DF raw data'!X9/100","v":-41.182558999999998},{"$f":"'DF raw data'!Y9/100","v":9.4755272999999995},null,{"$f":"F4","v":20},{"$f":"F21","v":-20.532708766666666},{"$f":"G21","v":10.404635166666667}]
[-8,{"$f":"'DF raw data'!D10/100","v":-1.5918003999999999},{"$f":"'DF raw data'!E10/100","v":-3.0520508},{"$f":"'DF raw data'!I10/100","v":-7.8375048999999999},{"$f":"'DF raw data'!J10/100","v":0.19099609000000001},{"$f":"'DF raw data'!N10/100","v":-17.499613},{"$f":"'DF raw data'!O10/100","v":1.5349999999999999},{"$f":"'DF raw data'!S10/100","v":-0.61748519999999996},{"$f":"'DF raw data'!T10/100","v":15.448965000000001},{"$f":"'DF raw data'!X10/100","v":-1.0588625},{"$f":"'DF raw data'!Y10/100","v":1.3641015999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":-9.2925497000000004},{"$f":"I21","v":-21.858645766666665}]
[-6,{"$f":"'DF raw data'!D11/100","v":-14.725137999999999},{"$f":"'DF raw data'!E11/100","v":28.286318000000001},{"$f":"'DF raw data'!I11/100","v":-28.885007000000002},{"$f":"'DF raw data'!J11/100","v":17.397197000000002},{"$f":"'DF raw data'!N11/100","v":-3.9839605999999996},{"$f":"'DF raw data'!O11/100","v":31.963895999999998},{"$f":"'DF raw data'!S11/100","v":-11.870145000000001},{"$f":"'DF raw data'!T11/100","v":24.328184},{"$f":"'DF raw data'!X11/100","v":-1.3361341999999998},{"$f":"'DF raw data'!Y11/100","v":8.3683104999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":-8.5730476333333332},{"$f":"K21","v":8.7195476666666671}]
[-4,{"$f":"'DF raw data'!D12/100","v":-13.437953},{"$f":"'DF raw data'!E12/100","v":17.992294999999999},{"$f":"'DF raw data'!I12/100","v":-28.077483000000001},{"$f":"'DF raw data'!J12/100","v":14.175889},{"$f":"'DF raw data'!N12/100","v":-31.384944000000001},{"$f":"'DF raw data'!O12/100","v":4.2919336000000001},{"$f":"'DF raw data'!S12/100","v":-35.453496000000001},{"$f":"'DF raw data'!T12/100","v":-46.643359000000004},{"$f":"'DF raw data'!X12/100","v":-17.042777000000001},{"$f":"'DF raw data'!Y12/100","v":-5.1358690999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":-14.002383},{"$f":"'DF raw data'!E13/100","v":30.321807},{"$f":"'DF raw data'!I13/100","v":-18.799886000000001},{"$f":"'DF raw data'!J13/100","v":14.826143},{"$f":"'DF raw data'!N13/100","v":-22.513607999999998},{"$f":"'DF raw data'!O13/100","v":12.80498},{"$f":"'DF raw data'!S13/100","v":-15.835735},{"$f":"'DF raw data'!T13/100","v":-2.5509472999999998},{"$f":"'DF raw data'!X13/100","v":-16.667750999999999},{"$f":"'DF raw data'!Y13/100","v":18.890352},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-39.560825000000001},{"$f":"'DF raw data'!E14/100","v":26.868895999999999},{"$f":"'DF raw data'!I14/100","v":-34.396161999999997},{"$f":"'DF raw data'!J14/100","v":11.616719},{"$f":"'DF raw data'!N14/100","v":-31.164102},{"$f":"'DF raw data'!O14/100","v":13.453320000000001},{"$f":"'DF raw data'!S14/100","v":-2.1425049},{"$f":"'DF raw data'!T14/100","v":-36.670586},{"$f":"'DF raw data'!X14/100","v":-4.2559575999999995},{"$f":"'DF raw data'!Y14/100","v":3.0558202999999997},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-15.308656000000001},{"$f":"'DF raw data'!E15/100","v":6.4649121000000003},{"$f":"'DF raw data'!I15/100","v":-5.6423217999999995},{"$f":"'DF raw data'!J15/100","v":18.117227},{"$f":"'DF raw data'!N15/100","v":-7.9204163000000003},{"$f":"'DF raw data'!O15/100","v":4.9556054999999999},{"$f":"'DF raw data'!S15/100","v":-9.8994092000000009},{"$f":"'DF raw data'!T15/100","v":-26.354403999999999},{"$f":"'DF raw data'!X15/100","v":-4.7954343000000001},{"$f":"'DF raw data'!Y15/100","v":4.2124706999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-12.566410000000001},{"$f":"'DF raw data'!E16/100","v":35.479609000000004},{"$f":"'DF raw data'!I16/100","v":-13.487263},{"$f":"'DF raw data'!J16/100","v":16.307431999999999},{"$f":"'DF raw data'!N16/100","v":-12.184757999999999},{"$f":"'DF raw data'!O16/100","v":14.632334},{"$f":"'DF raw data'!S16/100","v":-7.5992188000000001},{"$f":"'DF raw data'!T16/100","v":23.759110999999997},{"$f":"'DF raw data'!X16/100","v":-15.8133},{"$f":"'DF raw data'!Y16/100","v":21.066201}]
[6,{"$f":"'DF raw data'!D17/100","v":-18.626068},{"$f":"'DF raw data'!E17/100","v":28.509922000000003},{"$f":"'DF raw data'!I17/100","v":-17.685952999999998},{"$f":"'DF raw data'!J17/100","v":24.894434},{"$f":"'DF raw data'!N17/100","v":-12.421017999999998},{"$f":"'DF raw data'!O17/100","v":43.747109000000002},{"$f":"'DF raw data'!S17/100","v":-37.766554999999997},{"$f":"'DF raw data'!T17/100","v":17.617764000000001},{"$f":"'DF raw data'!X17/100","v":-8.9811461999999995},{"$f":"'DF raw data'!Y17/100","v":15.740215000000001}]
[8,{"$f":"'DF raw data'!D18/100","v":-31.99821},{"$f":"'DF raw data'!E18/100","v":14.950332000000001},{"$f":"'DF raw data'!I18/100","v":-32.180658999999999},{"$f":"'DF raw data'!J18/100","v":19.852725},{"$f":"'DF raw data'!N18/100","v":-24.710763999999998},{"$f":"'DF raw data'!O18/100","v":20.041240000000002},{"$f":"'DF raw data'!S18/100","v":-12.607068},{"$f":"'DF raw data'!T18/100","v":-3.3359668},{"$f":"'DF raw data'!X18/100","v":-7.2390954999999995},{"$f":"'DF raw data'!Y18/100","v":15.714306999999998}]
[10,{"$f":"'DF raw data'!D19/100","v":-30.460578999999999},{"$f":"'DF raw data'!E19/100","v":31.969971000000001},{"$f":"'DF raw data'!I19/100","v":-18.846696999999999},{"$f":"'DF raw data'!J19/100","v":24.515225000000001},{"$f":"'DF raw data'!N19/100","v":-22.810115},{"$f":"'DF raw data'!O19/100","v":24.825918000000001},{"$f":"'DF raw data'!S19/100","v":4.0347955000000004},{"$f":"'DF raw data'!T19/100","v":23.520273000000003},{"$f":"'DF raw data'!X19/100","v":-8.0082269000000004},{"$f":"'DF raw data'!Y19/100","v":24.744150000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":698},null,{"$f":"'DF raw data'!G4","v":699},null,{"$f":"'DF raw data'!L4","v":700},null,{"$f":"'DF raw data'!Q4","v":701},null,{"$f":"'DF raw data'!V4","v":702}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-22.957288000000002},{"$f":"AVERAGE(C11:C13)","v":21.218538366666667},-19.612789933333332,14.853363,-20.532708766666666,10.404635166666667,-9.2925497000000004,-21.858645766666665,-8.5730476333333332,8.7195476666666671]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16311, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16311, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16311, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16311, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16311, 100 mm gap"]
["Date:","23. aug 2017","19:32:26",null,null,"Date:","23. aug 2017","19:45:01",null,null,"Date:","23. aug 2017","19:57:40",null,null,"Date:","23. aug 2017","20:09:53",null,null,"Date:","23. aug 2017","20:23:53"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",698,null,null,null,"Indice:",699,null,null,null,"Indice:",700,null,null,null,"Indice:",701,null,null,null,"Indice:",702]
["Background Indice:",698,null,null,null,"Background Indice:",699,null,null,null,"Background Indice:",700,null,null,null,"Background Indice:",701,null,null,null,"Background Indice:",702]
["Temperature:",21.92,null,null,null,"Temperature:",21.82,null,null,null,"Temperature:",21.91,null,null,null,"Temperature:",21.93,null,null,null,"Temperature:",21.86]
["Encoder:",17223830,null,null,null,"Encoder:",17203783,null,null,null,"Encoder:",17183757,null,null,null,"Encoder:",17143759,null,null,null,"Encoder:",16863824]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-8.9007100999999995,3.0968100999999999,-1147.9204,733.35253999999998,-1,-5.0999799000000001,12.19251,-261.73428000000001,758.50098000000003,-1,-3.8068298999999999,4.0122198999999998,-448.81876,795.54785000000004,-1,-4.4894099000000001,3.8707299000000002,-1895.2971,1853.6992,-1,-13.81026,7.3996500999999997,-4118.2559000000001,947.55273]
[-0.80000000999999998,-8.4511994999999995,-1.5645601,-159.18003999999999,-305.20508000000001,-0.80000000999999998,-10.004580000000001,4.4597502000000002,-783.75049000000001,19.099609000000001,-0.80000000999999998,-13.654730000000001,2.6844698999999999,-1749.9612999999999,153.5,-0.80000000999999998,-13.77477,3.9579,-61.748519999999999,1544.8965000000001,-0.80000000999999998,-2.8476300000000001,2.3055899000000002,-105.88625,136.41015999999999]
[-0.60000001999999997,-8.6374396999999998,8.4064502999999995,-1472.5137999999999,2828.6318000000001,-0.60000001999999997,-11.151960000000001,10.965479999999999,-2888.5007000000001,1739.7197000000001,-0.60000001999999997,-8.4804200999999999,17.376349999999999,-398.39605999999998,3196.3896,-0.60000001999999997,-5.3592199999999997,7.6689800999999997,-1187.0145,2432.8184000000001,-0.60000001999999997,-2.68153,2.7684801000000001,-133.61341999999999,836.83105]
[-0.40000001000000002,-14.062569999999999,15.13763,-1343.7953,1799.2294999999999,-0.40000001000000002,-10.23232,7.9688600999999997,-2807.7483000000002,1417.5889,-0.40000001000000002,-15.23143,6.0974497999999997,-3138.4944,429.19335999999998,-0.40000001000000002,-10.68206,8.5982398999999994,-3545.3496,-4664.3359,-0.40000001000000002,-13.484080000000001,0.28973000999999998,-1704.2777000000001,-513.58690999999999]
[-0.2,-10.93055,16.295718999999998,-1400.2383,3032.1806999999999,-0.2,-7.3672199000000003,12.145770000000001,-1879.9885999999999,1482.6143,-0.2,-11.61788,11.3614,-2251.3607999999999,1280.498,-0.2,-6.6286997999999997,2.9102199,-1583.5735,-255.09473,-0.2,-11.301729999999999,10.075670000000001,-1666.7751000000001,1889.0352]
[0,-13.9009,15.41611,-3956.0825,2686.8896,0,-8.0986404000000007,8.4029998999999993,-3439.6161999999999,1161.6719000000001,0,-15.460839999999999,11.19684,-3116.4101999999998,1345.3320000000001,0,2.93573,-1.6899,-214.25049000000001,-3667.0585999999998,0,-3.57951,2.3546200000000002,-425.59575999999998,305.58202999999997]
[0.2,-13.933719999999999,6.1394299999999999,-1530.8656000000001,646.49121000000002,0.2,-8.8949203000000008,11.652559999999999,-564.23217999999997,1811.7227,0.2,-11.23664,6.4569101,-792.04163000000005,495.56054999999998,0.2,-20.446159000000002,1.8653200000000001,-989.94092000000001,-2635.4404,0.2,-8.0789899999999992,1.13401,-479.54343,421.24707000000001]
[0.40000001000000002,-16.361111000000001,21.691151000000001,-1256.6410000000001,3547.9609,0.40000001000000002,-14.22251,14.89589,-1348.7263,1630.7431999999999,0.40000001000000002,-13.83168,13.78973,-1218.4757999999999,1463.2334000000001,0.40000001000000002,-19.136040000000001,7.1746802000000001,-759.92187999999999,2375.9110999999998,0.40000001000000002,-10.768969999999999,9.2693501000000005,-1581.33,2106.6201000000001]
[0.60000001999999997,-12.05936,11.04419,-1862.6068,2850.9922000000001,0.60000001999999997,-8.8364601,13.29058,-1768.5953,2489.4434000000001,0.60000001999999997,-5.3401398999999996,15.27843,-1242.1017999999999,4374.7109,0.60000001999999997,-8.3539695999999992,8.6690102000000007,-3776.6554999999998,1761.7764,0.60000001999999997,-2.1698998999999999,8.8828297000000003,-898.11461999999995,1574.0215000000001]
[0.80000000999999998,-21.515920999999999,11.39953,-3199.8209999999999,1495.0332000000001,0.80000000999999998,-13.0307,13.09844,-3218.0659000000001,1985.2725,0.80000000999999998,-18.226579999999998,13.00825,-2471.0763999999999,2004.124,0.80000000999999998,-17.439329000000001,-9.2672997000000006,-1260.7067999999999,-333.59667999999999,0.80000000999999998,-2.6844001,8.0508699000000004,-723.90954999999997,1571.4306999999999]
[1,-19.391559999999998,13.680120000000001,-3046.0578999999998,3196.9971,1,-13.87011,9.2010802999999992,-1884.6696999999999,2451.5225,1,-17.339569000000001,7.4057497999999997,-2281.0115000000001,2482.5918000000001,1,-2.2480500000000001,11.425319999999999,403.47955000000002,2352.0273000000002,1,-0.79680996999999998,7.0063300000000002,-800.82268999999997,2474.415]