{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":523,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":8181368},null,{"$f":"'DF raw data'!G7","v":8161313},null,{"$f":"'DF raw data'!L7","v":8141286},null,{"$f":"'DF raw data'!Q7","v":8101270},null,{"$f":"'DF raw data'!V7","v":7821299},null,null,{"$f":"B4","v":10},{"$f":"B21","v":2.3630766666666667},{"$f":"C21","v":6.1241499666666668},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":7.8118300666666673},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":13.337933243333332}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":7.1930533666666667},{"$f":"E21","v":5.6882267000000004},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":7.8118300666666673},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":13.769593199999999}]
[-10,{"$f":"'DF raw data'!B9","v":21.44519},{"$f":"'DF raw data'!C9","v":9.1730803999999999},{"$f":"'DF raw data'!G9","v":20.745229999999999},{"$f":"'DF raw data'!H9","v":8.2173204000000002},{"$f":"'DF raw data'!L9","v":17.894449000000002},{"$f":"'DF raw data'!M9","v":8.2742395000000002},{"$f":"'DF raw data'!Q9","v":12.07056},{"$f":"'DF raw data'!R9","v":11.71294},{"$f":"'DF raw data'!V9","v":6.59023},{"$f":"'DF raw data'!W9","v":-4.2933002},null,{"$f":"F4","v":20},{"$f":"F21","v":10.174906733333334},{"$f":"G21","v":10.714756633333332}]
[-8,{"$f":"'DF raw data'!B10","v":21.01614},{"$f":"'DF raw data'!C10","v":6.4714197999999996},{"$f":"'DF raw data'!G10","v":18.7637},{"$f":"'DF raw data'!H10","v":7.8521399000000001},{"$f":"'DF raw data'!L10","v":16.773239},{"$f":"'DF raw data'!M10","v":-1.10365},{"$f":"'DF raw data'!Q10","v":18.952669},{"$f":"'DF raw data'!R10","v":3.2419997999999998E-2},{"$f":"'DF raw data'!V10","v":15.63105},{"$f":"'DF raw data'!W10","v":-4.2443499999999998},null,{"$f":"H4","v":30},{"$f":"H21","v":8.5918500333333334},{"$f":"I21","v":-2.6231766100000002},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":14.620010000000001},{"$f":"'DF raw data'!C11","v":5.9562100999999998},{"$f":"'DF raw data'!G11","v":17.065398999999999},{"$f":"'DF raw data'!H11","v":-5.2871598999999998},{"$f":"'DF raw data'!L11","v":16.020669999999999},{"$f":"'DF raw data'!M11","v":-5.5470499999999996},{"$f":"'DF raw data'!Q11","v":12.80728},{"$f":"'DF raw data'!R11","v":0.46461998999999998},{"$f":"'DF raw data'!V11","v":3.6080500999999998},{"$f":"'DF raw data'!W11","v":-8.2052297999999997},null,{"$f":"J4","v":100},{"$f":"J21","v":8.5720499666666665},{"$f":"K21","v":-3.0548365666666668},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":9.5793303999999999},{"$f":"'DF raw data'!C12","v":7.5729398999999997},{"$f":"'DF raw data'!G12","v":15.76511},{"$f":"'DF raw data'!H12","v":7.9678601999999996},{"$f":"'DF raw data'!L12","v":8.4185400000000001},{"$f":"'DF raw data'!M12","v":9.8442001000000001},{"$f":"'DF raw data'!Q12","v":17.099291000000001},{"$f":"'DF raw data'!R12","v":7.5938300999999999},{"$f":"'DF raw data'!V12","v":16.062519000000002},{"$f":"'DF raw data'!W12","v":-7.2665901000000002},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":4.3831300999999998},{"$f":"'DF raw data'!C13","v":6.1103101000000004},{"$f":"'DF raw data'!G13","v":3.4940500000000001},{"$f":"'DF raw data'!H13","v":6.6949401000000002},{"$f":"'DF raw data'!L13","v":12.07226},{"$f":"'DF raw data'!M13","v":4.7701200999999998},{"$f":"'DF raw data'!Q13","v":12.99309},{"$f":"'DF raw data'!R13","v":-3.5775499000000002},{"$f":"'DF raw data'!V13","v":9.3110399000000008},{"$f":"'DF raw data'!W13","v":-5.2749300000000003},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":4.5890798999999998},{"$f":"'DF raw data'!C14","v":6.8282398999999998},{"$f":"'DF raw data'!G14","v":11.266920000000001},{"$f":"'DF raw data'!H14","v":3.3550501000000001},{"$f":"'DF raw data'!L14","v":12.467420000000001},{"$f":"'DF raw data'!M14","v":7.6080097999999996},{"$f":"'DF raw data'!Q14","v":22.766960000000001},{"$f":"'DF raw data'!R14","v":-5.0852798999999997},{"$f":"'DF raw data'!V14","v":4.7570199999999998},{"$f":"'DF raw data'!W14","v":4.5744800999999997},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-1.8829800000000001},{"$f":"'DF raw data'!C15","v":5.4338999000000001},{"$f":"'DF raw data'!G15","v":6.8181900999999998},{"$f":"'DF raw data'!H15","v":7.0146898999999996},{"$f":"'DF raw data'!L15","v":5.9850402000000003},{"$f":"'DF raw data'!M15","v":19.76614},{"$f":"'DF raw data'!Q15","v":-9.9844998999999994},{"$f":"'DF raw data'!R15","v":0.79329996999999997},{"$f":"'DF raw data'!V15","v":11.64809},{"$f":"'DF raw data'!W15","v":-8.4640597999999994},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":5.3452101000000001},{"$f":"'DF raw data'!C16","v":12.83353},{"$f":"'DF raw data'!G16","v":8.7496203999999995},{"$f":"'DF raw data'!H16","v":0.76721001},{"$f":"'DF raw data'!L16","v":8.6827001999999993},{"$f":"'DF raw data'!M16","v":4.3780698999999998},{"$f":"'DF raw data'!Q16","v":8.1367197000000004},{"$f":"'DF raw data'!R16","v":-6.0690198000000004},{"$f":"'DF raw data'!V16","v":16.563369999999999},{"$f":"'DF raw data'!W16","v":-4.9127998000000002}]
[6,{"$f":"'DF raw data'!B17","v":4.9858498999999998},{"$f":"'DF raw data'!C17","v":19.175090999999998},{"$f":"'DF raw data'!G17","v":6.2193499000000001},{"$f":"'DF raw data'!H17","v":3.9676399},{"$f":"'DF raw data'!L17","v":6.6025600000000004},{"$f":"'DF raw data'!M17","v":7.5049400000000004},{"$f":"'DF raw data'!Q17","v":10.58273},{"$f":"'DF raw data'!R17","v":2.7832401},{"$f":"'DF raw data'!V17","v":6.2090702000000002},{"$f":"'DF raw data'!W17","v":0.68185001999999995}]
[8,{"$f":"'DF raw data'!B18","v":4.2974199999999998},{"$f":"'DF raw data'!C18","v":7.7019501000000004},{"$f":"'DF raw data'!G18","v":8.4223099000000001},{"$f":"'DF raw data'!H18","v":7.9107399000000003},{"$f":"'DF raw data'!L18","v":24.066130000000001},{"$f":"'DF raw data'!M18","v":3.2396900999999998},{"$f":"'DF raw data'!Q18","v":20.514900000000001},{"$f":"'DF raw data'!R18","v":-1.6544000000000001},{"$f":"'DF raw data'!V18","v":4.8639597999999999},{"$f":"'DF raw data'!W18","v":-12.281370000000001}]
[10,{"$f":"'DF raw data'!B19","v":11.210979999999999},{"$f":"'DF raw data'!C19","v":17.881920000000001},{"$f":"'DF raw data'!G19","v":25.170369999999998},{"$f":"'DF raw data'!H19","v":6.1203498999999999},{"$f":"'DF raw data'!L19","v":18.057639999999999},{"$f":"'DF raw data'!M19","v":8.0666904000000006},{"$f":"'DF raw data'!Q19","v":20.039128999999999},{"$f":"'DF raw data'!R19","v":1.60589},{"$f":"'DF raw data'!V19","v":7.7614197999999996},{"$f":"'DF raw data'!W19","v":-1.0513098999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":714},null,{"$f":"'DF raw data'!G4","v":715},null,{"$f":"'DF raw data'!L4","v":716},null,{"$f":"'DF raw data'!Q4","v":717},null,{"$f":"'DF raw data'!V4","v":718}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":2.3630766666666667},{"$f":"AVERAGE(C11:C13)","v":6.1241499666666668},7.1930533666666667,5.6882267000000004,10.174906733333334,10.714756633333332,8.5918500333333334,-2.6231766100000002,8.5720499666666665,-3.0548365666666668]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":8181368},null,{"$f":"'DF raw data'!G7","v":8161313},null,{"$f":"'DF raw data'!L7","v":8141286},null,{"$f":"'DF raw data'!Q7","v":8101270},null,{"$f":"'DF raw data'!V7","v":7821299},null,null,{"$f":"B4","v":10},{"$f":"B21","v":4.9284544000000006},{"$f":"C21","v":7.7092449333333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":21.953936600000002},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":26.634960733333333}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":26.882391000000002},{"$f":"E21","v":1.8312596666666667},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":21.953936600000002},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":26.634960733333333}]
[-10,{"$f":"'DF raw data'!D9/100","v":46.099975999999998},{"$f":"'DF raw data'!E9/100","v":16.202334},{"$f":"'DF raw data'!I9/100","v":22.316731000000001},{"$f":"'DF raw data'!J9/100","v":6.4213574000000007},{"$f":"'DF raw data'!N9/100","v":32.454045000000001},{"$f":"'DF raw data'!O9/100","v":7.7781444999999998},{"$f":"'DF raw data'!S9/100","v":38.339517000000001},{"$f":"'DF raw data'!T9/100","v":17.876239999999999},{"$f":"'DF raw data'!X9/100","v":8.5333740000000002},{"$f":"'DF raw data'!Y9/100","v":2.9134960999999997},null,{"$f":"F4","v":20},{"$f":"F21","v":6.6670471000000004},{"$f":"G21","v":22.075087666666665}]
[-8,{"$f":"'DF raw data'!D10/100","v":33.789332999999999},{"$f":"'DF raw data'!E10/100","v":15.854766},{"$f":"'DF raw data'!I10/100","v":40.969198999999996},{"$f":"'DF raw data'!J10/100","v":17.308916},{"$f":"'DF raw data'!N10/100","v":35.016033},{"$f":"'DF raw data'!O10/100","v":0.19510739999999999},{"$f":"'DF raw data'!S10/100","v":4.6739166000000001},{"$f":"'DF raw data'!T10/100","v":0.27321288999999999},{"$f":"'DF raw data'!X10/100","v":28.978811},{"$f":"'DF raw data'!Y10/100","v":-11.173798999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":13.347083300000001},{"$f":"I21","v":-4.5598730666666674}]
[-6,{"$f":"'DF raw data'!D11/100","v":22.880112},{"$f":"'DF raw data'!E11/100","v":14.100205000000001},{"$f":"'DF raw data'!I11/100","v":34.317450999999998},{"$f":"'DF raw data'!J11/100","v":-5.7032714999999996},{"$f":"'DF raw data'!N11/100","v":36.312824999999997},{"$f":"'DF raw data'!O11/100","v":7.2868262000000001},{"$f":"'DF raw data'!S11/100","v":25.155543999999999},{"$f":"'DF raw data'!T11/100","v":-6.9068456999999999},{"$f":"'DF raw data'!X11/100","v":2.2045081},{"$f":"'DF raw data'!Y11/100","v":3.3955468999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":10.1100157},{"$f":"K21","v":7.9708397333333325}]
[-4,{"$f":"'DF raw data'!D12/100","v":22.79543},{"$f":"'DF raw data'!E12/100","v":4.2790625000000002},{"$f":"'DF raw data'!I12/100","v":20.873227999999997},{"$f":"'DF raw data'!J12/100","v":11.294609000000001},{"$f":"'DF raw data'!N12/100","v":19.915413000000001},{"$f":"'DF raw data'!O12/100","v":36.033749999999998},{"$f":"'DF raw data'!S12/100","v":40.164895000000001},{"$f":"'DF raw data'!T12/100","v":3.2542968999999999},{"$f":"'DF raw data'!X12/100","v":43.759970999999993},{"$f":"'DF raw data'!Y12/100","v":-19.758251999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":7.9041283999999994},{"$f":"'DF raw data'!E13/100","v":5.8660449000000003},{"$f":"'DF raw data'!I13/100","v":21.607377999999997},{"$f":"'DF raw data'!J13/100","v":-19.784434000000001},{"$f":"'DF raw data'!N13/100","v":6.6472778000000003},{"$f":"'DF raw data'!O13/100","v":13.851561999999999},{"$f":"'DF raw data'!S13/100","v":1.9689246},{"$f":"'DF raw data'!T13/100","v":-7.3086328000000007},{"$f":"'DF raw data'!X13/100","v":16.580992000000002},{"$f":"'DF raw data'!Y13/100","v":-6.6037401999999998},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":8.6002173000000006},{"$f":"'DF raw data'!E14/100","v":7.2072949000000008},{"$f":"'DF raw data'!I14/100","v":45.346454999999999},{"$f":"'DF raw data'!J14/100","v":12.504072000000001},{"$f":"'DF raw data'!N14/100","v":5.7041222999999999},{"$f":"'DF raw data'!O14/100","v":25.136006000000002},{"$f":"'DF raw data'!S14/100","v":45.825869000000004},{"$f":"'DF raw data'!T14/100","v":-10.616045},{"$f":"'DF raw data'!X14/100","v":14.964591},{"$f":"'DF raw data'!Y14/100","v":36.633192999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-1.7189824999999999},{"$f":"'DF raw data'!E15/100","v":10.054395},{"$f":"'DF raw data'!I15/100","v":13.693340000000001},{"$f":"'DF raw data'!J15/100","v":12.774141},{"$f":"'DF raw data'!N15/100","v":7.6497411999999994},{"$f":"'DF raw data'!O15/100","v":27.237694999999999},{"$f":"'DF raw data'!S15/100","v":-7.7535436999999998},{"$f":"'DF raw data'!T15/100","v":4.2450586000000001},{"$f":"'DF raw data'!X15/100","v":-1.2155358999999999},{"$f":"'DF raw data'!Y15/100","v":-6.1169335999999994},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":10.011804},{"$f":"'DF raw data'!E16/100","v":20.231133},{"$f":"'DF raw data'!I16/100","v":2.1048175000000002},{"$f":"'DF raw data'!J16/100","v":17.684141},{"$f":"'DF raw data'!N16/100","v":1.4165954999999999},{"$f":"'DF raw data'!O16/100","v":14.357139},{"$f":"'DF raw data'!S16/100","v":11.462021},{"$f":"'DF raw data'!T16/100","v":-3.1936719},{"$f":"'DF raw data'!X16/100","v":16.800476},{"$f":"'DF raw data'!Y16/100","v":15.523857}]
[6,{"$f":"'DF raw data'!D17/100","v":9.0396117999999994},{"$f":"'DF raw data'!E17/100","v":35.697108999999998},{"$f":"'DF raw data'!I17/100","v":-5.7180102999999995},{"$f":"'DF raw data'!J17/100","v":5.1095996000000001},{"$f":"'DF raw data'!N17/100","v":11.505642},{"$f":"'DF raw data'!O17/100","v":24.714735999999998},{"$f":"'DF raw data'!S17/100","v":6.2389831999999998},{"$f":"'DF raw data'!T17/100","v":8.1167187999999992},{"$f":"'DF raw data'!X17/100","v":0.30928710999999998},{"$f":"'DF raw data'!Y17/100","v":6.0798046999999995}]
[8,{"$f":"'DF raw data'!D18/100","v":8.7935864000000006},{"$f":"'DF raw data'!E18/100","v":8.3416699000000012},{"$f":"'DF raw data'!I18/100","v":41.624053000000004},{"$f":"'DF raw data'!J18/100","v":2.4940625000000001},{"$f":"'DF raw data'!N18/100","v":34.013441999999998},{"$f":"'DF raw data'!O18/100","v":4.4529589999999999},{"$f":"'DF raw data'!S18/100","v":35.600884000000001},{"$f":"'DF raw data'!T18/100","v":8.6445214999999997},{"$f":"'DF raw data'!X18/100","v":4.8153005999999996},{"$f":"'DF raw data'!Y18/100","v":-22.446854999999999}]
[10,{"$f":"'DF raw data'!D19/100","v":18.783349999999999},{"$f":"'DF raw data'!E19/100","v":7.9393848},{"$f":"'DF raw data'!I19/100","v":48.064287000000007},{"$f":"'DF raw data'!J19/100","v":12.680957000000001},{"$f":"'DF raw data'!N19/100","v":12.679413},{"$f":"'DF raw data'!O19/100","v":2.5391406000000001},{"$f":"'DF raw data'!S19/100","v":36.414272000000004},{"$f":"'DF raw data'!T19/100","v":-0.38506839999999998},{"$f":"'DF raw data'!X19/100","v":14.801105999999999},{"$f":"'DF raw data'!Y19/100","v":-3.3632616999999998}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":714},null,{"$f":"'DF raw data'!G4","v":715},null,{"$f":"'DF raw data'!L4","v":716},null,{"$f":"'DF raw data'!Q4","v":717},null,{"$f":"'DF raw data'!V4","v":718}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":4.9284544000000006},{"$f":"AVERAGE(C11:C13)","v":7.7092449333333333},26.882391000000002,1.8312596666666667,6.6670471000000004,22.075087666666665,13.347083300000001,-4.5598730666666674,10.1100157,7.9708397333333325]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16312, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16312, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16312, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16312, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16312, 100 mm gap"]
["Date:","28. aug 2017","07:59:48",null,null,"Date:","28. aug 2017","07:59:52",null,null,"Date:","28. aug 2017","07:59:56",null,null,"Date:","28. aug 2017","08:00:00",null,null,"Date:","28. aug 2017","08:00:03"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",714,null,null,null,"Indice:",715,null,null,null,"Indice:",716,null,null,null,"Indice:",717,null,null,null,"Indice:",718]
["Background Indice:",714,null,null,null,"Background Indice:",715,null,null,null,"Background Indice:",716,null,null,null,"Background Indice:",717,null,null,null,"Background Indice:",718]
["Temperature:",22.04,null,null,null,"Temperature:",21.97,null,null,null,"Temperature:",21.94,null,null,null,"Temperature:",22.01,null,null,null,"Temperature:",21.96]
["Encoder:",8181368,null,null,null,"Encoder:",8161313,null,null,null,"Encoder:",8141286,null,null,null,"Encoder:",8101270,null,null,null,"Encoder:",7821299]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,21.44519,9.1730803999999999,4609.9975999999997,1620.2334000000001,-1,20.745229999999999,8.2173204000000002,2231.6731,642.13574000000006,-1,17.894449000000002,8.2742395000000002,3245.4045000000001,777.81444999999997,-1,12.07056,11.71294,3833.9517000000001,1787.624,-1,6.59023,-4.2933002,853.3374,291.34960999999998]
[-0.80000000999999998,21.01614,6.4714197999999996,3378.9333000000001,1585.4766,-0.80000000999999998,18.7637,7.8521399000000001,4096.9198999999999,1730.8915999999999,-0.80000000999999998,16.773239,-1.10365,3501.6033000000002,19.510739999999998,-0.80000000999999998,18.952669,3.2419997999999998E-2,467.39166,27.321289,-0.80000000999999998,15.63105,-4.2443499999999998,2897.8811000000001,-1117.3798999999999]
[-0.60000001999999997,14.620010000000001,5.9562100999999998,2288.0111999999999,1410.0205000000001,-0.60000001999999997,17.065398999999999,-5.2871598999999998,3431.7451000000001,-570.32714999999996,-0.60000001999999997,16.020669999999999,-5.5470499999999996,3631.2824999999998,728.68262000000004,-0.60000001999999997,12.80728,0.46461998999999998,2515.5544,-690.68457000000001,-0.60000001999999997,3.6080500999999998,-8.2052297999999997,220.45080999999999,339.55468999999999]
[-0.40000001000000002,9.5793303999999999,7.5729398999999997,2279.5430000000001,427.90625,-0.40000001000000002,15.76511,7.9678601999999996,2087.3227999999999,1129.4609,-0.40000001000000002,8.4185400000000001,9.8442001000000001,1991.5413000000001,3603.375,-0.40000001000000002,17.099291000000001,7.5938300999999999,4016.4895000000001,325.42968999999999,-0.40000001000000002,16.062519000000002,-7.2665901000000002,4375.9970999999996,-1975.8252]
[-0.2,4.3831300999999998,6.1103101000000004,790.41283999999996,586.60449000000006,-0.2,3.4940500000000001,6.6949401000000002,2160.7377999999999,-1978.4434000000001,-0.2,12.07226,4.7701200999999998,664.72778000000005,1385.1561999999999,-0.2,12.99309,-3.5775499000000002,196.89246,-730.86328000000003,-0.2,9.3110399000000008,-5.2749300000000003,1658.0992000000001,-660.37401999999997]
[0,4.5890798999999998,6.8282398999999998,860.02173000000005,720.72949000000006,0,11.266920000000001,3.3550501000000001,4534.6454999999996,1250.4072000000001,0,12.467420000000001,7.6080097999999996,570.41223000000002,2513.6006000000002,0,22.766960000000001,-5.0852798999999997,4582.5869000000002,-1061.6044999999999,0,4.7570199999999998,4.5744800999999997,1496.4591,3663.3193000000001]
[0.2,-1.8829800000000001,5.4338999000000001,-171.89824999999999,1005.4395,0.2,6.8181900999999998,7.0146898999999996,1369.3340000000001,1277.4141,0.2,5.9850402000000003,19.76614,764.97411999999997,2723.7694999999999,0.2,-9.9844998999999994,0.79329996999999997,-775.35437000000002,424.50585999999998,0.2,11.64809,-8.4640597999999994,-121.55359,-611.69335999999998]
[0.40000001000000002,5.3452101000000001,12.83353,1001.1804,2023.1133,0.40000001000000002,8.7496203999999995,0.76721001,210.48175000000001,1768.4141,0.40000001000000002,8.6827001999999993,4.3780698999999998,141.65955,1435.7139,0.40000001000000002,8.1367197000000004,-6.0690198000000004,1146.2021,-319.36718999999999,0.40000001000000002,16.563369999999999,-4.9127998000000002,1680.0476000000001,1552.3857]
[0.60000001999999997,4.9858498999999998,19.175090999999998,903.96118000000001,3569.7109,0.60000001999999997,6.2193499000000001,3.9676399,-571.80102999999997,510.95996000000002,0.60000001999999997,6.6025600000000004,7.5049400000000004,1150.5642,2471.4735999999998,0.60000001999999997,10.58273,2.7832401,623.89832000000001,811.67187999999999,0.60000001999999997,6.2090702000000002,0.68185001999999995,30.928711,607.98046999999997]
[0.80000000999999998,4.2974199999999998,7.7019501000000004,879.35864000000004,834.16699000000006,0.80000000999999998,8.4223099000000001,7.9107399000000003,4162.4053000000004,249.40625,0.80000000999999998,24.066130000000001,3.2396900999999998,3401.3442,445.29590000000002,0.80000000999999998,20.514900000000001,-1.6544000000000001,3560.0884000000001,864.45214999999996,0.80000000999999998,4.8639597999999999,-12.281370000000001,481.53005999999999,-2244.6855]
[1,11.210979999999999,17.881920000000001,1878.335,793.93848000000003,1,25.170369999999998,6.1203498999999999,4806.4287000000004,1268.0957000000001,1,18.057639999999999,8.0666904000000006,1267.9413,253.91406000000001,1,20.039128999999999,1.60589,3641.4272000000001,-38.506839999999997,1,7.7614197999999996,-1.0513098999999999,1480.1106,-336.32616999999999]