{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":528,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17543422},null,{"$f":"'DF raw data'!G7","v":17523371},null,{"$f":"'DF raw data'!L7","v":17503337},null,{"$f":"'DF raw data'!Q7","v":17463330},null,{"$f":"'DF raw data'!V7","v":17183357},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-10.28928},{"$f":"C21","v":-9.4822229666666669},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":9.5464567333333328},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":6.680139603333334}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-6.6491967333333335},{"$f":"E21","v":-6.5937799666666663},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":9.5464567333333328},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":10.268136200000001}]
[-10,{"$f":"'DF raw data'!B9","v":-0.33855998999999998},{"$f":"'DF raw data'!C9","v":-2.9459499999999998},{"$f":"'DF raw data'!G9","v":-16.489049999999999},{"$f":"'DF raw data'!H9","v":-4.0307497999999997},{"$f":"'DF raw data'!L9","v":6.3359598999999998},{"$f":"'DF raw data'!M9","v":14.245749999999999},{"$f":"'DF raw data'!Q9","v":-0.35464001000000001},{"$f":"'DF raw data'!R9","v":-6.6577000999999996},{"$f":"'DF raw data'!V9","v":-0.84351999},{"$f":"'DF raw data'!W9","v":8.0856600000000007},null,{"$f":"F4","v":20},{"$f":"F21","v":-0.74282326666666654},{"$f":"G21","v":-5.8905933333333342}]
[-8,{"$f":"'DF raw data'!B10","v":-0.45364000999999998},{"$f":"'DF raw data'!C10","v":-2.1582501000000001},{"$f":"'DF raw data'!G10","v":17.724648999999999},{"$f":"'DF raw data'!H10","v":-9.3504495999999993},{"$f":"'DF raw data'!L10","v":3.3640200999999998},{"$f":"'DF raw data'!M10","v":-6.5110897999999997},{"$f":"'DF raw data'!Q10","v":8.3209496000000005},{"$f":"'DF raw data'!R10","v":-5.1859001999999998},{"$f":"'DF raw data'!V10","v":6.1764001999999998},{"$f":"'DF raw data'!W10","v":-2.8438300999999999},null,{"$f":"H4","v":30},{"$f":"H21","v":-1.4439999966666666},{"$f":"I21","v":-2.8020833633333333},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-5.7554498000000001},{"$f":"'DF raw data'!C11","v":-4.1722102000000003},{"$f":"'DF raw data'!G11","v":-0.85798001000000002},{"$f":"'DF raw data'!H11","v":-4.8391099000000004},{"$f":"'DF raw data'!L11","v":8.1462401999999994},{"$f":"'DF raw data'!M11","v":-12.930960000000001},{"$f":"'DF raw data'!Q11","v":-2.0787201},{"$f":"'DF raw data'!R11","v":-8.6166496000000006},{"$f":"'DF raw data'!V11","v":-10.89518},{"$f":"'DF raw data'!W11","v":-8.3065003999999991},null,{"$f":"J4","v":100},{"$f":"J21","v":-1.2219067333333331},{"$f":"K21","v":0.78591323333333329},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-6.6612201000000004},{"$f":"'DF raw data'!C12","v":-11.57756},{"$f":"'DF raw data'!G12","v":-1.37639},{"$f":"'DF raw data'!H12","v":3.2100401000000001},{"$f":"'DF raw data'!L12","v":8.3080300999999999},{"$f":"'DF raw data'!M12","v":-4.0930600000000004},{"$f":"'DF raw data'!Q12","v":-0.37947999999999998},{"$f":"'DF raw data'!R12","v":1.0715899},{"$f":"'DF raw data'!V12","v":-3.5211000000000001},{"$f":"'DF raw data'!W12","v":-4.9725599000000003},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-10.882809999999999},{"$f":"'DF raw data'!C13","v":-17.363529},{"$f":"'DF raw data'!G13","v":-2.9650599999999998},{"$f":"'DF raw data'!H13","v":-14.073079999999999},{"$f":"'DF raw data'!L13","v":1.27441},{"$f":"'DF raw data'!M13","v":-8.9299297000000006},{"$f":"'DF raw data'!Q13","v":0.32905000000000001},{"$f":"'DF raw data'!R13","v":-15.64531},{"$f":"'DF raw data'!V13","v":3.8618100000000002},{"$f":"'DF raw data'!W13","v":-4.0166501999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-11.9011},{"$f":"'DF raw data'!C14","v":-4.4725999999999999},{"$f":"'DF raw data'!G14","v":-4.9401001999999998},{"$f":"'DF raw data'!H14","v":-6.0726199000000003},{"$f":"'DF raw data'!L14","v":1.6336401},{"$f":"'DF raw data'!M14","v":-8.8998203},{"$f":"'DF raw data'!Q14","v":-0.83013999000000005},{"$f":"'DF raw data'!R14","v":-0.18258999000000001},{"$f":"'DF raw data'!V14","v":-3.0632199999999998},{"$f":"'DF raw data'!W14","v":3.2620298999999999},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-8.0839300000000005},{"$f":"'DF raw data'!C15","v":-6.6105399},{"$f":"'DF raw data'!G15","v":-12.04243},{"$f":"'DF raw data'!H15","v":0.36436000000000002},{"$f":"'DF raw data'!L15","v":-5.1365198999999997},{"$f":"'DF raw data'!M15","v":0.15797},{"$f":"'DF raw data'!Q15","v":-3.8309099999999998},{"$f":"'DF raw data'!R15","v":7.4216499000000002},{"$f":"'DF raw data'!V15","v":-4.4643101999999999},{"$f":"'DF raw data'!W15","v":3.1123599999999998},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-5.0084099999999996},{"$f":"'DF raw data'!C16","v":-7.8836002000000001},{"$f":"'DF raw data'!G16","v":0.68532996999999996},{"$f":"'DF raw data'!H16","v":0.52250998999999998},{"$f":"'DF raw data'!L16","v":10.146129999999999},{"$f":"'DF raw data'!M16","v":-4.7847900000000001},{"$f":"'DF raw data'!Q16","v":-1.1773400000000001},{"$f":"'DF raw data'!R16","v":-3.5825700999999999},{"$f":"'DF raw data'!V16","v":-4.3589701999999999},{"$f":"'DF raw data'!W16","v":1.1890000000000001}]
[6,{"$f":"'DF raw data'!B17","v":-10.45894},{"$f":"'DF raw data'!C17","v":-4.7012701000000003},{"$f":"'DF raw data'!G17","v":-3.9390499999999999},{"$f":"'DF raw data'!H17","v":4.6278701},{"$f":"'DF raw data'!L17","v":-10.360810000000001},{"$f":"'DF raw data'!M17","v":9.8866900999999991},{"$f":"'DF raw data'!Q17","v":-5.3769897999999996},{"$f":"'DF raw data'!R17","v":9.9218902999999994},{"$f":"'DF raw data'!V17","v":-1.94763},{"$f":"'DF raw data'!W17","v":8.2158604000000004}]
[8,{"$f":"'DF raw data'!B18","v":-1.22035},{"$f":"'DF raw data'!C18","v":-7.2800598000000001},{"$f":"'DF raw data'!G18","v":-9.8428202000000002},{"$f":"'DF raw data'!H18","v":-11.35618},{"$f":"'DF raw data'!L18","v":-9.1577196000000001},{"$f":"'DF raw data'!M18","v":-0.54689001999999998},{"$f":"'DF raw data'!Q18","v":-0.4199},{"$f":"'DF raw data'!R18","v":3.1035199000000002},{"$f":"'DF raw data'!V18","v":-0.18049999999999999},{"$f":"'DF raw data'!W18","v":-2.5979999E-2}]
[10,{"$f":"'DF raw data'!B19","v":-6.0244698999999997},{"$f":"'DF raw data'!C19","v":-6.29636},{"$f":"'DF raw data'!G19","v":-7.0255698999999998},{"$f":"'DF raw data'!H19","v":1.76607},{"$f":"'DF raw data'!L19","v":-2.0955501000000001},{"$f":"'DF raw data'!M19","v":2.4345900999999999},{"$f":"'DF raw data'!Q19","v":-0.81003999999999998},{"$f":"'DF raw data'!R19","v":6.0070399999999999},{"$f":"'DF raw data'!V19","v":4.0557299000000002},{"$f":"'DF raw data'!W19","v":5.0915398999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":729},null,{"$f":"'DF raw data'!G4","v":730},null,{"$f":"'DF raw data'!L4","v":734},null,{"$f":"'DF raw data'!Q4","v":732},null,{"$f":"'DF raw data'!V4","v":733}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-10.28928},{"$f":"AVERAGE(C11:C13)","v":-9.4822229666666669},-6.6491967333333335,-6.5937799666666663,-0.74282326666666654,-5.8905933333333342,-1.4439999966666666,-2.8020833633333333,-1.2219067333333331,0.78591323333333329]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17543422},null,{"$f":"'DF raw data'!G7","v":17523371},null,{"$f":"'DF raw data'!L7","v":17503337},null,{"$f":"'DF raw data'!Q7","v":17463330},null,{"$f":"'DF raw data'!V7","v":17183357},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-28.133825666666667},{"$f":"C21","v":-19.748092333333336},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":22.9592779},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":18.620768233333337}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-24.584349166666666},{"$f":"E21","v":-14.119557433333334},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":24.543009333333334},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":21.238645866666669}]
[-10,{"$f":"'DF raw data'!D9/100","v":2.3871869999999999},{"$f":"'DF raw data'!E9/100","v":-9.9108690999999993},{"$f":"'DF raw data'!I9/100","v":-63.509633999999998},{"$f":"'DF raw data'!J9/100","v":2.0181249999999999},{"$f":"'DF raw data'!N9/100","v":-2.1001479999999999},{"$f":"'DF raw data'!O9/100","v":32.677959000000001},{"$f":"'DF raw data'!S9/100","v":16.721359},{"$f":"'DF raw data'!T9/100","v":-4.8165136999999998},{"$f":"'DF raw data'!X9/100","v":11.261140000000001},{"$f":"'DF raw data'!Y9/100","v":8.1398046999999991},null,{"$f":"F4","v":20},{"$f":"F21","v":-5.1745477666666666},{"$f":"G21","v":-12.604156966666666}]
[-8,{"$f":"'DF raw data'!D10/100","v":8.9690502999999993},{"$f":"'DF raw data'!E10/100","v":-14.151465},{"$f":"'DF raw data'!I10/100","v":18.237155999999999},{"$f":"'DF raw data'!J10/100","v":-9.3868358999999995},{"$f":"'DF raw data'!N10/100","v":16.912435000000002},{"$f":"'DF raw data'!O10/100","v":-21.240371},{"$f":"'DF raw data'!S10/100","v":22.782114},{"$f":"'DF raw data'!T10/100","v":-17.636328000000002},{"$f":"'DF raw data'!X10/100","v":-4.3826010000000002},{"$f":"'DF raw data'!Y10/100","v":-34.124141000000002},null,{"$f":"H4","v":30},{"$f":"H21","v":-19.816079333333334},{"$f":"I21","v":-1.1273241000000003}]
[-6,{"$f":"'DF raw data'!D11/100","v":-19.089480000000002},{"$f":"'DF raw data'!E11/100","v":-1.2255077999999999},{"$f":"'DF raw data'!I11/100","v":-2.8509476},{"$f":"'DF raw data'!J11/100","v":9.2706151999999999},{"$f":"'DF raw data'!N11/100","v":-33.907831999999999},{"$f":"'DF raw data'!O11/100","v":-14.582001999999999},{"$f":"'DF raw data'!S11/100","v":2.3691064000000002},{"$f":"'DF raw data'!T11/100","v":-5.4523925999999996},{"$f":"'DF raw data'!X11/100","v":-6.6113324000000002},{"$f":"'DF raw data'!Y11/100","v":-24.533456999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.5908163333333327},{"$f":"K21","v":1.4905535333333335}]
[-4,{"$f":"'DF raw data'!D12/100","v":-1.0021375000000001},{"$f":"'DF raw data'!E12/100","v":-2.8264551},{"$f":"'DF raw data'!I12/100","v":21.434906999999999},{"$f":"'DF raw data'!J12/100","v":3.1971289000000001},{"$f":"'DF raw data'!N12/100","v":20.217085999999998},{"$f":"'DF raw data'!O12/100","v":-6.3625585999999998},{"$f":"'DF raw data'!S12/100","v":9.6028545999999988},{"$f":"'DF raw data'!T12/100","v":-5.0146386999999999},{"$f":"'DF raw data'!X12/100","v":-14.313992000000001},{"$f":"'DF raw data'!Y12/100","v":-8.8662987999999991}]
[-2,{"$f":"'DF raw data'!D13/100","v":-35.579799999999999},{"$f":"'DF raw data'!E13/100","v":-27.139678},{"$f":"'DF raw data'!I13/100","v":-21.084761},{"$f":"'DF raw data'!J13/100","v":-35.647061000000001},{"$f":"'DF raw data'!N13/100","v":5.3186938000000001},{"$f":"'DF raw data'!O13/100","v":-14.499550999999999},{"$f":"'DF raw data'!S13/100","v":-37.930972000000004},{"$f":"'DF raw data'!T13/100","v":-35.341982000000002},{"$f":"'DF raw data'!X13/100","v":36.804785000000003},{"$f":"'DF raw data'!Y13/100","v":-33.211942999999998},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-19.303767000000001},{"$f":"'DF raw data'!E14/100","v":-13.530195000000001},{"$f":"'DF raw data'!I14/100","v":-8.0162694999999999},{"$f":"'DF raw data'!J14/100","v":2.4213379000000002},{"$f":"'DF raw data'!N14/100","v":-3.3373980999999997},{"$f":"'DF raw data'!O14/100","v":-15.574707},{"$f":"'DF raw data'!S14/100","v":-11.245972999999999},{"$f":"'DF raw data'!T14/100","v":4.0083887000000002},{"$f":"'DF raw data'!X14/100","v":-18.942136000000001},{"$f":"'DF raw data'!Y14/100","v":31.179335999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-29.517910000000001},{"$f":"'DF raw data'!E15/100","v":-18.574404000000001},{"$f":"'DF raw data'!I15/100","v":-44.652016999999994},{"$f":"'DF raw data'!J15/100","v":-9.1329492000000005},{"$f":"'DF raw data'!N15/100","v":-17.504939},{"$f":"'DF raw data'!O15/100","v":-7.7382128999999997},{"$f":"'DF raw data'!S15/100","v":-10.271293},{"$f":"'DF raw data'!T15/100","v":27.951620999999999},{"$f":"'DF raw data'!X15/100","v":-28.635097999999999},{"$f":"'DF raw data'!Y15/100","v":6.5042675999999995},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-15.300525},{"$f":"'DF raw data'!E16/100","v":-3.7861425999999998},{"$f":"'DF raw data'!I16/100","v":0.93530486999999995},{"$f":"'DF raw data'!J16/100","v":10.503271},{"$f":"'DF raw data'!N16/100","v":25.838229999999999},{"$f":"'DF raw data'!O16/100","v":-3.4487891000000004},{"$f":"'DF raw data'!S16/100","v":-1.6686242999999998},{"$f":"'DF raw data'!T16/100","v":5.1477734000000002},{"$f":"'DF raw data'!X16/100","v":-4.7470834000000002},{"$f":"'DF raw data'!Y16/100","v":6.4278808999999999}]
[6,{"$f":"'DF raw data'!D17/100","v":-38.667251},{"$f":"'DF raw data'!E17/100","v":-6.7725683999999999},{"$f":"'DF raw data'!I17/100","v":-18.227971},{"$f":"'DF raw data'!J17/100","v":31.142743999999997},{"$f":"'DF raw data'!N17/100","v":-30.727532},{"$f":"'DF raw data'!O17/100","v":22.063739999999999},{"$f":"'DF raw data'!S17/100","v":-14.800540000000002},{"$f":"'DF raw data'!T17/100","v":20.127793},{"$f":"'DF raw data'!X17/100","v":3.5638256999999998},{"$f":"'DF raw data'!Y17/100","v":29.499746000000002}]
[8,{"$f":"'DF raw data'!D18/100","v":-10.143101},{"$f":"'DF raw data'!E18/100","v":-17.349520999999999},{"$f":"'DF raw data'!I18/100","v":-24.795859},{"$f":"'DF raw data'!J18/100","v":-18.940742},{"$f":"'DF raw data'!N18/100","v":-11.318073},{"$f":"'DF raw data'!O18/100","v":0.92291991999999989},{"$f":"'DF raw data'!S18/100","v":-9.2932898000000002},{"$f":"'DF raw data'!T18/100","v":-0.19731449000000001},{"$f":"'DF raw data'!X18/100","v":-8.0070990000000002},{"$f":"'DF raw data'!Y18/100","v":-2.6239941}]
[10,{"$f":"'DF raw data'!D19/100","v":-7.1909229000000003},{"$f":"'DF raw data'!E19/100","v":8.4371875000000003},{"$f":"'DF raw data'!I19/100","v":-28.710183000000001},{"$f":"'DF raw data'!J19/100","v":18.134834000000001},{"$f":"'DF raw data'!N19/100","v":-4.9650167999999999},{"$f":"'DF raw data'!O19/100","v":0.96445312000000005},{"$f":"'DF raw data'!S19/100","v":11.669899000000001},{"$f":"'DF raw data'!T19/100","v":-4.0279688},{"$f":"'DF raw data'!X19/100","v":4.6853961000000002},{"$f":"'DF raw data'!Y19/100","v":-27.468153999999998}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":729},null,{"$f":"'DF raw data'!G4","v":730},null,{"$f":"'DF raw data'!L4","v":734},null,{"$f":"'DF raw data'!Q4","v":732},null,{"$f":"'DF raw data'!V4","v":733}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-28.133825666666667},{"$f":"AVERAGE(C11:C13)","v":-19.748092333333336},-24.584349166666666,-14.119557433333334,-5.1745477666666666,-12.604156966666666,-19.816079333333334,-1.1273241000000003,-3.5908163333333327,1.4905535333333335]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16313, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16313, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16313, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16313, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16313, 100 mm gap"]
["Date:","30. aug 2017","14:17:22",null,null,"Date:","30. aug 2017","13:22:39",null,null,"Date:","31. aug 2017","13:01:53",null,null,"Date:","30. aug 2017","13:49:22",null,null,"Date:","30. aug 2017","14:17:23"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",729,null,null,null,"Indice:",730,null,null,null,"Indice:",734,null,null,null,"Indice:",732,null,null,null,"Indice:",733]
["Background Indice:",729,null,null,null,"Background Indice:",730,null,null,null,"Background Indice:",734,null,null,null,"Background Indice:",732,null,null,null,"Background Indice:",733]
["Temperature:",22.1,null,null,null,"Temperature:",21.88,null,null,null,"Temperature:",21.91,null,null,null,"Temperature:",21.91,null,null,null,"Temperature:",21.91]
["Encoder:",17543422,null,null,null,"Encoder:",17523371,null,null,null,"Encoder:",17503337,null,null,null,"Encoder:",17463330,null,null,null,"Encoder:",17183357]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-0.33855998999999998,-2.9459499999999998,238.71870000000001,-991.08690999999999,-1,-16.489049999999999,-4.0307497999999997,-6350.9633999999996,201.8125,-1,6.3359598999999998,14.245749999999999,-210.01480000000001,3267.7959000000001,-1,-0.35464001000000001,-6.6577000999999996,1672.1359,-481.65136999999999,-1,-0.84351999,8.0856600000000007,1126.114,813.98046999999997]
[-0.80000000999999998,-0.45364000999999998,-2.1582501000000001,896.90503000000001,-1415.1465000000001,-0.80000000999999998,17.724648999999999,-9.3504495999999993,1823.7156,-938.68358999999998,-0.80000000999999998,3.3640200999999998,-6.5110897999999997,1691.2435,-2124.0371,-0.80000000999999998,8.3209496000000005,-5.1859001999999998,2278.2114000000001,-1763.6328000000001,-0.80000000999999998,6.1764001999999998,-2.8438300999999999,-438.26010000000002,-3412.4141]
[-0.60000001999999997,-5.7554498000000001,-4.1722102000000003,-1908.9480000000001,-122.55078,-0.60000001999999997,-0.85798001000000002,-4.8391099000000004,-285.09476000000001,927.06151999999997,-0.60000001999999997,8.1462401999999994,-12.930960000000001,-3390.7831999999999,-1458.2002,-0.60000001999999997,-2.0787201,-8.6166496000000006,236.91064,-545.23925999999994,-0.60000001999999997,-10.89518,-8.3065003999999991,-661.13324,-2453.3456999999999]
[-0.40000001000000002,-6.6612201000000004,-11.57756,-100.21375,-282.64551,-0.40000001000000002,-1.37639,3.2100401000000001,2143.4906999999998,319.71289000000002,-0.40000001000000002,8.3080300999999999,-4.0930600000000004,2021.7085999999999,-636.25585999999998,-0.40000001000000002,-0.37947999999999998,1.0715899,960.28545999999994,-501.46386999999999,-0.40000001000000002,-3.5211000000000001,-4.9725599000000003,-1431.3992000000001,-886.62987999999996]
[-0.2,-10.882809999999999,-17.363529,-3557.98,-2713.9677999999999,-0.2,-2.9650599999999998,-14.073079999999999,-2108.4760999999999,-3564.7060999999999,-0.2,1.27441,-8.9299297000000006,531.86937999999998,-1449.9550999999999,-0.2,0.32905000000000001,-15.64531,-3793.0972000000002,-3534.1981999999998,-0.2,3.8618100000000002,-4.0166501999999999,3680.4785000000002,-3321.1943000000001]
[0,-11.9011,-4.4725999999999999,-1930.3767,-1353.0195000000001,0,-4.9401001999999998,-6.0726199000000003,-801.62694999999997,242.13379,0,1.6336401,-8.8998203,-333.73980999999998,-1557.4707000000001,0,-0.83013999000000005,-0.18258999000000001,-1124.5972999999999,400.83886999999999,0,-3.0632199999999998,3.2620298999999999,-1894.2136,3117.9335999999998]
[0.2,-8.0839300000000005,-6.6105399,-2951.7910000000002,-1857.4404,0.2,-12.04243,0.36436000000000002,-4465.2016999999996,-913.29492000000005,0.2,-5.1365198999999997,0.15797,-1750.4938999999999,-773.82128999999998,0.2,-3.8309099999999998,7.4216499000000002,-1027.1293000000001,2795.1621,0.2,-4.4643101999999999,3.1123599999999998,-2863.5097999999998,650.42675999999994]
[0.40000001000000002,-5.0084099999999996,-7.8836002000000001,-1530.0525,-378.61426,0.40000001000000002,0.68532996999999996,0.52250998999999998,93.530486999999994,1050.3271,0.40000001000000002,10.146129999999999,-4.7847900000000001,2583.8229999999999,-344.87891000000002,0.40000001000000002,-1.1773400000000001,-3.5825700999999999,-166.86242999999999,514.77733999999998,0.40000001000000002,-4.3589701999999999,1.1890000000000001,-474.70834000000002,642.78809000000001]
[0.60000001999999997,-10.45894,-4.7012701000000003,-3866.7251000000001,-677.25684000000001,0.60000001999999997,-3.9390499999999999,4.6278701,-1822.7971,3114.2743999999998,0.60000001999999997,-10.360810000000001,9.8866900999999991,-3072.7532000000001,2206.3739999999998,0.60000001999999997,-5.3769897999999996,9.9218902999999994,-1480.0540000000001,2012.7792999999999,0.60000001999999997,-1.94763,8.2158604000000004,356.38256999999999,2949.9746]
[0.80000000999999998,-1.22035,-7.2800598000000001,-1014.3101,-1734.9521,0.80000000999999998,-9.8428202000000002,-11.35618,-2479.5859,-1894.0742,0.80000000999999998,-9.1577196000000001,-0.54689001999999998,-1131.8072999999999,92.291991999999993,0.80000000999999998,-0.4199,3.1035199000000002,-929.32898,-19.731449000000001,0.80000000999999998,-0.18049999999999999,-2.5979999E-2,-800.70989999999995,-262.39940999999999]
[1,-6.0244698999999997,-6.29636,-719.09229000000005,843.71875,1,-7.0255698999999998,1.76607,-2871.0183000000002,1813.4834000000001,1,-2.0955501000000001,2.4345900999999999,-496.50168000000002,96.445312000000001,1,-0.81003999999999998,6.0070399999999999,1166.9899,-402.79687999999999,1,4.0557299000000002,5.0915398999999999,468.53960999999998,-2746.8154]