{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":554,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17842366},null,{"$f":"'DF raw data'!G7","v":17822322},null,{"$f":"'DF raw data'!L7","v":17802293},null,{"$f":"'DF raw data'!Q7","v":17762274},null,{"$f":"'DF raw data'!V7","v":17482308},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-0.63052996666666672},{"$f":"C21","v":7.7629200666666662},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":2.6716366633333335},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":4.8520032666666655}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":1.0101199633333333},{"$f":"E21","v":12.614923333333332},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":5.9524633333333332},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":7.0592432333333317}]
[-10,{"$f":"'DF raw data'!B9","v":9.3208303000000008},{"$f":"'DF raw data'!C9","v":0.44112000000000001},{"$f":"'DF raw data'!G9","v":3.4060700000000002},{"$f":"'DF raw data'!H9","v":3.6587100000000001},{"$f":"'DF raw data'!L9","v":7.6890501999999996},{"$f":"'DF raw data'!M9","v":10.32161},{"$f":"'DF raw data'!Q9","v":-1.2083499},{"$f":"'DF raw data'!R9","v":6.1024899000000001},{"$f":"'DF raw data'!V9","v":9.3722800999999993},{"$f":"'DF raw data'!W9","v":-0.48019999000000002},null,{"$f":"F4","v":20},{"$f":"F21","v":0.56548989999999988},{"$f":"G21","v":9.2176065666666673}]
[-8,{"$f":"'DF raw data'!B10","v":6.1370300999999996},{"$f":"'DF raw data'!C10","v":6.8191299000000001},{"$f":"'DF raw data'!G10","v":9.7957000999999995},{"$f":"'DF raw data'!H10","v":16.690650999999999},{"$f":"'DF raw data'!L10","v":8.8227595999999995},{"$f":"'DF raw data'!M10","v":9.3650999000000006},{"$f":"'DF raw data'!Q10","v":-0.80615002000000002},{"$f":"'DF raw data'!R10","v":8.7248696999999993},{"$f":"'DF raw data'!V10","v":6.4730002999999994E-2},{"$f":"'DF raw data'!W10","v":8.7134894999999997},null,{"$f":"H4","v":30},{"$f":"H21","v":2.0411066966666667},{"$f":"I21","v":9.9315965999999989},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":5.0733699999999997},{"$f":"'DF raw data'!C11","v":-1.65412},{"$f":"'DF raw data'!G11","v":6.7346301000000004},{"$f":"'DF raw data'!H11","v":8.7563496000000001},{"$f":"'DF raw data'!L11","v":3.26966},{"$f":"'DF raw data'!M11","v":6.5250401},{"$f":"'DF raw data'!Q11","v":0.57238001000000005},{"$f":"'DF raw data'!R11","v":2.6145201},{"$f":"'DF raw data'!V11","v":5.9566498000000001},{"$f":"'DF raw data'!W11","v":6.7311902000000003},null,{"$f":"J4","v":100},{"$f":"J21","v":5.3219333666666664},{"$f":"K21","v":5.5556801},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":0.95644998999999997},{"$f":"'DF raw data'!C12","v":1.9544801000000001},{"$f":"'DF raw data'!G12","v":7.3263001000000001},{"$f":"'DF raw data'!H12","v":5.1712898999999997},{"$f":"'DF raw data'!L12","v":7.2519698000000004},{"$f":"'DF raw data'!M12","v":0.99844003000000003},{"$f":"'DF raw data'!Q12","v":0.19244},{"$f":"'DF raw data'!R12","v":16.590851000000001},{"$f":"'DF raw data'!V12","v":8.6490402},{"$f":"'DF raw data'!W12","v":1.67188},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-1.7002999999999999},{"$f":"'DF raw data'!C13","v":5.9343399999999997},{"$f":"'DF raw data'!G13","v":0.38418998999999998},{"$f":"'DF raw data'!H13","v":15.15962},{"$f":"'DF raw data'!L13","v":0.67911999999999995},{"$f":"'DF raw data'!M13","v":16.28013},{"$f":"'DF raw data'!Q13","v":-0.93276000000000003},{"$f":"'DF raw data'!R13","v":7.9956101999999998},{"$f":"'DF raw data'!V13","v":8.2332801999999994},{"$f":"'DF raw data'!W13","v":5.2156801000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":5.6833701000000003},{"$f":"'DF raw data'!C14","v":10.44064},{"$f":"'DF raw data'!G14","v":4.0641898999999997},{"$f":"'DF raw data'!H14","v":10.129989999999999},{"$f":"'DF raw data'!L14","v":8.0317898000000003},{"$f":"'DF raw data'!M14","v":3.9219799000000002},{"$f":"'DF raw data'!Q14","v":6.7142901000000004},{"$f":"'DF raw data'!R14","v":9.4954795999999995},{"$f":"'DF raw data'!V14","v":5.4262300000000003},{"$f":"'DF raw data'!W14","v":6.0588002000000003},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-5.8746600000000004},{"$f":"'DF raw data'!C15","v":6.9137801999999997},{"$f":"'DF raw data'!G15","v":-1.4180200000000001},{"$f":"'DF raw data'!H15","v":12.555160000000001},{"$f":"'DF raw data'!L15","v":-7.0144400999999998},{"$f":"'DF raw data'!M15","v":7.4507098000000003},{"$f":"'DF raw data'!Q15","v":0.34178998999999999},{"$f":"'DF raw data'!R15","v":12.303699999999999},{"$f":"'DF raw data'!V15","v":2.3062898999999999},{"$f":"'DF raw data'!W15","v":5.3925599999999996},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":12.78309},{"$f":"'DF raw data'!C16","v":10.62354},{"$f":"'DF raw data'!G16","v":-9.3064499000000005},{"$f":"'DF raw data'!H16","v":0.75528002000000005},{"$f":"'DF raw data'!L16","v":-9.4609498999999992},{"$f":"'DF raw data'!M16","v":18.8354},{"$f":"'DF raw data'!Q16","v":4.2277398000000002},{"$f":"'DF raw data'!R16","v":12.9221},{"$f":"'DF raw data'!V16","v":2.89886},{"$f":"'DF raw data'!W16","v":8.6486301000000001}]
[6,{"$f":"'DF raw data'!B17","v":5.6218399999999997},{"$f":"'DF raw data'!C17","v":9.3723496999999991},{"$f":"'DF raw data'!G17","v":0.78339999999999999},{"$f":"'DF raw data'!H17","v":3.3524200999999998},{"$f":"'DF raw data'!L17","v":15.58389},{"$f":"'DF raw data'!M17","v":0.84478003000000002},{"$f":"'DF raw data'!Q17","v":1.91371},{"$f":"'DF raw data'!R17","v":8.6376895999999999},{"$f":"'DF raw data'!V17","v":5.4394302000000003},{"$f":"'DF raw data'!W17","v":8.9556903999999999}]
[8,{"$f":"'DF raw data'!B18","v":-0.58820998999999996},{"$f":"'DF raw data'!C18","v":27.290279000000002},{"$f":"'DF raw data'!G18","v":-5.2730198000000001},{"$f":"'DF raw data'!H18","v":8.9783697},{"$f":"'DF raw data'!L18","v":-5.5685501000000004},{"$f":"'DF raw data'!M18","v":13.84679},{"$f":"'DF raw data'!Q18","v":2.4244599},{"$f":"'DF raw data'!R18","v":27.327850000000002},{"$f":"'DF raw data'!V18","v":-0.62132001000000003},{"$f":"'DF raw data'!W18","v":16.007959}]
[10,{"$f":"'DF raw data'!B19","v":-4.5724400999999997},{"$f":"'DF raw data'!C19","v":3.3575799000000002},{"$f":"'DF raw data'!G19","v":18.28126},{"$f":"'DF raw data'!H19","v":2.1994199999999999},{"$f":"'DF raw data'!L19","v":-12.51112},{"$f":"'DF raw data'!M19","v":10.37384},{"$f":"'DF raw data'!Q19","v":-6.74824},{"$f":"'DF raw data'!R19","v":-0.48910000999999997},{"$f":"'DF raw data'!V19","v":5.5015201999999999},{"$f":"'DF raw data'!W19","v":2.5143599999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":779},null,{"$f":"'DF raw data'!G4","v":775},null,{"$f":"'DF raw data'!L4","v":776},null,{"$f":"'DF raw data'!Q4","v":777},null,{"$f":"'DF raw data'!V4","v":778}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-0.63052996666666672},{"$f":"AVERAGE(C11:C13)","v":7.7629200666666662},1.0101199633333333,12.614923333333332,0.56548989999999988,9.2176065666666673,2.0411066966666667,9.9315965999999989,5.3219333666666664,5.5556801]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17842366},null,{"$f":"'DF raw data'!G7","v":17822322},null,{"$f":"'DF raw data'!L7","v":17802293},null,{"$f":"'DF raw data'!Q7","v":17762274},null,{"$f":"'DF raw data'!V7","v":17482308},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-5.792089466666666},{"$f":"C21","v":14.682597733333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":20.939562433333332},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":24.078886333333333}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-26.731651899999999},{"$f":"E21","v":33.771487333333333},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":20.939562433333332},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":24.432825333333334}]
[-10,{"$f":"'DF raw data'!D9/100","v":4.9433121},{"$f":"'DF raw data'!E9/100","v":3.0175800000000003E-2},{"$f":"'DF raw data'!I9/100","v":-20.493957999999999},{"$f":"'DF raw data'!J9/100","v":22.354452999999999},{"$f":"'DF raw data'!N9/100","v":20.307124999999999},{"$f":"'DF raw data'!O9/100","v":15.732773},{"$f":"'DF raw data'!S9/100","v":13.029806000000001},{"$f":"'DF raw data'!T9/100","v":3.8796386999999997},{"$f":"'DF raw data'!X9/100","v":1.7317180999999999},{"$f":"'DF raw data'!Y9/100","v":-16.106054999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":-10.002987966666668},{"$f":"G21","v":30.801653999999999}]
[-8,{"$f":"'DF raw data'!D10/100","v":-4.5072830000000002},{"$f":"'DF raw data'!E10/100","v":21.931777},{"$f":"'DF raw data'!I10/100","v":-4.7037677000000002},{"$f":"'DF raw data'!J10/100","v":57.811532999999997},{"$f":"'DF raw data'!N10/100","v":-1.1706555000000001},{"$f":"'DF raw data'!O10/100","v":29.515497999999997},{"$f":"'DF raw data'!S10/100","v":2.1870897999999999},{"$f":"'DF raw data'!T10/100","v":36.707568000000002},{"$f":"'DF raw data'!X10/100","v":-11.099088999999999},{"$f":"'DF raw data'!Y10/100","v":23.727471000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-7.6258895996666674},{"$f":"I21","v":9.6926009999999998}]
[-6,{"$f":"'DF raw data'!D11/100","v":4.5454094999999999},{"$f":"'DF raw data'!E11/100","v":-7.652793},{"$f":"'DF raw data'!I11/100","v":-13.190713000000001},{"$f":"'DF raw data'!J11/100","v":32.385331999999998},{"$f":"'DF raw data'!N11/100","v":-1.2189502000000001},{"$f":"'DF raw data'!O11/100","v":9.2491112999999991},{"$f":"'DF raw data'!S11/100","v":-0.44389277999999999},{"$f":"'DF raw data'!T11/100","v":-2.7436815999999999},{"$f":"'DF raw data'!X11/100","v":9.4928106999999997},{"$f":"'DF raw data'!Y11/100","v":19.495390999999998},null,{"$f":"J4","v":100},{"$f":"J21","v":-8.139330666666666},{"$f":"K21","v":9.3386620000000011}]
[-4,{"$f":"'DF raw data'!D12/100","v":-7.4571227999999996},{"$f":"'DF raw data'!E12/100","v":19.49999},{"$f":"'DF raw data'!I12/100","v":10.284509},{"$f":"'DF raw data'!J12/100","v":32.660302999999999},{"$f":"'DF raw data'!N12/100","v":40.476523},{"$f":"'DF raw data'!O12/100","v":9.8519042999999993},{"$f":"'DF raw data'!S12/100","v":21.973481000000003},{"$f":"'DF raw data'!T12/100","v":0.71578132999999999},{"$f":"'DF raw data'!X12/100","v":15.079957},{"$f":"'DF raw data'!Y12/100","v":15.630449}]
[-2,{"$f":"'DF raw data'!D13/100","v":-4.3454397999999994},{"$f":"'DF raw data'!E13/100","v":21.329423999999999},{"$f":"'DF raw data'!I13/100","v":-40.599792000000001},{"$f":"'DF raw data'!J13/100","v":26.086981999999999},{"$f":"'DF raw data'!N13/100","v":3.6305950999999999},{"$f":"'DF raw data'!O13/100","v":19.631104000000001},{"$f":"'DF raw data'!S13/100","v":-3.2007324000000001},{"$f":"'DF raw data'!T13/100","v":-15.909365000000001},{"$f":"'DF raw data'!X13/100","v":16.952931},{"$f":"'DF raw data'!Y13/100","v":21.035693000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-11.146217999999999},{"$f":"'DF raw data'!E14/100","v":-3.3331738},{"$f":"'DF raw data'!I14/100","v":-9.4154906999999994},{"$f":"'DF raw data'!J14/100","v":28.716190999999998},{"$f":"'DF raw data'!N14/100","v":-17.374037000000001},{"$f":"'DF raw data'!O14/100","v":23.312431999999998},{"$f":"'DF raw data'!S14/100","v":8.5506601000000002E-2},{"$f":"'DF raw data'!T14/100","v":-1.6517089999999999},{"$f":"'DF raw data'!X14/100","v":-25.494052999999997},{"$f":"'DF raw data'!Y14/100","v":-14.540215},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-1.8846106},{"$f":"'DF raw data'!E15/100","v":26.051543000000002},{"$f":"'DF raw data'!I15/100","v":-30.179672999999998},{"$f":"'DF raw data'!J15/100","v":46.511288999999998},{"$f":"'DF raw data'!N15/100","v":-16.265522000000001},{"$f":"'DF raw data'!O15/100","v":49.461426000000003},{"$f":"'DF raw data'!S15/100","v":-19.762443000000001},{"$f":"'DF raw data'!T15/100","v":46.638877000000001},{"$f":"'DF raw data'!X15/100","v":-15.876869999999998},{"$f":"'DF raw data'!Y15/100","v":21.520508},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":37.781708999999999},{"$f":"'DF raw data'!E16/100","v":23.924268000000001},{"$f":"'DF raw data'!I16/100","v":-6.9595910999999999},{"$f":"'DF raw data'!J16/100","v":3.5466797000000003},{"$f":"'DF raw data'!N16/100","v":-7.2840142999999999},{"$f":"'DF raw data'!O16/100","v":34.127637},{"$f":"'DF raw data'!S16/100","v":-9.4061132999999995},{"$f":"'DF raw data'!T16/100","v":38.461475},{"$f":"'DF raw data'!X16/100","v":-8.6624987999999998},{"$f":"'DF raw data'!Y16/100","v":37.041221}]
[6,{"$f":"'DF raw data'!D17/100","v":-26.558069},{"$f":"'DF raw data'!E17/100","v":6.3145703000000006},{"$f":"'DF raw data'!I17/100","v":25.665521999999999},{"$f":"'DF raw data'!J17/100","v":12.978261999999999},{"$f":"'DF raw data'!N17/100","v":-12.288955000000001},{"$f":"'DF raw data'!O17/100","v":6.5387890999999998},{"$f":"'DF raw data'!S17/100","v":-0.62018070000000003},{"$f":"'DF raw data'!T17/100","v":15.197011999999999},{"$f":"'DF raw data'!X17/100","v":-1.6776078999999999},{"$f":"'DF raw data'!Y17/100","v":40.222304999999999}]
[8,{"$f":"'DF raw data'!D18/100","v":-25.401584},{"$f":"'DF raw data'!E18/100","v":83.439374999999998},{"$f":"'DF raw data'!I18/100","v":-32.721876999999999},{"$f":"'DF raw data'!J18/100","v":22.883467},{"$f":"'DF raw data'!N18/100","v":-15.221575},{"$f":"'DF raw data'!O18/100","v":35.173642999999998},{"$f":"'DF raw data'!S18/100","v":-7.3526140999999994},{"$f":"'DF raw data'!T18/100","v":33.752148000000005},{"$f":"'DF raw data'!X18/100","v":-10.724128},{"$f":"'DF raw data'!Y18/100","v":19.972832}]
[10,{"$f":"'DF raw data'!D19/100","v":0.18599001000000001},{"$f":"'DF raw data'!E19/100","v":19.380399999999998},{"$f":"'DF raw data'!I19/100","v":18.468036999999999},{"$f":"'DF raw data'!J19/100","v":5.0873242000000003},{"$f":"'DF raw data'!N19/100","v":-42.373599000000006},{"$f":"'DF raw data'!O19/100","v":30.174364999999998},{"$f":"'DF raw data'!S19/100","v":10.041136999999999},{"$f":"'DF raw data'!T19/100","v":10.484931999999999},{"$f":"'DF raw data'!X19/100","v":8.4366228999999997},{"$f":"'DF raw data'!Y19/100","v":24.527294999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":779},null,{"$f":"'DF raw data'!G4","v":775},null,{"$f":"'DF raw data'!L4","v":776},null,{"$f":"'DF raw data'!Q4","v":777},null,{"$f":"'DF raw data'!V4","v":778}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-5.792089466666666},{"$f":"AVERAGE(C11:C13)","v":14.682597733333333},-26.731651899999999,33.771487333333333,-10.002987966666668,30.801653999999999,-7.6258895996666674,9.6926009999999998,-8.139330666666666,9.3386620000000011]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16317, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16317, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16317, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16317, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16317, 100 mm gap"]
["Date:","12. sep 2017","10:28:53",null,null,"Date:","12. sep 2017","09:55:18",null,null,"Date:","12. sep 2017","09:55:22",null,null,"Date:","12. sep 2017","09:55:26",null,null,"Date:","12. sep 2017","09:55:29"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",779,null,null,null,"Indice:",775,null,null,null,"Indice:",776,null,null,null,"Indice:",777,null,null,null,"Indice:",778]
["Background Indice:",779,null,null,null,"Background Indice:",775,null,null,null,"Background Indice:",776,null,null,null,"Background Indice:",777,null,null,null,"Background Indice:",778]
["Temperature:",21.8,null,null,null,"Temperature:",21.69,null,null,null,"Temperature:",21.71,null,null,null,"Temperature:",21.76,null,null,null,"Temperature:",21.74]
["Encoder:",17842366,null,null,null,"Encoder:",17822322,null,null,null,"Encoder:",17802293,null,null,null,"Encoder:",17762274,null,null,null,"Encoder:",17482308]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,9.3208303000000008,0.44112000000000001,494.33121,3.0175800000000002,-1,3.4060700000000002,3.6587100000000001,-2049.3957999999998,2235.4452999999999,-1,7.6890501999999996,10.32161,2030.7125000000001,1573.2773,-1,-1.2083499,6.1024899000000001,1302.9806000000001,387.96386999999999,-1,9.3722800999999993,-0.48019999000000002,173.17180999999999,-1610.6054999999999]
[-0.80000000999999998,6.1370300999999996,6.8191299000000001,-450.72829999999999,2193.1777000000002,-0.80000000999999998,9.7957000999999995,16.690650999999999,-470.37677000000002,5781.1532999999999,-0.80000000999999998,8.8227595999999995,9.3650999000000006,-117.06555,2951.5497999999998,-0.80000000999999998,-0.80615002000000002,8.7248696999999993,218.70898,3670.7568000000001,-0.80000000999999998,6.4730002999999994E-2,8.7134894999999997,-1109.9088999999999,2372.7471]
[-0.60000001999999997,5.0733699999999997,-1.65412,454.54095000000001,-765.27930000000003,-0.60000001999999997,6.7346301000000004,8.7563496000000001,-1319.0713000000001,3238.5331999999999,-0.60000001999999997,3.26966,6.5250401,-121.89502,924.91112999999996,-0.60000001999999997,0.57238001000000005,2.6145201,-44.389277999999997,-274.36815999999999,-0.60000001999999997,5.9566498000000001,6.7311902000000003,949.28107,1949.5391]
[-0.40000001000000002,0.95644998999999997,1.9544801000000001,-745.71227999999996,1949.999,-0.40000001000000002,7.3263001000000001,5.1712898999999997,1028.4509,3266.0302999999999,-0.40000001000000002,7.2519698000000004,0.99844003000000003,4047.6523000000002,985.19042999999999,-0.40000001000000002,0.19244,16.590851000000001,2197.3481000000002,71.578132999999994,-0.40000001000000002,8.6490402,1.67188,1507.9956999999999,1563.0449000000001]
[-0.2,-1.7002999999999999,5.9343399999999997,-434.54397999999998,2132.9423999999999,-0.2,0.38418998999999998,15.15962,-4059.9792000000002,2608.6981999999998,-0.2,0.67911999999999995,16.28013,363.05950999999999,1963.1104,-0.2,-0.93276000000000003,7.9956101999999998,-320.07324,-1590.9365,-0.2,8.2332801999999994,5.2156801000000002,1695.2931000000001,2103.5693000000001]
[0,5.6833701000000003,10.44064,-1114.6217999999999,-333.31738000000001,0,4.0641898999999997,10.129989999999999,-941.54907000000003,2871.6190999999999,0,8.0317898000000003,3.9219799000000002,-1737.4037000000001,2331.2431999999999,0,6.7142901000000004,9.4954795999999995,8.5506601,-165.17089999999999,0,5.4262300000000003,6.0588002000000003,-2549.4052999999999,-1454.0215000000001]
[0.2,-5.8746600000000004,6.9137801999999997,-188.46106,2605.1543000000001,0.2,-1.4180200000000001,12.555160000000001,-3017.9672999999998,4651.1288999999997,0.2,-7.0144400999999998,7.4507098000000003,-1626.5522000000001,4946.1426000000001,0.2,0.34178998999999999,12.303699999999999,-1976.2443000000001,4663.8877000000002,0.2,2.3062898999999999,5.3925599999999996,-1587.6869999999999,2152.0508]
[0.40000001000000002,12.78309,10.62354,3778.1709000000001,2392.4268000000002,0.40000001000000002,-9.3064499000000005,0.75528002000000005,-695.95911000000001,354.66797000000003,0.40000001000000002,-9.4609498999999992,18.8354,-728.40143,3412.7637,0.40000001000000002,4.2277398000000002,12.9221,-940.61132999999995,3846.1475,0.40000001000000002,2.89886,8.6486301000000001,-866.24987999999996,3704.1221]
[0.60000001999999997,5.6218399999999997,9.3723496999999991,-2655.8069,631.45703000000003,0.60000001999999997,0.78339999999999999,3.3524200999999998,2566.5522000000001,1297.8262,0.60000001999999997,15.58389,0.84478003000000002,-1228.8955000000001,653.87891000000002,0.60000001999999997,1.91371,8.6376895999999999,-62.018070000000002,1519.7012,0.60000001999999997,5.4394302000000003,8.9556903999999999,-167.76078999999999,4022.2305000000001]
[0.80000000999999998,-0.58820998999999996,27.290279000000002,-2540.1583999999998,8343.9375,0.80000000999999998,-5.2730198000000001,8.9783697,-3272.1876999999999,2288.3467000000001,0.80000000999999998,-5.5685501000000004,13.84679,-1522.1575,3517.3643000000002,0.80000000999999998,2.4244599,27.327850000000002,-735.26140999999996,3375.2148000000002,0.80000000999999998,-0.62132001000000003,16.007959,-1072.4128000000001,1997.2832000000001]
[1,-4.5724400999999997,3.3575799000000002,18.599001000000001,1938.04,1,18.28126,2.1994199999999999,1846.8036999999999,508.73241999999999,1,-12.51112,10.37384,-4237.3599000000004,3017.4364999999998,1,-6.74824,-0.48910000999999997,1004.1137,1048.4931999999999,1,5.5015201999999999,2.5143599999999999,843.66228999999998,2452.7294999999999]