{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":204,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":14640945},null,{"$f":"'DF raw data'!G7","v":14620894},null,{"$f":"'DF raw data'!L7","v":14600878},null,{"$f":"'DF raw data'!Q7","v":14560855},null,{"$f":"'DF raw data'!V7","v":14280903},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-2.8827966000000003},{"$f":"C21","v":-16.875626666666665},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":1.7081032333333337},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":10.838056666666667}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-2.7737666666666665},{"$f":"E21","v":-9.1799065999999989},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":2.0896432666666662},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":15.006223303333332}]
[-10,{"$f":"'DF raw data'!B9","v":0.81928003000000005},{"$f":"'DF raw data'!C9","v":7.1227001999999997},{"$f":"'DF raw data'!G9","v":2.1438600999999999},{"$f":"'DF raw data'!H9","v":5.0084800999999999},{"$f":"'DF raw data'!L9","v":13.08488},{"$f":"'DF raw data'!M9","v":8.7616195999999995},{"$f":"'DF raw data'!Q9","v":17.360689000000001},{"$f":"'DF raw data'!R9","v":4.4696698000000001},{"$f":"'DF raw data'!V9","v":-3.6776399999999998},{"$f":"'DF raw data'!W9","v":-13.37556},null,{"$f":"F4","v":20},{"$f":"F21","v":-1.1746933666666666},{"$f":"G21","v":-7.7598133666666671}]
[-8,{"$f":"'DF raw data'!B10","v":1.5812299999999999},{"$f":"'DF raw data'!C10","v":2.0959401},{"$f":"'DF raw data'!G10","v":19.376659},{"$f":"'DF raw data'!H10","v":4.9342598999999998},{"$f":"'DF raw data'!L10","v":-1.5348299999999999},{"$f":"'DF raw data'!M10","v":-12.60122},{"$f":"'DF raw data'!Q10","v":3.7920200999999998},{"$f":"'DF raw data'!R10","v":2.0483501},{"$f":"'DF raw data'!V10","v":11.918329999999999},{"$f":"'DF raw data'!W10","v":-2.3686099},null,{"$f":"H4","v":30},{"$f":"H21","v":-2.5549766666666667},{"$f":"I21","v":-6.0375699999999997},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":4.7165599},{"$f":"'DF raw data'!C11","v":1.13652},{"$f":"'DF raw data'!G11","v":-2.4973198999999999},{"$f":"'DF raw data'!H11","v":6.5220199000000001},{"$f":"'DF raw data'!L11","v":11.538460000000001},{"$f":"'DF raw data'!M11","v":0.23049},{"$f":"'DF raw data'!Q11","v":4.7482600000000001},{"$f":"'DF raw data'!R11","v":-15.10426},{"$f":"'DF raw data'!V11","v":6.1964401999999996},{"$f":"'DF raw data'!W11","v":3.9807301000000002},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.264336633333333},{"$f":"K21","v":-1.8694033633333333},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-3.3922501},{"$f":"'DF raw data'!C12","v":0.17946999999999999},{"$f":"'DF raw data'!G12","v":7.1367998000000004},{"$f":"'DF raw data'!H12","v":-13.664},{"$f":"'DF raw data'!L12","v":-4.9471002000000004},{"$f":"'DF raw data'!M12","v":-8.3260698000000009},{"$f":"'DF raw data'!Q12","v":14.627890000000001},{"$f":"'DF raw data'!R12","v":-2.7216000999999999},{"$f":"'DF raw data'!V12","v":1.9650000000000001},{"$f":"'DF raw data'!W12","v":-3.3030700999999998},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-5.8739699999999999},{"$f":"'DF raw data'!C13","v":-16.645399000000001},{"$f":"'DF raw data'!G13","v":-7.8859199999999996},{"$f":"'DF raw data'!H13","v":-10.71471},{"$f":"'DF raw data'!L13","v":-7.6270198999999996},{"$f":"'DF raw data'!M13","v":-6.3686398999999998},{"$f":"'DF raw data'!Q13","v":2.82125},{"$f":"'DF raw data'!R13","v":-4.6948800000000004},{"$f":"'DF raw data'!V13","v":-2.7128999},{"$f":"'DF raw data'!W13","v":-0.52395999000000004},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":3.35114},{"$f":"'DF raw data'!C14","v":-10.57549},{"$f":"'DF raw data'!G14","v":-7.6240601999999997},{"$f":"'DF raw data'!H14","v":-11.034409999999999},{"$f":"'DF raw data'!L14","v":4.7972397999999998},{"$f":"'DF raw data'!M14","v":-4.2483902000000002},{"$f":"'DF raw data'!Q14","v":-6.9199599999999997},{"$f":"'DF raw data'!R14","v":-1.7263900000000001},{"$f":"'DF raw data'!V14","v":-5.2594799999999999},{"$f":"'DF raw data'!W14","v":1.8301799999999999},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-6.1255598000000004},{"$f":"'DF raw data'!C15","v":-23.405991},{"$f":"'DF raw data'!G15","v":7.1886802000000003},{"$f":"'DF raw data'!H15","v":-5.7905997999999999},{"$f":"'DF raw data'!L15","v":-0.69430000000000003},{"$f":"'DF raw data'!M15","v":-12.662409999999999},{"$f":"'DF raw data'!Q15","v":-3.5662199999999999},{"$f":"'DF raw data'!R15","v":-11.69144},{"$f":"'DF raw data'!V15","v":-1.82063},{"$f":"'DF raw data'!W15","v":-6.9144300999999997},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-2.80565},{"$f":"'DF raw data'!C16","v":-17.206140999999999},{"$f":"'DF raw data'!G16","v":7.5327101000000001},{"$f":"'DF raw data'!H16","v":-14.369429999999999},{"$f":"'DF raw data'!L16","v":-6.9880800000000001},{"$f":"'DF raw data'!M16","v":-13.27075},{"$f":"'DF raw data'!Q16","v":5.6663398999999997},{"$f":"'DF raw data'!R16","v":-10.31542},{"$f":"'DF raw data'!V16","v":-2.3294499000000002},{"$f":"'DF raw data'!W16","v":-5.5697799000000003}]
[6,{"$f":"'DF raw data'!B17","v":-1.52644},{"$f":"'DF raw data'!C17","v":-24.367849},{"$f":"'DF raw data'!G17","v":-13.05045},{"$f":"'DF raw data'!H17","v":-29.007351},{"$f":"'DF raw data'!L17","v":-4.1361299000000002},{"$f":"'DF raw data'!M17","v":-10.000529999999999},{"$f":"'DF raw data'!Q17","v":13.417949999999999},{"$f":"'DF raw data'!R17","v":-22.605599999999999},{"$f":"'DF raw data'!V17","v":0.52919000000000005},{"$f":"'DF raw data'!W17","v":-4.0539398000000002}]
[8,{"$f":"'DF raw data'!B18","v":-1.143},{"$f":"'DF raw data'!C18","v":-13.363569999999999},{"$f":"'DF raw data'!G18","v":-5.2357702000000002},{"$f":"'DF raw data'!H18","v":-29.955759},{"$f":"'DF raw data'!L18","v":10.06073},{"$f":"'DF raw data'!M18","v":-23.685261000000001},{"$f":"'DF raw data'!Q18","v":4.1650200000000002},{"$f":"'DF raw data'!R18","v":-0.59064000999999999},{"$f":"'DF raw data'!V18","v":3.6301100000000002},{"$f":"'DF raw data'!W18","v":1.1906199}]
[10,{"$f":"'DF raw data'!B19","v":3.3633101000000001},{"$f":"'DF raw data'!C19","v":-15.745279999999999},{"$f":"'DF raw data'!G19","v":-1.5589999999999999},{"$f":"'DF raw data'!H19","v":-39.755519999999997},{"$f":"'DF raw data'!L19","v":11.676069999999999},{"$f":"'DF raw data'!M19","v":-15.26731},{"$f":"'DF raw data'!Q19","v":4.4831900999999998},{"$f":"'DF raw data'!R19","v":-0.24981998999999999},{"$f":"'DF raw data'!V19","v":-1.0403399},{"$f":"'DF raw data'!W19","v":-0.93725997000000005}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":551},null,{"$f":"'DF raw data'!G4","v":552},null,{"$f":"'DF raw data'!L4","v":553},null,{"$f":"'DF raw data'!Q4","v":554},null,{"$f":"'DF raw data'!V4","v":555}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-2.8827966000000003},{"$f":"AVERAGE(C11:C13)","v":-16.875626666666665},-2.7737666666666665,-9.1799065999999989,-1.1746933666666666,-7.7598133666666671,-2.5549766666666667,-6.0375699999999997,-3.264336633333333,-1.8694033633333333]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":14640945},null,{"$f":"'DF raw data'!G7","v":14620894},null,{"$f":"'DF raw data'!L7","v":14600878},null,{"$f":"'DF raw data'!Q7","v":14560855},null,{"$f":"'DF raw data'!V7","v":14280903},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-9.3241074766666667},{"$f":"C21","v":-16.073915866666663},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":8.251262999999998},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":25.769312933333328}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-16.409770599999998},{"$f":"E21","v":-19.317838599999998},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":20.949650633333331},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":25.769312933333328}]
[-10,{"$f":"'DF raw data'!D9/100","v":-38.997720000000001},{"$f":"'DF raw data'!E9/100","v":35.227412000000001},{"$f":"'DF raw data'!I9/100","v":13.589893999999999},{"$f":"'DF raw data'!J9/100","v":-18.280566},{"$f":"'DF raw data'!N9/100","v":34.754491999999999},{"$f":"'DF raw data'!O9/100","v":26.186239999999998},{"$f":"'DF raw data'!S9/100","v":19.487390999999999},{"$f":"'DF raw data'!T9/100","v":48.483350000000002},{"$f":"'DF raw data'!X9/100","v":-11.031965},{"$f":"'DF raw data'!Y9/100","v":-7.2210058999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":-8.1585076000000001},{"$f":"G21","v":-17.133062866666666}]
[-8,{"$f":"'DF raw data'!D10/100","v":2.2306211999999999},{"$f":"'DF raw data'!E10/100","v":2.9691797000000002},{"$f":"'DF raw data'!I10/100","v":35.094032999999996},{"$f":"'DF raw data'!J10/100","v":12.773203000000001},{"$f":"'DF raw data'!N10/100","v":28.942917000000001},{"$f":"'DF raw data'!O10/100","v":-44.529306999999996},{"$f":"'DF raw data'!S10/100","v":42.859603999999997},{"$f":"'DF raw data'!T10/100","v":29.018193},{"$f":"'DF raw data'!X10/100","v":5.5343255999999998},{"$f":"'DF raw data'!Y10/100","v":-29.062988000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-12.129181000000001},{"$f":"I21","v":6.4514743333333309}]
[-6,{"$f":"'DF raw data'!D11/100","v":0.88999267999999998},{"$f":"'DF raw data'!E11/100","v":2.7532323999999999},{"$f":"'DF raw data'!I11/100","v":-6.0045947000000002},{"$f":"'DF raw data'!J11/100","v":-18.536034999999998},{"$f":"'DF raw data'!N11/100","v":32.693917999999996},{"$f":"'DF raw data'!O11/100","v":-5.7525390999999999},{"$f":"'DF raw data'!S11/100","v":2.9855045000000002},{"$f":"'DF raw data'!T11/100","v":-0.82089844000000001},{"$f":"'DF raw data'!X11/100","v":8.6748119999999993},{"$f":"'DF raw data'!Y11/100","v":8.051513700000001},null,{"$f":"J4","v":100},{"$f":"J21","v":4.5398800333333336},{"$f":"K21","v":-4.1044920336666664}]
[-4,{"$f":"'DF raw data'!D12/100","v":-25.530087999999999},{"$f":"'DF raw data'!E12/100","v":5.5749805000000006},{"$f":"'DF raw data'!I12/100","v":-38.207759000000003},{"$f":"'DF raw data'!J12/100","v":-41.226679999999995},{"$f":"'DF raw data'!N12/100","v":-39.803110000000004},{"$f":"'DF raw data'!O12/100","v":9.5392675999999987},{"$f":"'DF raw data'!S12/100","v":33.714634000000004},{"$f":"'DF raw data'!T12/100","v":-0.19279301000000001},{"$f":"'DF raw data'!X12/100","v":-2.1775525},{"$f":"'DF raw data'!Y12/100","v":9.1300390999999994}]
[-2,{"$f":"'DF raw data'!D13/100","v":-15.953989},{"$f":"'DF raw data'!E13/100","v":-2.9803223000000001},{"$f":"'DF raw data'!I13/100","v":-36.05668},{"$f":"'DF raw data'!J13/100","v":4.2282422000000004},{"$f":"'DF raw data'!N13/100","v":-27.663928000000002},{"$f":"'DF raw data'!O13/100","v":-35.835234},{"$f":"'DF raw data'!S13/100","v":17.900361},{"$f":"'DF raw data'!T13/100","v":-6.8120410000000007},{"$f":"'DF raw data'!X13/100","v":1.0931848},{"$f":"'DF raw data'!Y13/100","v":6.0907520000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":0.85648857000000012},{"$f":"'DF raw data'!E14/100","v":-2.5064453000000002},{"$f":"'DF raw data'!I14/100","v":-20.999352999999999},{"$f":"'DF raw data'!J14/100","v":-22.825762000000001},{"$f":"'DF raw data'!N14/100","v":5.1741132000000007},{"$f":"'DF raw data'!O14/100","v":9.6388183999999999},{"$f":"'DF raw data'!S14/100","v":-38.274253000000002},{"$f":"'DF raw data'!T14/100","v":51.473397999999996},{"$f":"'DF raw data'!X14/100","v":20.867112000000002},{"$f":"'DF raw data'!Y14/100","v":3.4960898999999997E-2},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-12.874822},{"$f":"'DF raw data'!E15/100","v":-42.734979999999993},{"$f":"'DF raw data'!I15/100","v":7.8267211999999997},{"$f":"'DF raw data'!J15/100","v":-39.355995999999998},{"$f":"'DF raw data'!N15/100","v":-1.9857079999999998},{"$f":"'DF raw data'!O15/100","v":-25.202773000000001},{"$f":"'DF raw data'!S15/100","v":-16.013650999999999},{"$f":"'DF raw data'!T15/100","v":-25.306934000000002},{"$f":"'DF raw data'!X15/100","v":-8.3406567000000003},{"$f":"'DF raw data'!Y15/100","v":-18.439188999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":6.5591443000000007},{"$f":"'DF raw data'!E16/100","v":-21.331309000000001},{"$f":"'DF raw data'!I16/100","v":36.569724000000001},{"$f":"'DF raw data'!J16/100","v":-9.1063574000000003},{"$f":"'DF raw data'!N16/100","v":-6.9572778000000008},{"$f":"'DF raw data'!O16/100","v":-24.270907999999999},{"$f":"'DF raw data'!S16/100","v":2.0894225},{"$f":"'DF raw data'!T16/100","v":-10.930273},{"$f":"'DF raw data'!X16/100","v":0.61568789999999995},{"$f":"'DF raw data'!Y16/100","v":12.462998000000001}]
[6,{"$f":"'DF raw data'!D17/100","v":-6.4279126},{"$f":"'DF raw data'!E17/100","v":-22.577031000000002},{"$f":"'DF raw data'!I17/100","v":-27.111992000000001},{"$f":"'DF raw data'!J17/100","v":-63.908408000000001},{"$f":"'DF raw data'!N17/100","v":-10.443359000000001},{"$f":"'DF raw data'!O17/100","v":-43.460107000000001},{"$f":"'DF raw data'!S17/100","v":31.306646000000001},{"$f":"'DF raw data'!T17/100","v":-32.533760000000001},{"$f":"'DF raw data'!X17/100","v":4.3687433000000002},{"$f":"'DF raw data'!Y17/100","v":-23.328330000000001}]
[8,{"$f":"'DF raw data'!D18/100","v":-38.517415},{"$f":"'DF raw data'!E18/100","v":-28.297529000000001},{"$f":"'DF raw data'!I18/100","v":-7.7834009000000002},{"$f":"'DF raw data'!J18/100","v":-69.27373},{"$f":"'DF raw data'!N18/100","v":10.907128999999999},{"$f":"'DF raw data'!O18/100","v":-51.699179999999998},{"$f":"'DF raw data'!S18/100","v":19.29271},{"$f":"'DF raw data'!T18/100","v":19.396836},{"$f":"'DF raw data'!X18/100","v":-4.6704333},{"$f":"'DF raw data'!Y18/100","v":-9.4687598000000008}]
[10,{"$f":"'DF raw data'!D19/100","v":7.7953168000000002},{"$f":"'DF raw data'!E19/100","v":-21.897881000000002},{"$f":"'DF raw data'!I19/100","v":37.102907999999999},{"$f":"'DF raw data'!J19/100","v":-76.891582},{"$f":"'DF raw data'!N19/100","v":20.936387},{"$f":"'DF raw data'!O19/100","v":-14.352157999999999},{"$f":"'DF raw data'!S19/100","v":-23.752292000000001},{"$f":"'DF raw data'!T19/100","v":-3.7078516000000001},{"$f":"'DF raw data'!X19/100","v":-4.1757628999999996},{"$f":"'DF raw data'!Y19/100","v":-4.0527246000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":551},null,{"$f":"'DF raw data'!G4","v":552},null,{"$f":"'DF raw data'!L4","v":553},null,{"$f":"'DF raw data'!Q4","v":554},null,{"$f":"'DF raw data'!V4","v":555}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-9.3241074766666667},{"$f":"AVERAGE(C11:C13)","v":-16.073915866666663},-16.409770599999998,-19.317838599999998,-8.1585076000000001,-17.133062866666666,-12.129181000000001,6.4514743333333309,4.5398800333333336,-4.1044920336666664]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16319, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16319, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16319, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16319, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16319, 100 mm gap"]
["Date:","15. jun 2017","15:09:22",null,null,"Date:","15. jun 2017","15:09:29",null,null,"Date:","15. jun 2017","15:09:35",null,null,"Date:","15. jun 2017","15:09:41",null,null,"Date:","15. jun 2017","15:09:56"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",551,null,null,null,"Indice:",552,null,null,null,"Indice:",553,null,null,null,"Indice:",554,null,null,null,"Indice:",555]
["Background Indice:",551,null,null,null,"Background Indice:",552,null,null,null,"Background Indice:",553,null,null,null,"Background Indice:",554,null,null,null,"Background Indice:",555]
["Temperature:",21.78,null,null,null,"Temperature:",21.78,null,null,null,"Temperature:",21.89,null,null,null,"Temperature:",21.85,null,null,null,"Temperature:",21.87]
["Encoder:",14640945,null,null,null,"Encoder:",14620894,null,null,null,"Encoder:",14600878,null,null,null,"Encoder:",14560855,null,null,null,"Encoder:",14280903]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,0.81928003000000005,7.1227001999999997,-3899.7719999999999,3522.7411999999999,-1,2.1438600999999999,5.0084800999999999,1358.9893999999999,-1828.0565999999999,-1,13.08488,8.7616195999999995,3475.4492,2618.6239999999998,-1,17.360689000000001,4.4696698000000001,1948.7391,4848.335,-1,-3.6776399999999998,-13.37556,-1103.1965,-722.10059000000001]
[-0.80000000999999998,1.5812299999999999,2.0959401,223.06211999999999,296.91797000000003,-0.80000000999999998,19.376659,4.9342598999999998,3509.4032999999999,1277.3203000000001,-0.80000000999999998,-1.5348299999999999,-12.60122,2894.2917000000002,-4452.9306999999999,-0.80000000999999998,3.7920200999999998,2.0483501,4285.9603999999999,2901.8193000000001,-0.80000000999999998,11.918329999999999,-2.3686099,553.43255999999997,-2906.2988]
[-0.60000001999999997,4.7165599,1.13652,88.999268000000001,275.32324,-0.60000001999999997,-2.4973198999999999,6.5220199000000001,-600.45947000000001,-1853.6034999999999,-0.60000001999999997,11.538460000000001,0.23049,3269.3917999999999,-575.25391000000002,-0.60000001999999997,4.7482600000000001,-15.10426,298.55045000000001,-82.089843999999999,-0.60000001999999997,6.1964401999999996,3.9807301000000002,867.48119999999994,805.15137000000004]
[-0.40000001000000002,-3.3922501,0.17946999999999999,-2553.0088000000001,557.49805000000003,-0.40000001000000002,7.1367998000000004,-13.664,-3820.7759000000001,-4122.6679999999997,-0.40000001000000002,-4.9471002000000004,-8.3260698000000009,-3980.3110000000001,953.92675999999994,-0.40000001000000002,14.627890000000001,-2.7216000999999999,3371.4634000000001,-19.279301,-0.40000001000000002,1.9650000000000001,-3.3030700999999998,-217.75524999999999,913.00391000000002]
[-0.2,-5.8739699999999999,-16.645399000000001,-1595.3988999999999,-298.03223000000003,-0.2,-7.8859199999999996,-10.71471,-3605.6680000000001,422.82422000000003,-0.2,-7.6270198999999996,-6.3686398999999998,-2766.3928000000001,-3583.5234,-0.2,2.82125,-4.6948800000000004,1790.0361,-681.20410000000004,-0.2,-2.7128999,-0.52395999000000004,109.31847999999999,609.0752]
[0,3.35114,-10.57549,85.648857000000007,-250.64453,0,-7.6240601999999997,-11.034409999999999,-2099.9353000000001,-2282.5762,0,4.7972397999999998,-4.2483902000000002,517.41132000000005,963.88184000000001,0,-6.9199599999999997,-1.7263900000000001,-3827.4252999999999,5147.3397999999997,0,-5.2594799999999999,1.8301799999999999,2086.7112000000002,3.4960898999999999]
[0.2,-6.1255598000000004,-23.405991,-1287.4821999999999,-4273.4979999999996,0.2,7.1886802000000003,-5.7905997999999999,782.67211999999995,-3935.5996,0.2,-0.69430000000000003,-12.662409999999999,-198.57079999999999,-2520.2773000000002,0.2,-3.5662199999999999,-11.69144,-1601.3651,-2530.6934000000001,0.2,-1.82063,-6.9144300999999997,-834.06566999999995,-1843.9188999999999]
[0.40000001000000002,-2.80565,-17.206140999999999,655.91443000000004,-2133.1309000000001,0.40000001000000002,7.5327101000000001,-14.369429999999999,3656.9724000000001,-910.63574000000006,0.40000001000000002,-6.9880800000000001,-13.27075,-695.72778000000005,-2427.0907999999999,0.40000001000000002,5.6663398999999997,-10.31542,208.94225,-1093.0273,0.40000001000000002,-2.3294499000000002,-5.5697799000000003,61.56879,1246.2998]
[0.60000001999999997,-1.52644,-24.367849,-642.79125999999997,-2257.7031000000002,0.60000001999999997,-13.05045,-29.007351,-2711.1992,-6390.8407999999999,0.60000001999999997,-4.1361299000000002,-10.000529999999999,-1044.3359,-4346.0106999999998,0.60000001999999997,13.417949999999999,-22.605599999999999,3130.6646000000001,-3253.3760000000002,0.60000001999999997,0.52919000000000005,-4.0539398000000002,436.87432999999999,-2332.8330000000001]
[0.80000000999999998,-1.143,-13.363569999999999,-3851.7415000000001,-2829.7529,0.80000000999999998,-5.2357702000000002,-29.955759,-778.34009000000003,-6927.3729999999996,0.80000000999999998,10.06073,-23.685261000000001,1090.7129,-5169.9179999999997,0.80000000999999998,4.1650200000000002,-0.59064000999999999,1929.271,1939.6836000000001,0.80000000999999998,3.6301100000000002,1.1906199,-467.04333000000003,-946.87598000000003]
[1,3.3633101000000001,-15.745279999999999,779.53168000000005,-2189.7881000000002,1,-1.5589999999999999,-39.755519999999997,3710.2908000000002,-7689.1581999999999,1,11.676069999999999,-15.26731,2093.6387,-1435.2157999999999,1,4.4831900999999998,-0.24981998999999999,-2375.2292000000002,-370.78516000000002,1,-1.0403399,-0.93725997000000005,-417.57628999999997,-405.27246000000002]