{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":619,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17205685},null,{"$f":"'DF raw data'!G7","v":17185623},null,{"$f":"'DF raw data'!L7","v":17165619},null,{"$f":"'DF raw data'!Q7","v":17125612},null,{"$f":"'DF raw data'!V7","v":16845622},null,null,{"$f":"B4","v":10},{"$f":"B21","v":8.9443300666666659},{"$f":"C21","v":0.48921336999999993},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":12.466903366666665},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":5.6043365599999992}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":3.0834566000000003},{"$f":"E21","v":0.92108993000000006},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":12.830903266666667},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":5.6043365599999992}]
[-10,{"$f":"'DF raw data'!B9","v":-7.0297298000000001},{"$f":"'DF raw data'!C9","v":14.59043},{"$f":"'DF raw data'!G9","v":-13.445639999999999},{"$f":"'DF raw data'!H9","v":15.46918},{"$f":"'DF raw data'!L9","v":-8.8763199000000004},{"$f":"'DF raw data'!M9","v":13.511369999999999},{"$f":"'DF raw data'!Q9","v":-12.115959999999999},{"$f":"'DF raw data'!R9","v":11.2265},{"$f":"'DF raw data'!V9","v":-9.1050701000000007},{"$f":"'DF raw data'!W9","v":1.2463599000000001},null,{"$f":"F4","v":20},{"$f":"F21","v":2.8315265600000004},{"$f":"G21","v":4.6737166666666665}]
[-8,{"$f":"'DF raw data'!B10","v":14.54021},{"$f":"'DF raw data'!C10","v":0.72017998000000005},{"$f":"'DF raw data'!G10","v":0.81238001999999998},{"$f":"'DF raw data'!H10","v":8.9456100000000003},{"$f":"'DF raw data'!L10","v":-13.189629999999999},{"$f":"'DF raw data'!M10","v":16.76886},{"$f":"'DF raw data'!Q10","v":-13.23414},{"$f":"'DF raw data'!R10","v":7.2397799000000003},{"$f":"'DF raw data'!V10","v":-0.63593},{"$f":"'DF raw data'!W10","v":3.0084499999999998},null,{"$f":"H4","v":30},{"$f":"H21","v":-3.5225732999999999},{"$f":"I21","v":6.0935499299999991},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-0.93553001000000002},{"$f":"'DF raw data'!C11","v":3.3741099999999999},{"$f":"'DF raw data'!G11","v":-6.0019102000000002},{"$f":"'DF raw data'!H11","v":3.6424601000000001},{"$f":"'DF raw data'!L11","v":2.49546},{"$f":"'DF raw data'!M11","v":6.9374399000000002},{"$f":"'DF raw data'!Q11","v":-4.9493698999999998},{"$f":"'DF raw data'!R11","v":8.10989},{"$f":"'DF raw data'!V11","v":2.4826199999999998},{"$f":"'DF raw data'!W11","v":1.83544},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.8865732000000004},{"$f":"K21","v":2.7903933333333328},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":6.7049899000000002},{"$f":"'DF raw data'!C12","v":6.0610499000000004},{"$f":"'DF raw data'!G12","v":8.9014702000000003},{"$f":"'DF raw data'!H12","v":7.7463999000000001},{"$f":"'DF raw data'!L12","v":4.6663198000000001},{"$f":"'DF raw data'!M12","v":10.49724},{"$f":"'DF raw data'!Q12","v":-5.4356097999999999},{"$f":"'DF raw data'!R12","v":3.7202899},{"$f":"'DF raw data'!V12","v":-2.1681499},{"$f":"'DF raw data'!W12","v":6.5700402000000002},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":9.0926303999999991},{"$f":"'DF raw data'!C13","v":-0.50156999000000002},{"$f":"'DF raw data'!G13","v":3.9437399000000002},{"$f":"'DF raw data'!H13","v":-0.44890001000000002},{"$f":"'DF raw data'!L13","v":9.2033795999999999},{"$f":"'DF raw data'!M13","v":10.99869},{"$f":"'DF raw data'!Q13","v":-6.3983498000000001},{"$f":"'DF raw data'!R13","v":0.89473999000000004},{"$f":"'DF raw data'!V13","v":-8.4109697000000008},{"$f":"'DF raw data'!W13","v":-2.4237899999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":7.9745001999999996},{"$f":"'DF raw data'!C14","v":3.62656},{"$f":"'DF raw data'!G14","v":3.0701599000000002},{"$f":"'DF raw data'!H14","v":5.9427199000000002},{"$f":"'DF raw data'!L14","v":-1.2577699},{"$f":"'DF raw data'!M14","v":1.8323400000000001},{"$f":"'DF raw data'!Q14","v":-1.9063399999999999},{"$f":"'DF raw data'!R14","v":4.1329497999999996},{"$f":"'DF raw data'!V14","v":-2.0826498999999998},{"$f":"'DF raw data'!W14","v":6.7555499000000001},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":9.7658596000000006},{"$f":"'DF raw data'!C15","v":-1.6573499},{"$f":"'DF raw data'!G15","v":2.2364700000000002},{"$f":"'DF raw data'!H15","v":-2.7305500999999999},{"$f":"'DF raw data'!L15","v":0.54896997999999997},{"$f":"'DF raw data'!M15","v":1.1901200000000001},{"$f":"'DF raw data'!Q15","v":-2.2630300999999999},{"$f":"'DF raw data'!R15","v":13.25296},{"$f":"'DF raw data'!V15","v":-1.1660999999999999},{"$f":"'DF raw data'!W15","v":4.0394201000000001},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":13.826879999999999},{"$f":"'DF raw data'!C16","v":-0.46002999},{"$f":"'DF raw data'!G16","v":-2.1826601000000001},{"$f":"'DF raw data'!H16","v":-3.2020298999999999},{"$f":"'DF raw data'!L16","v":5.7380098999999998},{"$f":"'DF raw data'!M16","v":1.7561800000000001},{"$f":"'DF raw data'!Q16","v":-5.2564301000000002},{"$f":"'DF raw data'!R16","v":5.3543500999999996},{"$f":"'DF raw data'!V16","v":-1.07908},{"$f":"'DF raw data'!W16","v":-1.9423600000000001}]
[6,{"$f":"'DF raw data'!B17","v":4.13028},{"$f":"'DF raw data'!C17","v":1.18191},{"$f":"'DF raw data'!G17","v":2.8110198999999998},{"$f":"'DF raw data'!H17","v":7.0091299999999999},{"$f":"'DF raw data'!L17","v":-5.1872501},{"$f":"'DF raw data'!M17","v":3.1780700999999998},{"$f":"'DF raw data'!Q17","v":-2.2429299},{"$f":"'DF raw data'!R17","v":5.0250801999999997},{"$f":"'DF raw data'!V17","v":-0.27968999999999999},{"$f":"'DF raw data'!W17","v":3.0829200999999999}]
[8,{"$f":"'DF raw data'!B18","v":-0.67919998999999998},{"$f":"'DF raw data'!C18","v":6.0840997999999997},{"$f":"'DF raw data'!G18","v":-15.56076},{"$f":"'DF raw data'!H18","v":7.2852898000000001},{"$f":"'DF raw data'!L18","v":-2.0215399000000001},{"$f":"'DF raw data'!M18","v":30.534019000000001},{"$f":"'DF raw data'!Q18","v":-5.8400698000000002},{"$f":"'DF raw data'!R18","v":8.5669699000000001},{"$f":"'DF raw data'!V18","v":1.4005300000000001},{"$f":"'DF raw data'!W18","v":-0.26840999999999998}]
[10,{"$f":"'DF raw data'!B19","v":-8.9175196000000003},{"$f":"'DF raw data'!C19","v":19.817471000000001},{"$f":"'DF raw data'!G19","v":-13.6793},{"$f":"'DF raw data'!H19","v":26.502800000000001},{"$f":"'DF raw data'!L19","v":3.3620100000000002},{"$f":"'DF raw data'!M19","v":11.49295},{"$f":"'DF raw data'!Q19","v":-11.20093},{"$f":"'DF raw data'!R19","v":-2.3977100999999998},{"$f":"'DF raw data'!V19","v":-13.519830000000001},{"$f":"'DF raw data'!W19","v":-3.8279399999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":865},null,{"$f":"'DF raw data'!G4","v":866},null,{"$f":"'DF raw data'!L4","v":867},null,{"$f":"'DF raw data'!Q4","v":868},null,{"$f":"'DF raw data'!V4","v":869}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":8.9443300666666659},{"$f":"AVERAGE(C11:C13)","v":0.48921336999999993},3.0834566000000003,0.92108993000000006,2.8315265600000004,4.6737166666666665,-3.5225732999999999,6.0935499299999991,-3.8865732000000004,2.7903933333333328]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17205685},null,{"$f":"'DF raw data'!G7","v":17185623},null,{"$f":"'DF raw data'!L7","v":17165619},null,{"$f":"'DF raw data'!Q7","v":17125612},null,{"$f":"'DF raw data'!V7","v":16845622},null,null,{"$f":"B4","v":10},{"$f":"B21","v":6.5762311333333336},{"$f":"C21","v":-5.110602206666667},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":25.138407666666666},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":18.026471473333334}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":4.3927537000000001},{"$f":"E21","v":1.0375292000000005},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":27.820276366666668},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":18.766972576666664}]
[-10,{"$f":"'DF raw data'!D9/100","v":-3.1642621000000002},{"$f":"'DF raw data'!E9/100","v":12.899258},{"$f":"'DF raw data'!I9/100","v":-2.0784253000000001},{"$f":"'DF raw data'!J9/100","v":16.539541},{"$f":"'DF raw data'!N9/100","v":1.8820956000000002},{"$f":"'DF raw data'!O9/100","v":-2.4158105000000001},{"$f":"'DF raw data'!S9/100","v":-7.8960442999999998},{"$f":"'DF raw data'!T9/100","v":7.1720507999999992},{"$f":"'DF raw data'!X9/100","v":-10.304429000000001},{"$f":"'DF raw data'!Y9/100","v":-0.38460940999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":16.505407533333333},{"$f":"G21","v":12.393222633333332}]
[-8,{"$f":"'DF raw data'!D10/100","v":21.796187},{"$f":"'DF raw data'!E10/100","v":-8.9812694999999998},{"$f":"'DF raw data'!I10/100","v":-3.3916739000000002},{"$f":"'DF raw data'!J10/100","v":28.933145},{"$f":"'DF raw data'!N10/100","v":-21.458595999999996},{"$f":"'DF raw data'!O10/100","v":45.669970999999997},{"$f":"'DF raw data'!S10/100","v":-23.040199999999999},{"$f":"'DF raw data'!T10/100","v":23.364756000000003},{"$f":"'DF raw data'!X10/100","v":-1.1249921000000001},{"$f":"'DF raw data'!Y10/100","v":11.733184000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-8.6330001333333328},{"$f":"I21","v":12.915869266666668}]
[-6,{"$f":"'DF raw data'!D11/100","v":-4.5984875000000001},{"$f":"'DF raw data'!E11/100","v":-8.1088867000000011},{"$f":"'DF raw data'!I11/100","v":10.652079000000001},{"$f":"'DF raw data'!J11/100","v":9.6030663999999994},{"$f":"'DF raw data'!N11/100","v":5.8928601},{"$f":"'DF raw data'!O11/100","v":3.5270996000000001},{"$f":"'DF raw data'!S11/100","v":-9.7617969000000002},{"$f":"'DF raw data'!T11/100","v":17.202822000000001},{"$f":"'DF raw data'!X11/100","v":2.5291063999999999},{"$f":"'DF raw data'!Y11/100","v":1.8133594},null,{"$f":"J4","v":100},{"$f":"J21","v":-11.314868833333334},{"$f":"K21","v":13.656370369999998}]
[-4,{"$f":"'DF raw data'!D12/100","v":9.6791210999999997},{"$f":"'DF raw data'!E12/100","v":17.102578000000001},{"$f":"'DF raw data'!I12/100","v":10.183052},{"$f":"'DF raw data'!J12/100","v":12.583534999999999},{"$f":"'DF raw data'!N12/100","v":3.3236816},{"$f":"'DF raw data'!O12/100","v":20.983222999999999},{"$f":"'DF raw data'!S12/100","v":-9.9588342000000001},{"$f":"'DF raw data'!T12/100","v":14.934951000000002},{"$f":"'DF raw data'!X12/100","v":-15.275778000000001},{"$f":"'DF raw data'!Y12/100","v":-26.827002000000004}]
[-2,{"$f":"'DF raw data'!D13/100","v":8.5257982999999999},{"$f":"'DF raw data'!E13/100","v":-0.98448241999999997},{"$f":"'DF raw data'!I13/100","v":-2.8357697000000002},{"$f":"'DF raw data'!J13/100","v":-3.8300781000000002},{"$f":"'DF raw data'!N13/100","v":6.1789282000000005},{"$f":"'DF raw data'!O13/100","v":21.906621000000001},{"$f":"'DF raw data'!S13/100","v":-14.245326},{"$f":"'DF raw data'!T13/100","v":7.3743164000000005},{"$f":"'DF raw data'!X13/100","v":-28.737983},{"$f":"'DF raw data'!Y13/100","v":0.85769530999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-4.5102319},{"$f":"'DF raw data'!E14/100","v":-12.206582000000001},{"$f":"'DF raw data'!I14/100","v":9.7922399999999996},{"$f":"'DF raw data'!J14/100","v":13.338027},{"$f":"'DF raw data'!N14/100","v":5.4316063999999997},{"$f":"'DF raw data'!O14/100","v":5.3713183999999998},{"$f":"'DF raw data'!S14/100","v":-8.5683703999999992},{"$f":"'DF raw data'!T14/100","v":5.9344824000000003},{"$f":"'DF raw data'!X14/100","v":-7.0865746999999999},{"$f":"'DF raw data'!Y14/100","v":36.122323999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":15.713127},{"$f":"'DF raw data'!E15/100","v":-2.1407422},{"$f":"'DF raw data'!I15/100","v":6.2217908},{"$f":"'DF raw data'!J15/100","v":-6.3953612999999994},{"$f":"'DF raw data'!N15/100","v":37.905687999999998},{"$f":"'DF raw data'!O15/100","v":9.9017285000000008},{"$f":"'DF raw data'!S15/100","v":-3.0853039999999998},{"$f":"'DF raw data'!T15/100","v":25.438809000000003},{"$f":"'DF raw data'!X15/100","v":1.8799511999999998},{"$f":"'DF raw data'!Y15/100","v":3.9890917999999997},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":15.222213999999999},{"$f":"'DF raw data'!E16/100","v":-3.8895508000000003},{"$f":"'DF raw data'!I16/100","v":14.040657000000001},{"$f":"'DF raw data'!J16/100","v":-0.55015628999999999},{"$f":"'DF raw data'!N16/100","v":-1.183457},{"$f":"'DF raw data'!O16/100","v":-38.303632999999998},{"$f":"'DF raw data'!S16/100","v":9.5992516999999999},{"$f":"'DF raw data'!T16/100","v":11.181357},{"$f":"'DF raw data'!X16/100","v":2.3304781999999999},{"$f":"'DF raw data'!Y16/100","v":-7.0575976999999996}]
[6,{"$f":"'DF raw data'!D17/100","v":57.068857000000001},{"$f":"'DF raw data'!E17/100","v":-4.7670117000000003},{"$f":"'DF raw data'!I17/100","v":25.002336},{"$f":"'DF raw data'!J17/100","v":-3.1921776999999998},{"$f":"'DF raw data'!N17/100","v":1.3524411000000001},{"$f":"'DF raw data'!O17/100","v":-9.7759180000000008},{"$f":"'DF raw data'!S17/100","v":-3.2096312999999999},{"$f":"'DF raw data'!T17/100","v":8.0036523000000006},{"$f":"'DF raw data'!X17/100","v":-0.27083370000000001},{"$f":"'DF raw data'!Y17/100","v":-0.19957028999999998}]
[8,{"$f":"'DF raw data'!D18/100","v":-2.0202062999999999},{"$f":"'DF raw data'!E18/100","v":25.985166},{"$f":"'DF raw data'!I18/100","v":-30.810164},{"$f":"'DF raw data'!J18/100","v":37.171171999999999},{"$f":"'DF raw data'!N18/100","v":14.293239000000002},{"$f":"'DF raw data'!O18/100","v":53.104530999999994},{"$f":"'DF raw data'!S18/100","v":-8.1064685000000001},{"$f":"'DF raw data'!T18/100","v":7.3844335999999995},{"$f":"'DF raw data'!X18/100","v":18.997021},{"$f":"'DF raw data'!Y18/100","v":-3.2749901999999995}]
[10,{"$f":"'DF raw data'!D19/100","v":-38.807856000000001},{"$f":"'DF raw data'!E19/100","v":32.128770000000003},{"$f":"'DF raw data'!I19/100","v":-28.084320999999999},{"$f":"'DF raw data'!J19/100","v":29.683672000000001},{"$f":"'DF raw data'!N19/100","v":32.923732999999999},{"$f":"'DF raw data'!O19/100","v":8.2482714999999995},{"$f":"'DF raw data'!S19/100","v":-9.0606793000000003},{"$f":"'DF raw data'!T19/100","v":-5.8714648},{"$f":"'DF raw data'!X19/100","v":-15.811923999999999},{"$f":"'DF raw data'!Y19/100","v":-6.0759766000000006}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":865},null,{"$f":"'DF raw data'!G4","v":866},null,{"$f":"'DF raw data'!L4","v":867},null,{"$f":"'DF raw data'!Q4","v":868},null,{"$f":"'DF raw data'!V4","v":869}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":6.5762311333333336},{"$f":"AVERAGE(C11:C13)","v":-5.110602206666667},4.3927537000000001,1.0375292000000005,16.505407533333333,12.393222633333332,-8.6330001333333328,12.915869266666668,-11.314868833333334,13.656370369999998]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16321, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16321, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16321, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16321, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16321, 100 mm gap"]
["Date:","28. sep 2017","16:35:50",null,null,"Date:","28. sep 2017","16:49:10",null,null,"Date:","28. sep 2017","17:01:50",null,null,"Date:","28. sep 2017","17:26:25",null,null,"Date:","28. sep 2017","17:41:54"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",865,null,null,null,"Indice:",866,null,null,null,"Indice:",867,null,null,null,"Indice:",868,null,null,null,"Indice:",869]
["Background Indice:",865,null,null,null,"Background Indice:",866,null,null,null,"Background Indice:",867,null,null,null,"Background Indice:",868,null,null,null,"Background Indice:",869]
["Temperature:",22.03,null,null,null,"Temperature:",22,null,null,null,"Temperature:",21.96,null,null,null,"Temperature:",22.01,null,null,null,"Temperature:",21.96]
["Encoder:",17205685,null,null,null,"Encoder:",17185623,null,null,null,"Encoder:",17165619,null,null,null,"Encoder:",17125612,null,null,null,"Encoder:",16845622]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-7.0297298000000001,14.59043,-316.42621000000003,1289.9258,-1,-13.445639999999999,15.46918,-207.84253000000001,1653.9540999999999,-1,-8.8763199000000004,13.511369999999999,188.20956000000001,-241.58105,-1,-12.115959999999999,11.2265,-789.60442999999998,717.20507999999995,-1,-9.1050701000000007,1.2463599000000001,-1030.4429,-38.460940999999998]
[-0.80000000999999998,14.54021,0.72017998000000005,2179.6187,-898.12694999999997,-0.80000000999999998,0.81238001999999998,8.9456100000000003,-339.16739000000001,2893.3145,-0.80000000999999998,-13.189629999999999,16.76886,-2145.8595999999998,4566.9970999999996,-0.80000000999999998,-13.23414,7.2397799000000003,-2304.02,2336.4756000000002,-0.80000000999999998,-0.63593,3.0084499999999998,-112.49921000000001,1173.3184000000001]
[-0.60000001999999997,-0.93553001000000002,3.3741099999999999,-459.84875,-810.88867000000005,-0.60000001999999997,-6.0019102000000002,3.6424601000000001,1065.2079000000001,960.30664000000002,-0.60000001999999997,2.49546,6.9374399000000002,589.28601000000003,352.70996000000002,-0.60000001999999997,-4.9493698999999998,8.10989,-976.17969000000005,1720.2822000000001,-0.60000001999999997,2.4826199999999998,1.83544,252.91064,181.33593999999999]
[-0.40000001000000002,6.7049899000000002,6.0610499000000004,967.91210999999998,1710.2578000000001,-0.40000001000000002,8.9014702000000003,7.7463999000000001,1018.3052,1258.3534999999999,-0.40000001000000002,4.6663198000000001,10.49724,332.36815999999999,2098.3222999999998,-0.40000001000000002,-5.4356097999999999,3.7202899,-995.88342,1493.4951000000001,-0.40000001000000002,-2.1681499,6.5700402000000002,-1527.5778,-2682.7002000000002]
[-0.2,9.0926303999999991,-0.50156999000000002,852.57983000000002,-98.448241999999993,-0.2,3.9437399000000002,-0.44890001000000002,-283.57697000000002,-383.00781000000001,-0.2,9.2033795999999999,10.99869,617.89282000000003,2190.6621,-0.2,-6.3983498000000001,0.89473999000000004,-1424.5326,737.43164000000002,-0.2,-8.4109697000000008,-2.4237899999999999,-2873.7982999999999,85.769531000000001]
[0,7.9745001999999996,3.62656,-451.02319,-1220.6582000000001,0,3.0701599000000002,5.9427199000000002,979.22400000000005,1333.8027,0,-1.2577699,1.8323400000000001,543.16063999999994,537.13184000000001,0,-1.9063399999999999,4.1329497999999996,-856.83704,593.44824000000006,0,-2.0826498999999998,6.7555499000000001,-708.65746999999999,3612.2323999999999]
[0.2,9.7658596000000006,-1.6573499,1571.3126999999999,-214.07422,0.2,2.2364700000000002,-2.7305500999999999,622.17908,-639.53612999999996,0.2,0.54896997999999997,1.1901200000000001,3790.5688,990.17285000000004,0.2,-2.2630300999999999,13.25296,-308.53039999999999,2543.8809000000001,0.2,-1.1660999999999999,4.0394201000000001,187.99511999999999,398.90917999999999]
[0.40000001000000002,13.826879999999999,-0.46002999,1522.2213999999999,-388.95508000000001,0.40000001000000002,-2.1826601000000001,-3.2020298999999999,1404.0657000000001,-55.015628999999997,0.40000001000000002,5.7380098999999998,1.7561800000000001,-118.34569999999999,-3830.3633,0.40000001000000002,-5.2564301000000002,5.3543500999999996,959.92516999999998,1118.1357,0.40000001000000002,-1.07908,-1.9423600000000001,233.04782,-705.75977]
[0.60000001999999997,4.13028,1.18191,5706.8856999999998,-476.70116999999999,0.60000001999999997,2.8110198999999998,7.0091299999999999,2500.2336,-319.21776999999997,0.60000001999999997,-5.1872501,3.1780700999999998,135.24411000000001,-977.59180000000003,0.60000001999999997,-2.2429299,5.0250801999999997,-320.96312999999998,800.36523,0.60000001999999997,-0.27968999999999999,3.0829200999999999,-27.083369999999999,-19.957028999999999]
[0.80000000999999998,-0.67919998999999998,6.0840997999999997,-202.02063000000001,2598.5165999999999,0.80000000999999998,-15.56076,7.2852898000000001,-3081.0164,3717.1172000000001,0.80000000999999998,-2.0215399000000001,30.534019000000001,1429.3239000000001,5310.4530999999997,0.80000000999999998,-5.8400698000000002,8.5669699000000001,-810.64684999999997,738.44335999999998,0.80000000999999998,1.4005300000000001,-0.26840999999999998,1899.7021,-327.49901999999997]
[1,-8.9175196000000003,19.817471000000001,-3880.7856000000002,3212.877,1,-13.6793,26.502800000000001,-2808.4321,2968.3672000000001,1,3.3620100000000002,11.49295,3292.3733000000002,824.82714999999996,1,-11.20093,-2.3977100999999998,-906.06793000000005,-587.14648,1,-13.519830000000001,-3.8279399999999999,-1581.1923999999999,-607.59766000000002]