{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":569,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":25018409},null,{"$f":"'DF raw data'!G7","v":24998365},null,{"$f":"'DF raw data'!L7","v":24978350},null,{"$f":"'DF raw data'!Q7","v":24938332},null,{"$f":"'DF raw data'!V7","v":24658334},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-3.7530666666666668},{"$f":"C21","v":-8.8066399999999998},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":3.3692299666666665},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":6.5460833666666662}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-3.8808333333333334},{"$f":"E21","v":-7.4130931000000002},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":3.3692299666666665},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":6.5460833666666662}]
[-10,{"$f":"'DF raw data'!B9","v":-19.533190000000001},{"$f":"'DF raw data'!C9","v":-4.9932499000000004},{"$f":"'DF raw data'!G9","v":-11.289149999999999},{"$f":"'DF raw data'!H9","v":6.5816998},{"$f":"'DF raw data'!L9","v":-11.37425},{"$f":"'DF raw data'!M9","v":5.6534700000000004},{"$f":"'DF raw data'!Q9","v":-6.2302999000000003},{"$f":"'DF raw data'!R9","v":2.8783099999999999},{"$f":"'DF raw data'!V9","v":-12.17393},{"$f":"'DF raw data'!W9","v":-4.7520099},null,{"$f":"F4","v":20},{"$f":"F21","v":-0.51160336666666684},{"$f":"G21","v":-2.2605566333333331}]
[-8,{"$f":"'DF raw data'!B10","v":-18.725071},{"$f":"'DF raw data'!C10","v":-5.1249399000000002},{"$f":"'DF raw data'!G10","v":-16.430820000000001},{"$f":"'DF raw data'!H10","v":1.50665},{"$f":"'DF raw data'!L10","v":-21.665929999999999},{"$f":"'DF raw data'!M10","v":-3.1786699},{"$f":"'DF raw data'!Q10","v":-3.1273700999999998},{"$f":"'DF raw data'!R10","v":-10.56099},{"$f":"'DF raw data'!V10","v":-1.8562799999999999},{"$f":"'DF raw data'!W10","v":-8.8237400000000008},null,{"$f":"H4","v":30},{"$f":"H21","v":-0.81609996666666673},{"$f":"I21","v":-2.5776466433333334},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-21.920850999999999},{"$f":"'DF raw data'!C11","v":-9.8650999000000006},{"$f":"'DF raw data'!G11","v":-13.58942},{"$f":"'DF raw data'!H11","v":-1.9942299999999999},{"$f":"'DF raw data'!L11","v":-2.7846500999999999},{"$f":"'DF raw data'!M11","v":-1.5103899999999999},{"$f":"'DF raw data'!Q11","v":-3.0009301000000002},{"$f":"'DF raw data'!R11","v":-3.3293300000000001},{"$f":"'DF raw data'!V11","v":-6.3056597999999999},{"$f":"'DF raw data'!W11","v":-3.91751},null,{"$f":"J4","v":100},{"$f":"J21","v":-2.00502},{"$f":"K21","v":-7.4460266333333331},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-12.020110000000001},{"$f":"'DF raw data'!C12","v":-11.173579999999999},{"$f":"'DF raw data'!G12","v":-4.0741700999999999},{"$f":"'DF raw data'!H12","v":-4.7596698000000002},{"$f":"'DF raw data'!L12","v":-2.24593},{"$f":"'DF raw data'!M12","v":-13.5321},{"$f":"'DF raw data'!Q12","v":-0.45738000000000001},{"$f":"'DF raw data'!R12","v":-6.8550601000000002},{"$f":"'DF raw data'!V12","v":2.9849000000000001},{"$f":"'DF raw data'!W12","v":-11.01285},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-7.3610601000000004},{"$f":"'DF raw data'!C13","v":-10.43186},{"$f":"'DF raw data'!G13","v":-12.777010000000001},{"$f":"'DF raw data'!H13","v":-8.9787196999999992},{"$f":"'DF raw data'!L13","v":-3.1629200000000002},{"$f":"'DF raw data'!M13","v":2.5627},{"$f":"'DF raw data'!Q13","v":-6.5287499000000002},{"$f":"'DF raw data'!R13","v":-6.0571799000000004},{"$f":"'DF raw data'!V13","v":-1.3869100000000001},{"$f":"'DF raw data'!W13","v":-9.3009194999999991},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":3.9044599999999998},{"$f":"'DF raw data'!C14","v":-8.7065000999999995},{"$f":"'DF raw data'!G14","v":4.5299500999999998},{"$f":"'DF raw data'!H14","v":-7.8070798000000003},{"$f":"'DF raw data'!L14","v":3.7257699999999998},{"$f":"'DF raw data'!M14","v":-2.2187399999999999},{"$f":"'DF raw data'!Q14","v":5.3381499999999997},{"$f":"'DF raw data'!R14","v":-0.69875001999999997},{"$f":"'DF raw data'!V14","v":-1.43784},{"$f":"'DF raw data'!W14","v":-4.8821301000000004},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-7.8025998999999997},{"$f":"'DF raw data'!C15","v":-7.2815599000000004},{"$f":"'DF raw data'!G15","v":-3.3954401000000001},{"$f":"'DF raw data'!H15","v":-5.4534798000000002},{"$f":"'DF raw data'!L15","v":-2.0976601000000001},{"$f":"'DF raw data'!M15","v":-7.1256298999999999},{"$f":"'DF raw data'!Q15","v":-1.2577},{"$f":"'DF raw data'!R15","v":-0.97701000999999998},{"$f":"'DF raw data'!V15","v":-3.1903100000000002},{"$f":"'DF raw data'!W15","v":-8.1550303},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-5.4329599999999996},{"$f":"'DF raw data'!C16","v":2.9762298999999999},{"$f":"'DF raw data'!G16","v":-6.4324398},{"$f":"'DF raw data'!H16","v":2.1565398999999998},{"$f":"'DF raw data'!L16","v":-3.84653},{"$f":"'DF raw data'!M16","v":-7.1699099999999998},{"$f":"'DF raw data'!Q16","v":-4.3070997999999996},{"$f":"'DF raw data'!R16","v":-1.1812400000000001},{"$f":"'DF raw data'!V16","v":-13.79401},{"$f":"'DF raw data'!W16","v":-5.3373599}]
[6,{"$f":"'DF raw data'!B17","v":5.4994301999999999},{"$f":"'DF raw data'!C17","v":-0.27605998999999998},{"$f":"'DF raw data'!G17","v":6.2117801000000004},{"$f":"'DF raw data'!H17","v":-1.84988},{"$f":"'DF raw data'!L17","v":1.1713800000000001},{"$f":"'DF raw data'!M17","v":-3.9751699},{"$f":"'DF raw data'!Q17","v":2.30707},{"$f":"'DF raw data'!R17","v":-4.0368298999999999},{"$f":"'DF raw data'!V17","v":-4.1105399},{"$f":"'DF raw data'!W17","v":-4.8217502000000003}]
[8,{"$f":"'DF raw data'!B18","v":22.64612},{"$f":"'DF raw data'!C18","v":-5.9265099000000001},{"$f":"'DF raw data'!G18","v":9.1174698000000003},{"$f":"'DF raw data'!H18","v":-5.9156199000000003},{"$f":"'DF raw data'!L18","v":10.20116},{"$f":"'DF raw data'!M18","v":-8.8188896000000003},{"$f":"'DF raw data'!Q18","v":8.5224600000000006},{"$f":"'DF raw data'!R18","v":-7.6382298000000004},{"$f":"'DF raw data'!V18","v":-0.34147999000000001},{"$f":"'DF raw data'!W18","v":-16.028690000000001}]
[10,{"$f":"'DF raw data'!B19","v":19.556619999999999},{"$f":"'DF raw data'!C19","v":12.892860000000001},{"$f":"'DF raw data'!G19","v":12.397259999999999},{"$f":"'DF raw data'!H19","v":3.1667299},{"$f":"'DF raw data'!L19","v":13.459390000000001},{"$f":"'DF raw data'!M19","v":2.2527001000000002},{"$f":"'DF raw data'!Q19","v":7.7291898999999997},{"$f":"'DF raw data'!R19","v":-0.64303999999999994},{"$f":"'DF raw data'!V19","v":-2.89059},{"$f":"'DF raw data'!W19","v":-7.4324002}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":802},null,{"$f":"'DF raw data'!G4","v":803},null,{"$f":"'DF raw data'!L4","v":804},null,{"$f":"'DF raw data'!Q4","v":805},null,{"$f":"'DF raw data'!V4","v":806}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-3.7530666666666668},{"$f":"AVERAGE(C11:C13)","v":-8.8066399999999998},-3.8808333333333334,-7.4130931000000002,-0.51160336666666684,-2.2605566333333331,-0.81609996666666673,-2.5776466433333334,-2.00502,-7.4460266333333331]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":25018409},null,{"$f":"'DF raw data'!G7","v":24998365},null,{"$f":"'DF raw data'!L7","v":24978350},null,{"$f":"'DF raw data'!Q7","v":24938332},null,{"$f":"'DF raw data'!V7","v":24658334},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-8.7974799666666659},{"$f":"C21","v":-14.533554466666667},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":14.889601300000002},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":7.2920408333333331}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-20.158763666666669},{"$f":"E21","v":-14.354319433333332},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":15.206365403333336},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":7.2920408333333331}]
[-10,{"$f":"'DF raw data'!D9/100","v":-28.189099000000002},{"$f":"'DF raw data'!E9/100","v":-21.587422},{"$f":"'DF raw data'!I9/100","v":-15.362693999999999},{"$f":"'DF raw data'!J9/100","v":-9.7716016000000003},{"$f":"'DF raw data'!N9/100","v":-18.637613999999999},{"$f":"'DF raw data'!O9/100","v":-9.3385155999999991},{"$f":"'DF raw data'!S9/100","v":-1.0719327000000001},{"$f":"'DF raw data'!T9/100","v":-19.239795000000001},{"$f":"'DF raw data'!X9/100","v":-11.527904000000001},{"$f":"'DF raw data'!Y9/100","v":-22.972753999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":-13.034907099999998},{"$f":"G21","v":-10.638851033333333}]
[-8,{"$f":"'DF raw data'!D10/100","v":-40.459828999999999},{"$f":"'DF raw data'!E10/100","v":-16.342227000000001},{"$f":"'DF raw data'!I10/100","v":-20.934823999999999},{"$f":"'DF raw data'!J10/100","v":-5.7902148000000002},{"$f":"'DF raw data'!N10/100","v":-28.683872000000001},{"$f":"'DF raw data'!O10/100","v":-13.58541},{"$f":"'DF raw data'!S10/100","v":-8.2873424999999994},{"$f":"'DF raw data'!T10/100","v":-21.209268000000002},{"$f":"'DF raw data'!X10/100","v":-6.7122601},{"$f":"'DF raw data'!Y10/100","v":-16.52957},null,{"$f":"H4","v":30},{"$f":"H21","v":-5.2691623666666665},{"$f":"I21","v":-17.930891866666666}]
[-6,{"$f":"'DF raw data'!D11/100","v":-37.425736999999998},{"$f":"'DF raw data'!E11/100","v":-23.542559000000001},{"$f":"'DF raw data'!I11/100","v":-25.892786000000001},{"$f":"'DF raw data'!J11/100","v":-14.340897999999999},{"$f":"'DF raw data'!N11/100","v":0.44625229},{"$f":"'DF raw data'!O11/100","v":-21.835868999999999},{"$f":"'DF raw data'!S11/100","v":-6.1601287999999998},{"$f":"'DF raw data'!T11/100","v":-9.7334375000000009},{"$f":"'DF raw data'!X11/100","v":-13.083557000000001},{"$f":"'DF raw data'!Y11/100","v":-23.339931999999997},null,{"$f":"J4","v":100},{"$f":"J21","v":-4.9523982633333334},{"$f":"K21","v":-15.607236300000002}]
[-4,{"$f":"'DF raw data'!D12/100","v":-30.616680000000002},{"$f":"'DF raw data'!E12/100","v":-25.522627000000004},{"$f":"'DF raw data'!I12/100","v":-9.6030151000000004},{"$f":"'DF raw data'!J12/100","v":-9.9560253999999997},{"$f":"'DF raw data'!N12/100","v":-5.8152759000000005},{"$f":"'DF raw data'!O12/100","v":-26.677177999999998},{"$f":"'DF raw data'!S12/100","v":-8.0703223000000008},{"$f":"'DF raw data'!T12/100","v":-14.546894999999999},{"$f":"'DF raw data'!X12/100","v":4.2738727000000001},{"$f":"'DF raw data'!Y12/100","v":-24.844502000000002}]
[-2,{"$f":"'DF raw data'!D13/100","v":-14.327230999999999},{"$f":"'DF raw data'!E13/100","v":-20.422606999999999},{"$f":"'DF raw data'!I13/100","v":-46.634799999999998},{"$f":"'DF raw data'!J13/100","v":-1.8219042999999999},{"$f":"'DF raw data'!N13/100","v":-10.641105},{"$f":"'DF raw data'!O13/100","v":-4.1154883},{"$f":"'DF raw data'!S13/100","v":-16.557718999999999},{"$f":"'DF raw data'!T13/100","v":-22.152568000000002},{"$f":"'DF raw data'!X13/100","v":-9.9852739999999987},{"$f":"'DF raw data'!Y13/100","v":-12.256982000000001},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":2.0196270999999997},{"$f":"'DF raw data'!E14/100","v":3.6250586},{"$f":"'DF raw data'!I14/100","v":4.4731180000000004},{"$f":"'DF raw data'!J14/100","v":-11.041523},{"$f":"'DF raw data'!N14/100","v":-7.0562652999999997},{"$f":"'DF raw data'!O14/100","v":-0.73659180000000002},{"$f":"'DF raw data'!S14/100","v":-7.2925860999999994},{"$f":"'DF raw data'!T14/100","v":-5.4406835999999998},{"$f":"'DF raw data'!X14/100","v":0.37238831},{"$f":"'DF raw data'!Y14/100","v":-8.1447949000000008},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-14.084836000000001},{"$f":"'DF raw data'!E15/100","v":-26.803114999999998},{"$f":"'DF raw data'!I15/100","v":-18.314609000000001},{"$f":"'DF raw data'!J15/100","v":-30.199531},{"$f":"'DF raw data'!N15/100","v":-21.407350999999998},{"$f":"'DF raw data'!O15/100","v":-27.064473},{"$f":"'DF raw data'!S15/100","v":8.0428180000000005},{"$f":"'DF raw data'!T15/100","v":-26.199424},{"$f":"'DF raw data'!X15/100","v":-5.2443091000000006},{"$f":"'DF raw data'!Y15/100","v":-26.419931999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-14.890251000000001},{"$f":"'DF raw data'!E16/100","v":-9.8212499999999991},{"$f":"'DF raw data'!I16/100","v":-13.808595},{"$f":"'DF raw data'!J16/100","v":-13.711074000000002},{"$f":"'DF raw data'!N16/100","v":-7.5692596000000005},{"$f":"'DF raw data'!O16/100","v":-31.267617000000001},{"$f":"'DF raw data'!S16/100","v":-15.599172000000001},{"$f":"'DF raw data'!T16/100","v":-18.932050999999998},{"$f":"'DF raw data'!X16/100","v":-12.894304},{"$f":"'DF raw data'!Y16/100","v":-21.759110999999997}]
[6,{"$f":"'DF raw data'!D17/100","v":11.429926},{"$f":"'DF raw data'!E17/100","v":7.9465918000000002},{"$f":"'DF raw data'!I17/100","v":3.0082938000000001},{"$f":"'DF raw data'!J17/100","v":6.0315918000000002},{"$f":"'DF raw data'!N17/100","v":-22.089169999999999},{"$f":"'DF raw data'!O17/100","v":5.0878027000000001},{"$f":"'DF raw data'!S17/100","v":1.9681237999999999},{"$f":"'DF raw data'!T17/100","v":11.875996000000001},{"$f":"'DF raw data'!X17/100","v":-2.6970348999999998},{"$f":"'DF raw data'!Y17/100","v":2.5368163999999997}]
[8,{"$f":"'DF raw data'!D18/100","v":47.348154000000001},{"$f":"'DF raw data'!E18/100","v":-17.243213000000001},{"$f":"'DF raw data'!I18/100","v":15.502901999999999},{"$f":"'DF raw data'!J18/100","v":-11.583486000000001},{"$f":"'DF raw data'!N18/100","v":21.450706},{"$f":"'DF raw data'!O18/100","v":-16.466875000000002},{"$f":"'DF raw data'!S18/100","v":26.772158000000001},{"$f":"'DF raw data'!T18/100","v":-16.161152000000001},{"$f":"'DF raw data'!X18/100","v":5.4375336000000001},{"$f":"'DF raw data'!Y18/100","v":-23.375565999999999}]
[10,{"$f":"'DF raw data'!D19/100","v":24.049469999999999},{"$f":"'DF raw data'!E19/100","v":29.384257999999999},{"$f":"'DF raw data'!I19/100","v":9.2027576},{"$f":"'DF raw data'!J19/100","v":7.8411425999999995},{"$f":"'DF raw data'!N19/100","v":3.0744860999999997},{"$f":"'DF raw data'!O19/100","v":3.6913770000000001},{"$f":"'DF raw data'!S19/100","v":14.637222},{"$f":"'DF raw data'!T19/100","v":10.837725000000001},{"$f":"'DF raw data'!X19/100","v":-23.164936999999998},{"$f":"'DF raw data'!Y19/100","v":-5.3504687999999998}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":802},null,{"$f":"'DF raw data'!G4","v":803},null,{"$f":"'DF raw data'!L4","v":804},null,{"$f":"'DF raw data'!Q4","v":805},null,{"$f":"'DF raw data'!V4","v":806}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-8.7974799666666659},{"$f":"AVERAGE(C11:C13)","v":-14.533554466666667},-20.158763666666669,-14.354319433333332,-13.034907099999998,-10.638851033333333,-5.2691623666666665,-17.930891866666666,-4.9523982633333334,-15.607236300000002]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16326, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16326, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16326, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16326, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16326, 100 mm gap"]
["Date:","18. sep 2017","08:05:45",null,null,"Date:","18. sep 2017","08:05:52",null,null,"Date:","18. sep 2017","08:05:57",null,null,"Date:","18. sep 2017","08:06:02",null,null,"Date:","18. sep 2017","08:06:31"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",802,null,null,null,"Indice:",803,null,null,null,"Indice:",804,null,null,null,"Indice:",805,null,null,null,"Indice:",806]
["Background Indice:",802,null,null,null,"Background Indice:",803,null,null,null,"Background Indice:",804,null,null,null,"Background Indice:",805,null,null,null,"Background Indice:",806]
["Temperature:",21.78,null,null,null,"Temperature:",21.65,null,null,null,"Temperature:",21.64,null,null,null,"Temperature:",21.57,null,null,null,"Temperature:",21.57]
["Encoder:",25018409,null,null,null,"Encoder:",24998365,null,null,null,"Encoder:",24978350,null,null,null,"Encoder:",24938332,null,null,null,"Encoder:",24658334]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-19.533190000000001,-4.9932499000000004,-2818.9099000000001,-2158.7422000000001,-1,-11.289149999999999,6.5816998,-1536.2693999999999,-977.16016000000002,-1,-11.37425,5.6534700000000004,-1863.7614000000001,-933.85155999999995,-1,-6.2302999000000003,2.8783099999999999,-107.19327,-1923.9794999999999,-1,-12.17393,-4.7520099,-1152.7904000000001,-2297.2754]
[-0.80000000999999998,-18.725071,-5.1249399000000002,-4045.9829,-1634.2227,-0.80000000999999998,-16.430820000000001,1.50665,-2093.4823999999999,-579.02148,-0.80000000999999998,-21.665929999999999,-3.1786699,-2868.3872000000001,-1358.5409999999999,-0.80000000999999998,-3.1273700999999998,-10.56099,-828.73424999999997,-2120.9268000000002,-0.80000000999999998,-1.8562799999999999,-8.8237400000000008,-671.22600999999997,-1652.9570000000001]
[-0.60000001999999997,-21.920850999999999,-9.8650999000000006,-3742.5736999999999,-2354.2559000000001,-0.60000001999999997,-13.58942,-1.9942299999999999,-2589.2786000000001,-1434.0898,-0.60000001999999997,-2.7846500999999999,-1.5103899999999999,44.625228999999997,-2183.5868999999998,-0.60000001999999997,-3.0009301000000002,-3.3293300000000001,-616.01288,-973.34375,-0.60000001999999997,-6.3056597999999999,-3.91751,-1308.3557000000001,-2333.9931999999999]
[-0.40000001000000002,-12.020110000000001,-11.173579999999999,-3061.6680000000001,-2552.2627000000002,-0.40000001000000002,-4.0741700999999999,-4.7596698000000002,-960.30151000000001,-995.60253999999998,-0.40000001000000002,-2.24593,-13.5321,-581.52759000000003,-2667.7177999999999,-0.40000001000000002,-0.45738000000000001,-6.8550601000000002,-807.03223000000003,-1454.6895,-0.40000001000000002,2.9849000000000001,-11.01285,427.38727,-2484.4502000000002]
[-0.2,-7.3610601000000004,-10.43186,-1432.7230999999999,-2042.2607,-0.2,-12.777010000000001,-8.9787196999999992,-4663.4799999999996,-182.19042999999999,-0.2,-3.1629200000000002,2.5627,-1064.1105,-411.54883000000001,-0.2,-6.5287499000000002,-6.0571799000000004,-1655.7719,-2215.2568000000001,-0.2,-1.3869100000000001,-9.3009194999999991,-998.52739999999994,-1225.6982]
[0,3.9044599999999998,-8.7065000999999995,201.96270999999999,362.50585999999998,0,4.5299500999999998,-7.8070798000000003,447.31180000000001,-1104.1523,0,3.7257699999999998,-2.2187399999999999,-705.62653,-73.659180000000006,0,5.3381499999999997,-0.69875001999999997,-729.25860999999998,-544.06835999999998,0,-1.43784,-4.8821301000000004,37.238830999999998,-814.47949000000006]
[0.2,-7.8025998999999997,-7.2815599000000004,-1408.4836,-2680.3114999999998,0.2,-3.3954401000000001,-5.4534798000000002,-1831.4609,-3019.9531000000002,0.2,-2.0976601000000001,-7.1256298999999999,-2140.7350999999999,-2706.4472999999998,0.2,-1.2577,-0.97701000999999998,804.28179999999998,-2619.9423999999999,0.2,-3.1903100000000002,-8.1550303,-524.43091000000004,-2641.9931999999999]
[0.40000001000000002,-5.4329599999999996,2.9762298999999999,-1489.0251000000001,-982.125,0.40000001000000002,-6.4324398,2.1565398999999998,-1380.8595,-1371.1074000000001,0.40000001000000002,-3.84653,-7.1699099999999998,-756.92596000000003,-3126.7617,0.40000001000000002,-4.3070997999999996,-1.1812400000000001,-1559.9172000000001,-1893.2050999999999,0.40000001000000002,-13.79401,-5.3373599,-1289.4304,-2175.9110999999998]
[0.60000001999999997,5.4994301999999999,-0.27605998999999998,1142.9926,794.65917999999999,0.60000001999999997,6.2117801000000004,-1.84988,300.82938000000001,603.15917999999999,0.60000001999999997,1.1713800000000001,-3.9751699,-2208.9169999999999,508.78026999999997,0.60000001999999997,2.30707,-4.0368298999999999,196.81237999999999,1187.5996,0.60000001999999997,-4.1105399,-4.8217502000000003,-269.70348999999999,253.68163999999999]
[0.80000000999999998,22.64612,-5.9265099000000001,4734.8154000000004,-1724.3213000000001,0.80000000999999998,9.1174698000000003,-5.9156199000000003,1550.2901999999999,-1158.3486,0.80000000999999998,10.20116,-8.8188896000000003,2145.0706,-1646.6875,0.80000000999999998,8.5224600000000006,-7.6382298000000004,2677.2157999999999,-1616.1152,0.80000000999999998,-0.34147999000000001,-16.028690000000001,543.75336000000004,-2337.5565999999999]
[1,19.556619999999999,12.892860000000001,2404.9470000000001,2938.4258,1,12.397259999999999,3.1667299,920.27575999999999,784.11425999999994,1,13.459390000000001,2.2527001000000002,307.44860999999997,369.1377,1,7.7291898999999997,-0.64303999999999994,1463.7221999999999,1083.7725,1,-2.89059,-7.4324002,-2316.4937,-535.04687999999999]