{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":649,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["HXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17939678},null,{"$f":"'DF raw data'!G7","v":17919621},null,{"$f":"'DF raw data'!L7","v":17899603},null,{"$f":"'DF raw data'!Q7","v":17859589},null,{"$f":"'DF raw data'!V7","v":17579610},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-4.8853133333333334},{"$f":"C21","v":4.4007066999999997},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":4.1362299333333326},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":4.1947199666666668}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-3.5992833000000002},{"$f":"E21","v":8.2830266666666663},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":4.1362299333333326},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":7.5900232999999995}]
[-10,{"$f":"'DF raw data'!B9","v":-10.18117},{"$f":"'DF raw data'!C9","v":21.840250000000001},{"$f":"'DF raw data'!G9","v":-14.68524},{"$f":"'DF raw data'!H9","v":19.762329000000001},{"$f":"'DF raw data'!L9","v":-0.12296},{"$f":"'DF raw data'!M9","v":13.81663},{"$f":"'DF raw data'!Q9","v":-7.6605401000000004},{"$f":"'DF raw data'!R9","v":13.02844},{"$f":"'DF raw data'!V9","v":-5.8134598999999998},{"$f":"'DF raw data'!W9","v":9.2018603999999993},null,{"$f":"F4","v":20},{"$f":"F21","v":-7.7355132333333332},{"$f":"G21","v":4.8248932333333334}]
[-8,{"$f":"'DF raw data'!B10","v":-3.1907301000000001},{"$f":"'DF raw data'!C10","v":14.876099999999999},{"$f":"'DF raw data'!G10","v":-3.6362700000000001},{"$f":"'DF raw data'!H10","v":16.747510999999999},{"$f":"'DF raw data'!L10","v":-4.8966699},{"$f":"'DF raw data'!M10","v":18.202739999999999},{"$f":"'DF raw data'!Q10","v":-2.1602098999999999},{"$f":"'DF raw data'!R10","v":11.63129},{"$f":"'DF raw data'!V10","v":-0.81142998},{"$f":"'DF raw data'!W10","v":8.5706596000000008},null,{"$f":"H4","v":30},{"$f":"H21","v":-5.1379301333333336},{"$f":"I21","v":4.0883066999999995},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-4.1653298999999997},{"$f":"'DF raw data'!C11","v":15.36242},{"$f":"'DF raw data'!G11","v":-3.3341099999999999},{"$f":"'DF raw data'!H11","v":17.52467},{"$f":"'DF raw data'!L11","v":-9.4863701000000002},{"$f":"'DF raw data'!M11","v":19.773371000000001},{"$f":"'DF raw data'!Q11","v":-4.0014200000000004},{"$f":"'DF raw data'!R11","v":10.487170000000001},{"$f":"'DF raw data'!V11","v":-1.4019200000000001},{"$f":"'DF raw data'!W11","v":1.25587},null,{"$f":"J4","v":100},{"$f":"J21","v":-3.7930666666666659},{"$f":"K21","v":0.69300336666666651},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-7.1566600999999999},{"$f":"'DF raw data'!C12","v":14.70187},{"$f":"'DF raw data'!G12","v":-9.1537103999999996},{"$f":"'DF raw data'!H12","v":18.6404},{"$f":"'DF raw data'!L12","v":-5.6403398999999999},{"$f":"'DF raw data'!M12","v":10.657780000000001},{"$f":"'DF raw data'!Q12","v":-12.374510000000001},{"$f":"'DF raw data'!R12","v":14.507680000000001},{"$f":"'DF raw data'!V12","v":-2.6352500999999999},{"$f":"'DF raw data'!W12","v":5.9015899000000003},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-7.3569002000000001},{"$f":"'DF raw data'!C13","v":3.1037499999999998},{"$f":"'DF raw data'!G13","v":-5.8696399000000001},{"$f":"'DF raw data'!H13","v":5.2433399999999999},{"$f":"'DF raw data'!L13","v":-9.6208896999999993},{"$f":"'DF raw data'!M13","v":2.1641699999999999},{"$f":"'DF raw data'!Q13","v":-4.7800202000000001},{"$f":"'DF raw data'!R13","v":4.1339502000000001},{"$f":"'DF raw data'!V13","v":-4.1063599999999996},{"$f":"'DF raw data'!W13","v":-6.1054200999999999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":-3.9079598999999998},{"$f":"'DF raw data'!C14","v":5.8103800000000003},{"$f":"'DF raw data'!G14","v":-3.7340599999999999},{"$f":"'DF raw data'!H14","v":14.396100000000001},{"$f":"'DF raw data'!L14","v":-7.2706099000000002},{"$f":"'DF raw data'!M14","v":7.1684298999999996},{"$f":"'DF raw data'!Q14","v":-7.5724301000000001},{"$f":"'DF raw data'!R14","v":7.1213799},{"$f":"'DF raw data'!V14","v":-5.0867199999999997},{"$f":"'DF raw data'!W14","v":5.8635301999999996},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-3.3910798999999998},{"$f":"'DF raw data'!C15","v":4.2879901},{"$f":"'DF raw data'!G15","v":-1.19415},{"$f":"'DF raw data'!H15","v":5.2096400000000003},{"$f":"'DF raw data'!L15","v":-6.3150401},{"$f":"'DF raw data'!M15","v":5.1420798000000003},{"$f":"'DF raw data'!Q15","v":-3.0613400999999998},{"$f":"'DF raw data'!R15","v":1.00959},{"$f":"'DF raw data'!V15","v":-2.1861199999999998},{"$f":"'DF raw data'!W15","v":2.3209},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-6.8969997999999997},{"$f":"'DF raw data'!C16","v":-0.34983998999999999},{"$f":"'DF raw data'!G16","v":0.18415999},{"$f":"'DF raw data'!H16","v":11.45584},{"$f":"'DF raw data'!L16","v":1.9878},{"$f":"'DF raw data'!M16","v":0.89085000999999997},{"$f":"'DF raw data'!Q16","v":-11.458729999999999},{"$f":"'DF raw data'!R16","v":2.3814199},{"$f":"'DF raw data'!V16","v":-5.9247798999999999},{"$f":"'DF raw data'!W16","v":1.09995}]
[6,{"$f":"'DF raw data'!B17","v":-6.3484100999999997},{"$f":"'DF raw data'!C17","v":-4.2293301000000003},{"$f":"'DF raw data'!G17","v":-6.1003499000000003},{"$f":"'DF raw data'!H17","v":-2.0729001},{"$f":"'DF raw data'!L17","v":-6.0591201999999997},{"$f":"'DF raw data'!M17","v":-3.79813},{"$f":"'DF raw data'!Q17","v":0.43332999999999999},{"$f":"'DF raw data'!R17","v":4.7149000000000001},{"$f":"'DF raw data'!V17","v":-16.290379999999999},{"$f":"'DF raw data'!W17","v":8.2811202999999995}]
[8,{"$f":"'DF raw data'!B18","v":-2.1036201000000001},{"$f":"'DF raw data'!C18","v":-6.1202202000000003},{"$f":"'DF raw data'!G18","v":-13.21744},{"$f":"'DF raw data'!H18","v":9.2884101999999995},{"$f":"'DF raw data'!L18","v":-5.3147501999999998},{"$f":"'DF raw data'!M18","v":-0.34786999000000002},{"$f":"'DF raw data'!Q18","v":-9.4898395999999998},{"$f":"'DF raw data'!R18","v":1.0096499999999999},{"$f":"'DF raw data'!V18","v":-16.501418999999999},{"$f":"'DF raw data'!W18","v":14.14306}]
[10,{"$f":"'DF raw data'!B19","v":-17.435869},{"$f":"'DF raw data'!C19","v":-8.7192898000000003},{"$f":"'DF raw data'!G19","v":-20.34572},{"$f":"'DF raw data'!H19","v":-9.5851097000000003},{"$f":"'DF raw data'!L19","v":-14.423679999999999},{"$f":"'DF raw data'!M19","v":2.4797201000000002},{"$f":"'DF raw data'!Q19","v":-8.3556498999999995},{"$f":"'DF raw data'!R19","v":0.25519999999999998},{"$f":"'DF raw data'!V19","v":-4.1394501000000004},{"$f":"'DF raw data'!W19","v":7.1519399000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":919},null,{"$f":"'DF raw data'!G4","v":920},null,{"$f":"'DF raw data'!L4","v":921},null,{"$f":"'DF raw data'!Q4","v":922},null,{"$f":"'DF raw data'!V4","v":923}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-4.8853133333333334},{"$f":"AVERAGE(C11:C13)","v":4.4007066999999997},-3.5992833000000002,8.2830266666666663,-7.7355132333333332,4.8248932333333334,-5.1379301333333336,4.0883066999999995,-3.7930666666666659,0.69300336666666651]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["HXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":17939678},null,{"$f":"'DF raw data'!G7","v":17919621},null,{"$f":"'DF raw data'!L7","v":17899603},null,{"$f":"'DF raw data'!Q7","v":17859589},null,{"$f":"'DF raw data'!V7","v":17579610},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-12.747837666666667},{"$f":"C21","v":6.9910872666666668},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":7.9235126666666655},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":16.135488416666664}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-16.722161666666668},{"$f":"E21","v":18.639469509999998},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":10.265181733333334},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":16.677369909999996}]
[-10,{"$f":"'DF raw data'!D9/100","v":-19.715886000000001},{"$f":"'DF raw data'!E9/100","v":41.962568000000005},{"$f":"'DF raw data'!I9/100","v":-25.237978999999999},{"$f":"'DF raw data'!J9/100","v":28.681846},{"$f":"'DF raw data'!N9/100","v":-11.562505000000002},{"$f":"'DF raw data'!O9/100","v":20.069482000000001},{"$f":"'DF raw data'!S9/100","v":-16.815275999999997},{"$f":"'DF raw data'!T9/100","v":12.480469000000001},{"$f":"'DF raw data'!X9/100","v":-4.9960759999999995},{"$f":"'DF raw data'!Y9/100","v":11.769520999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":-20.671350333333333},{"$f":"G21","v":5.923346266666667}]
[-8,{"$f":"'DF raw data'!D10/100","v":-5.5576318000000002},{"$f":"'DF raw data'!E10/100","v":18.480907999999999},{"$f":"'DF raw data'!I10/100","v":-20.069230000000001},{"$f":"'DF raw data'!J10/100","v":29.410654000000001},{"$f":"'DF raw data'!N10/100","v":-21.916487000000004},{"$f":"'DF raw data'!O10/100","v":26.670673999999998},{"$f":"'DF raw data'!S10/100","v":-10.243622},{"$f":"'DF raw data'!T10/100","v":13.237832000000001},{"$f":"'DF raw data'!X10/100","v":-8.6393787},{"$f":"'DF raw data'!Y10/100","v":16.550615000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-17.913154599999999},{"$f":"I21","v":2.5039810933333335}]
[-6,{"$f":"'DF raw data'!D11/100","v":-10.254994999999999},{"$f":"'DF raw data'!E11/100","v":22.689589999999999},{"$f":"'DF raw data'!I11/100","v":-14.615380999999999},{"$f":"'DF raw data'!J11/100","v":15.860723},{"$f":"'DF raw data'!N11/100","v":-15.001206},{"$f":"'DF raw data'!O11/100","v":31.035868999999998},{"$f":"'DF raw data'!S11/100","v":-8.2217846999999988},{"$f":"'DF raw data'!T11/100","v":16.288798999999997},{"$f":"'DF raw data'!X11/100","v":7.7075702000000001},{"$f":"'DF raw data'!Y11/100","v":2.9858496000000003},null,{"$f":"J4","v":100},{"$f":"J21","v":-10.406168599999999},{"$f":"K21","v":1.9620996000000002}]
[-4,{"$f":"'DF raw data'!D12/100","v":-10.149319999999999},{"$f":"'DF raw data'!E12/100","v":27.193798999999999},{"$f":"'DF raw data'!I12/100","v":-48.868003000000002},{"$f":"'DF raw data'!J12/100","v":36.588154000000003},{"$f":"'DF raw data'!N12/100","v":-20.315293},{"$f":"'DF raw data'!O12/100","v":18.772763999999999},{"$f":"'DF raw data'!S12/100","v":-35.234141000000001},{"$f":"'DF raw data'!T12/100","v":25.959229000000001},{"$f":"'DF raw data'!X12/100","v":-16.502510999999998},{"$f":"'DF raw data'!Y12/100","v":11.975048999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":-11.601331999999999},{"$f":"'DF raw data'!E13/100","v":4.6562597999999999},{"$f":"'DF raw data'!I13/100","v":-26.322217000000002},{"$f":"'DF raw data'!J13/100","v":0.74418952999999999},{"$f":"'DF raw data'!N13/100","v":-21.954459999999997},{"$f":"'DF raw data'!O13/100","v":5.8296289000000003},{"$f":"'DF raw data'!S13/100","v":-21.770893999999998},{"$f":"'DF raw data'!T13/100","v":-0.84948241999999996},{"$f":"'DF raw data'!X13/100","v":-11.942621999999998},{"$f":"'DF raw data'!Y13/100","v":-5.572959},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-14.419592},{"$f":"'DF raw data'!E14/100","v":15.316376999999999},{"$f":"'DF raw data'!I14/100","v":-34.877488},{"$f":"'DF raw data'!J14/100","v":29.350351999999997},{"$f":"'DF raw data'!N14/100","v":-25.751493999999997},{"$f":"'DF raw data'!O14/100","v":17.606699000000003},{"$f":"'DF raw data'!S14/100","v":-25.355999000000001},{"$f":"'DF raw data'!T14/100","v":12.013496},{"$f":"'DF raw data'!X14/100","v":-9.0112047999999998},{"$f":"'DF raw data'!Y14/100","v":9.5593164000000002},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-12.222589000000001},{"$f":"'DF raw data'!E15/100","v":1.0006250000000001},{"$f":"'DF raw data'!I15/100","v":11.033219999999998},{"$f":"'DF raw data'!J15/100","v":25.823867},{"$f":"'DF raw data'!N15/100","v":-14.308097},{"$f":"'DF raw data'!O15/100","v":-5.6662891000000002},{"$f":"'DF raw data'!S15/100","v":-6.6125707999999994},{"$f":"'DF raw data'!T15/100","v":-3.6520702999999997},{"$f":"'DF raw data'!X15/100","v":-10.264679000000001},{"$f":"'DF raw data'!Y15/100","v":1.8999413999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-11.749591000000001},{"$f":"'DF raw data'!E16/100","v":-0.78911133000000011},{"$f":"'DF raw data'!I16/100","v":-16.760653999999999},{"$f":"'DF raw data'!J16/100","v":24.181903999999999},{"$f":"'DF raw data'!N16/100","v":1.2229657},{"$f":"'DF raw data'!O16/100","v":8.7341797000000003},{"$f":"'DF raw data'!S16/100","v":-34.726891999999999},{"$f":"'DF raw data'!T16/100","v":19.993164},{"$f":"'DF raw data'!X16/100","v":-10.11631},{"$f":"'DF raw data'!Y16/100","v":15.747998000000001}]
[6,{"$f":"'DF raw data'!D17/100","v":-16.989471000000002},{"$f":"'DF raw data'!E17/100","v":-7.0107128999999997},{"$f":"'DF raw data'!I17/100","v":-19.805486999999999},{"$f":"'DF raw data'!J17/100","v":-4.3334960999999996},{"$f":"'DF raw data'!N17/100","v":-11.820314000000002},{"$f":"'DF raw data'!O17/100","v":-12.364746},{"$f":"'DF raw data'!S17/100","v":-7.2559472999999999},{"$f":"'DF raw data'!T17/100","v":-7.7438574000000004},{"$f":"'DF raw data'!X17/100","v":-26.801327999999998},{"$f":"'DF raw data'!Y17/100","v":14.388936000000001}]
[8,{"$f":"'DF raw data'!D18/100","v":-13.495433},{"$f":"'DF raw data'!E18/100","v":-14.296445},{"$f":"'DF raw data'!I18/100","v":-20.191434000000001},{"$f":"'DF raw data'!J18/100","v":19.715713000000001},{"$f":"'DF raw data'!N18/100","v":-22.273235},{"$f":"'DF raw data'!O18/100","v":5.5629785000000007},{"$f":"'DF raw data'!S18/100","v":-31.198049000000001},{"$f":"'DF raw data'!T18/100","v":8.2171973000000005},{"$f":"'DF raw data'!X18/100","v":-42.275048999999996},{"$f":"'DF raw data'!Y18/100","v":25.655839999999998}]
[10,{"$f":"'DF raw data'!D19/100","v":-33.572437000000001},{"$f":"'DF raw data'!E19/100","v":-10.047666},{"$f":"'DF raw data'!I19/100","v":-47.634467999999998},{"$f":"'DF raw data'!J19/100","v":-29.128975000000001},{"$f":"'DF raw data'!N19/100","v":-26.027698000000001},{"$f":"'DF raw data'!O19/100","v":13.429824000000002},{"$f":"'DF raw data'!S19/100","v":-34.544145999999998},{"$f":"'DF raw data'!T19/100","v":-4.6129883000000005},{"$f":"'DF raw data'!X19/100","v":-14.120676000000001},{"$f":"'DF raw data'!Y19/100","v":10.527461000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":919},null,{"$f":"'DF raw data'!G4","v":920},null,{"$f":"'DF raw data'!L4","v":921},null,{"$f":"'DF raw data'!Q4","v":922},null,{"$f":"'DF raw data'!V4","v":923}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-12.747837666666667},{"$f":"AVERAGE(C11:C13)","v":6.9910872666666668},-16.722161666666668,18.639469509999998,-20.671350333333333,5.923346266666667,-17.913154599999999,2.5039810933333335,-10.406168599999999,1.9620996000000002]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16333, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16333, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16333, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16333, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16333, 100 mm gap"]
["Date:","11. okt 2017","07:49:12",null,null,"Date:","11. okt 2017","07:48:52",null,null,"Date:","11. okt 2017","07:49:00",null,null,"Date:","11. okt 2017","07:49:06",null,null,"Date:","11. okt 2017","07:49:13"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",919,null,null,null,"Indice:",920,null,null,null,"Indice:",921,null,null,null,"Indice:",922,null,null,null,"Indice:",923]
["Background Indice:",919,null,null,null,"Background Indice:",920,null,null,null,"Background Indice:",921,null,null,null,"Background Indice:",922,null,null,null,"Background Indice:",923]
["Temperature:",21.68,null,null,null,"Temperature:",21.68,null,null,null,"Temperature:",21.65,null,null,null,"Temperature:",21.67,null,null,null,"Temperature:",21.63]
["Encoder:",17939678,null,null,null,"Encoder:",17919621,null,null,null,"Encoder:",17899603,null,null,null,"Encoder:",17859589,null,null,null,"Encoder:",17579610]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-10.18117,21.840250000000001,-1971.5886,4196.2568000000001,-1,-14.68524,19.762329000000001,-2523.7979,2868.1846,-1,-0.12296,13.81663,-1156.2505000000001,2006.9482,-1,-7.6605401000000004,13.02844,-1681.5275999999999,1248.0469000000001,-1,-5.8134598999999998,9.2018603999999993,-499.60759999999999,1176.9521]
[-0.80000000999999998,-3.1907301000000001,14.876099999999999,-555.76318000000003,1848.0907999999999,-0.80000000999999998,-3.6362700000000001,16.747510999999999,-2006.923,2941.0654,-0.80000000999999998,-4.8966699,18.202739999999999,-2191.6487000000002,2667.0673999999999,-0.80000000999999998,-2.1602098999999999,11.63129,-1024.3622,1323.7832000000001,-0.80000000999999998,-0.81142998,8.5706596000000008,-863.93786999999998,1655.0615]
[-0.60000001999999997,-4.1653298999999997,15.36242,-1025.4994999999999,2268.9589999999998,-0.60000001999999997,-3.3341099999999999,17.52467,-1461.5381,1586.0723,-0.60000001999999997,-9.4863701000000002,19.773371000000001,-1500.1206,3103.5868999999998,-0.60000001999999997,-4.0014200000000004,10.487170000000001,-822.17846999999995,1628.8798999999999,-0.60000001999999997,-1.4019200000000001,1.25587,770.75702000000001,298.58496000000002]
[-0.40000001000000002,-7.1566600999999999,14.70187,-1014.932,2719.3798999999999,-0.40000001000000002,-9.1537103999999996,18.6404,-4886.8002999999999,3658.8154,-0.40000001000000002,-5.6403398999999999,10.657780000000001,-2031.5292999999999,1877.2764,-0.40000001000000002,-12.374510000000001,14.507680000000001,-3523.4141,2595.9229,-0.40000001000000002,-2.6352500999999999,5.9015899000000003,-1650.2511,1197.5048999999999]
[-0.2,-7.3569002000000001,3.1037499999999998,-1160.1332,465.62598000000003,-0.2,-5.8696399000000001,5.2433399999999999,-2632.2217000000001,74.418953000000002,-0.2,-9.6208896999999993,2.1641699999999999,-2195.4459999999999,582.96289000000002,-0.2,-4.7800202000000001,4.1339502000000001,-2177.0893999999998,-84.948241999999993,-0.2,-4.1063599999999996,-6.1054200999999999,-1194.2621999999999,-557.29589999999996]
[0,-3.9079598999999998,5.8103800000000003,-1441.9592,1531.6377,0,-3.7340599999999999,14.396100000000001,-3487.7487999999998,2935.0351999999998,0,-7.2706099000000002,7.1684298999999996,-2575.1493999999998,1760.6699000000001,0,-7.5724301000000001,7.1213799,-2535.5999000000002,1201.3496,0,-5.0867199999999997,5.8635301999999996,-901.12048000000004,955.93164000000002]
[0.2,-3.3910798999999998,4.2879901,-1222.2589,100.0625,0.2,-1.19415,5.2096400000000003,1103.3219999999999,2582.3867,0.2,-6.3150401,5.1420798000000003,-1430.8097,-566.62891000000002,0.2,-3.0613400999999998,1.00959,-661.25707999999997,-365.20702999999997,0.2,-2.1861199999999998,2.3209,-1026.4679000000001,189.99413999999999]
[0.40000001000000002,-6.8969997999999997,-0.34983998999999999,-1174.9591,-78.911133000000007,0.40000001000000002,0.18415999,11.45584,-1676.0654,2418.1904,0.40000001000000002,1.9878,0.89085000999999997,122.29657,873.41796999999997,0.40000001000000002,-11.458729999999999,2.3814199,-3472.6891999999998,1999.3163999999999,0.40000001000000002,-5.9247798999999999,1.09995,-1011.631,1574.7998]
[0.60000001999999997,-6.3484100999999997,-4.2293301000000003,-1698.9471000000001,-701.07128999999998,0.60000001999999997,-6.1003499000000003,-2.0729001,-1980.5487000000001,-433.34960999999998,0.60000001999999997,-6.0591201999999997,-3.79813,-1182.0314000000001,-1236.4746,0.60000001999999997,0.43332999999999999,4.7149000000000001,-725.59473000000003,-774.38574000000006,0.60000001999999997,-16.290379999999999,8.2811202999999995,-2680.1327999999999,1438.8936000000001]
[0.80000000999999998,-2.1036201000000001,-6.1202202000000003,-1349.5433,-1429.6445000000001,0.80000000999999998,-13.21744,9.2884101999999995,-2019.1433999999999,1971.5713000000001,0.80000000999999998,-5.3147501999999998,-0.34786999000000002,-2227.3235,556.29785000000004,0.80000000999999998,-9.4898395999999998,1.0096499999999999,-3119.8049000000001,821.71973000000003,0.80000000999999998,-16.501418999999999,14.14306,-4227.5048999999999,2565.5839999999998]
[1,-17.435869,-8.7192898000000003,-3357.2437,-1004.7666,1,-20.34572,-9.5851097000000003,-4763.4467999999997,-2912.8975,1,-14.423679999999999,2.4797201000000002,-2602.7698,1342.9824000000001,1,-8.3556498999999995,0.25519999999999998,-3454.4146000000001,-461.29883000000001,1,-4.1394501000000004,7.1519399000000003,-1412.0676000000001,1052.7461000000001]