{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":31,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386573},null,{"$f":"'DF raw data'!G7","v":487201},null,{"$f":"'DF raw data'!L7","v":587536},null,{"$f":"'DF raw data'!Q7","v":787807},null,{"$f":"'DF raw data'!V7","v":2187952},null,null,{"$f":"B4","v":10},{"$f":"B21","v":1.0724566933333335},{"$f":"C21","v":-3.6177633},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":3.6110733266666664},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":4.0539166029999993}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-1.2303699666666665},{"$f":"E21","v":-5.3330966666666662},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":7.956379993333333},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":4.8962567033333322}]
[-10,{"$f":"'DF raw data'!B9","v":5.1468600999999996},{"$f":"'DF raw data'!C9","v":-6.7997499000000001},{"$f":"'DF raw data'!G9","v":7.5170002999999999E-2},{"$f":"'DF raw data'!H9","v":-11.024990000000001},{"$f":"'DF raw data'!L9","v":7.6790700000000003},{"$f":"'DF raw data'!M9","v":-20.422661000000002},{"$f":"'DF raw data'!Q9","v":-9.6972103000000001},{"$f":"'DF raw data'!R9","v":-7.8865800000000004},{"$f":"'DF raw data'!V9","v":-0.86945998999999996},{"$f":"'DF raw data'!W9","v":-6.7975501999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":-1.4155901},{"$f":"G21","v":-2.3524866633333334}]
[-8,{"$f":"'DF raw data'!B10","v":0.66267997000000001},{"$f":"'DF raw data'!C10","v":-4.6005402000000002},{"$f":"'DF raw data'!G10","v":3.8782499000000001},{"$f":"'DF raw data'!H10","v":-10.466760000000001},{"$f":"'DF raw data'!L10","v":1.28687},{"$f":"'DF raw data'!M10","v":-15.7202},{"$f":"'DF raw data'!Q10","v":2.7096300000000002},{"$f":"'DF raw data'!R10","v":-4.8853698000000003},{"$f":"'DF raw data'!V10","v":-2.3414999999999999},{"$f":"'DF raw data'!W10","v":-8.1977100000000007},null,{"$f":"H4","v":30},{"$f":"H21","v":-2.5386166333333331},{"$f":"I21","v":-6.4064032663333323},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-2.5943000000000001},{"$f":"'DF raw data'!C11","v":-4.0891599999999997},{"$f":"'DF raw data'!G11","v":-1.19045},{"$f":"'DF raw data'!H11","v":-3.3604701000000001},{"$f":"'DF raw data'!L11","v":-4.2116097999999997},{"$f":"'DF raw data'!M11","v":-3.5898199000000002},{"$f":"'DF raw data'!Q11","v":-1.4494100000000001},{"$f":"'DF raw data'!R11","v":2.4734001000000001},{"$f":"'DF raw data'!V11","v":-2.7331300000000001},{"$f":"'DF raw data'!W11","v":1.8844799999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":-6.8839232999999993},{"$f":"K21","v":-7.2487433666666661},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":1.05077},{"$f":"'DF raw data'!C12","v":-3.3866700999999999},{"$f":"'DF raw data'!G12","v":-2.6096499},{"$f":"'DF raw data'!H12","v":-9.7984895999999999},{"$f":"'DF raw data'!L12","v":-3.8149099},{"$f":"'DF raw data'!M12","v":-1.1129100000000001},{"$f":"'DF raw data'!Q12","v":-2.4322701000000002},{"$f":"'DF raw data'!R12","v":1.7782100000000001},{"$f":"'DF raw data'!V12","v":-5.6578102000000001},{"$f":"'DF raw data'!W12","v":-1.54958},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":-0.11337},{"$f":"'DF raw data'!C13","v":-5.8978000000000002},{"$f":"'DF raw data'!G13","v":-1.76197},{"$f":"'DF raw data'!H13","v":-6.3244300000000004},{"$f":"'DF raw data'!L13","v":2.0547198999999998},{"$f":"'DF raw data'!M13","v":-3.33846},{"$f":"'DF raw data'!Q13","v":1.6110199999999999},{"$f":"'DF raw data'!R13","v":-5.3479999E-2},{"$f":"'DF raw data'!V13","v":-2.8109099999999998},{"$f":"'DF raw data'!W13","v":-4.6232901000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":4.0902801000000002},{"$f":"'DF raw data'!C14","v":-1.06423},{"$f":"'DF raw data'!G14","v":3.8299601000000001},{"$f":"'DF raw data'!H14","v":-3.5168900000000001},{"$f":"'DF raw data'!L14","v":-1.9351499999999999},{"$f":"'DF raw data'!M14","v":0.48884000999999999},{"$f":"'DF raw data'!Q14","v":-1.87862},{"$f":"'DF raw data'!R14","v":-5.9078397999999996},{"$f":"'DF raw data'!V14","v":-5.9774498999999999},{"$f":"'DF raw data'!W14","v":-9.3016299999999994},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-0.75954001999999998},{"$f":"'DF raw data'!C15","v":-3.8912599000000001},{"$f":"'DF raw data'!G15","v":-5.7591000000000001},{"$f":"'DF raw data'!H15","v":-6.1579699999999997},{"$f":"'DF raw data'!L15","v":-4.3663401999999998},{"$f":"'DF raw data'!M15","v":-4.20784},{"$f":"'DF raw data'!Q15","v":-7.3482498999999999},{"$f":"'DF raw data'!R15","v":-13.25789},{"$f":"'DF raw data'!V15","v":-11.86341},{"$f":"'DF raw data'!W15","v":-7.8213100000000004},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":5.3397198000000001},{"$f":"'DF raw data'!C16","v":11.23719},{"$f":"'DF raw data'!G16","v":2.3629999000000002},{"$f":"'DF raw data'!H16","v":1.6575899999999999},{"$f":"'DF raw data'!L16","v":-0.77370000000000005},{"$f":"'DF raw data'!M16","v":5.5411099999999998},{"$f":"'DF raw data'!Q16","v":4.2098398000000001},{"$f":"'DF raw data'!R16","v":6.3607798000000004},{"$f":"'DF raw data'!V16","v":-0.32512000000000002},{"$f":"'DF raw data'!W16","v":-0.71147000999999999}]
[6,{"$f":"'DF raw data'!B17","v":-6.0624900000000004},{"$f":"'DF raw data'!C17","v":5.8829298000000003},{"$f":"'DF raw data'!G17","v":-0.26941000999999998},{"$f":"'DF raw data'!H17","v":1.7905},{"$f":"'DF raw data'!L17","v":0.54984999000000001},{"$f":"'DF raw data'!M17","v":-1.3202499999999999},{"$f":"'DF raw data'!Q17","v":-7.8935699000000001},{"$f":"'DF raw data'!R17","v":2.4807301000000002},{"$f":"'DF raw data'!V17","v":-4.4871100999999998},{"$f":"'DF raw data'!W17","v":-2.09849}]
[8,{"$f":"'DF raw data'!B18","v":-2.3843200000000002},{"$f":"'DF raw data'!C18","v":6.5428499999999996},{"$f":"'DF raw data'!G18","v":-1.0101500000000001},{"$f":"'DF raw data'!H18","v":6.6595402000000004},{"$f":"'DF raw data'!L18","v":-3.5115699999999999},{"$f":"'DF raw data'!M18","v":4.0817098999999999},{"$f":"'DF raw data'!Q18","v":-1.9806299999999999},{"$f":"'DF raw data'!R18","v":2.2995999},{"$f":"'DF raw data'!V18","v":-10.90578},{"$f":"'DF raw data'!W18","v":-6.0621400000000003}]
[10,{"$f":"'DF raw data'!B19","v":-3.0720000000000001},{"$f":"'DF raw data'!C19","v":18.849139999999998},{"$f":"'DF raw data'!G19","v":-1.83717},{"$f":"'DF raw data'!H19","v":13.95645},{"$f":"'DF raw data'!L19","v":-8.9072703999999998},{"$f":"'DF raw data'!M19","v":11.881169999999999},{"$f":"'DF raw data'!Q19","v":-5.2515897999999996},{"$f":"'DF raw data'!R19","v":13.952109999999999},{"$f":"'DF raw data'!V19","v":-2.8818901000000001},{"$f":"'DF raw data'!W19","v":1.1548001000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":232},null,{"$f":"'DF raw data'!G4","v":233},null,{"$f":"'DF raw data'!L4","v":229},null,{"$f":"'DF raw data'!Q4","v":230},null,{"$f":"'DF raw data'!V4","v":231}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":1.0724566933333335},{"$f":"AVERAGE(C11:C13)","v":-3.6177633},-1.2303699666666665,-5.3330966666666662,-1.4155901,-2.3524866633333334,-2.5386166333333331,-6.4064032663333323,-6.8839232999999993,-7.2487433666666661]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386573},null,{"$f":"'DF raw data'!G7","v":487201},null,{"$f":"'DF raw data'!L7","v":587536},null,{"$f":"'DF raw data'!Q7","v":787807},null,{"$f":"'DF raw data'!V7","v":2187952},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-6.7697853666666674},{"$f":"C21","v":11.551087199999998},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":6.5760115433333333},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":8.3081445333333317}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-1.6651326666666666},{"$f":"E21","v":4.8622721333333336},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":10.092296500000002},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":9.606868409999997}]
[-10,{"$f":"'DF raw data'!D9/100","v":-8.5479132},{"$f":"'DF raw data'!E9/100","v":-8.0870508000000001},{"$f":"'DF raw data'!I9/100","v":-35.264868},{"$f":"'DF raw data'!J9/100","v":-3.8337108999999998},{"$f":"'DF raw data'!N9/100","v":7.0318250000000004},{"$f":"'DF raw data'!O9/100","v":-26.401895},{"$f":"'DF raw data'!S9/100","v":-24.868662},{"$f":"'DF raw data'!T9/100","v":-0.46530270000000001},{"$f":"'DF raw data'!X9/100","v":-11.897815000000001},{"$f":"'DF raw data'!Y9/100","v":2.2774315999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":-1.8165128666666668},{"$f":"G21","v":9.589557366666666}]
[-8,{"$f":"'DF raw data'!D10/100","v":-21.360214999999997},{"$f":"'DF raw data'!E10/100","v":-13.801602000000001},{"$f":"'DF raw data'!I10/100","v":3.3291577000000001},{"$f":"'DF raw data'!J10/100","v":1.2066602},{"$f":"'DF raw data'!N10/100","v":7.2409942999999997},{"$f":"'DF raw data'!O10/100","v":-4.2184473000000002},{"$f":"'DF raw data'!S10/100","v":-0.86630981000000007},{"$f":"'DF raw data'!T10/100","v":-1.4495898},{"$f":"'DF raw data'!X10/100","v":1.8741759999999998},{"$f":"'DF raw data'!Y10/100","v":-11.895352000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":-8.2411442099999999},{"$f":"I21","v":3.2429426666666665}]
[-6,{"$f":"'DF raw data'!D11/100","v":-7.1197790999999997},{"$f":"'DF raw data'!E11/100","v":2.8924512},{"$f":"'DF raw data'!I11/100","v":-3.3811309999999999},{"$f":"'DF raw data'!J11/100","v":10.945752000000001},{"$f":"'DF raw data'!N11/100","v":-14.190183999999999},{"$f":"'DF raw data'!O11/100","v":5.2477050999999992},{"$f":"'DF raw data'!S11/100","v":-9.0246887000000005},{"$f":"'DF raw data'!T11/100","v":11.604844},{"$f":"'DF raw data'!X11/100","v":-12.687443},{"$f":"'DF raw data'!Y11/100","v":4.0627734000000002},null,{"$f":"J4","v":100},{"$f":"J21","v":-11.757429166666668},{"$f":"K21","v":1.9442187900000005}]
[-4,{"$f":"'DF raw data'!D12/100","v":-2.2128899E-2},{"$f":"'DF raw data'!E12/100","v":9.9634473000000003},{"$f":"'DF raw data'!I12/100","v":3.2122046000000002},{"$f":"'DF raw data'!J12/100","v":-14.675673999999999},{"$f":"'DF raw data'!N12/100","v":-3.5485937999999999},{"$f":"'DF raw data'!O12/100","v":8.686328099999999},{"$f":"'DF raw data'!S12/100","v":-4.1320519999999998},{"$f":"'DF raw data'!T12/100","v":19.286846000000001},{"$f":"'DF raw data'!X12/100","v":-16.505399999999998},{"$f":"'DF raw data'!Y12/100","v":2.5052734000000001}]
[-2,{"$f":"'DF raw data'!D13/100","v":3.8180786000000002},{"$f":"'DF raw data'!E13/100","v":12.397440999999999},{"$f":"'DF raw data'!I13/100","v":1.7224994},{"$f":"'DF raw data'!J13/100","v":9.5751270000000002},{"$f":"'DF raw data'!N13/100","v":5.4298339999999996},{"$f":"'DF raw data'!O13/100","v":10.846034999999999},{"$f":"'DF raw data'!S13/100","v":-6.6905652},{"$f":"'DF raw data'!T13/100","v":20.078524999999999},{"$f":"'DF raw data'!X13/100","v":-11.142682000000001},{"$f":"'DF raw data'!Y13/100","v":13.513340000000001},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":3.8823993000000003},{"$f":"'DF raw data'!E14/100","v":4.5280176000000001},{"$f":"'DF raw data'!I14/100","v":4.9501276000000001},{"$f":"'DF raw data'!J14/100","v":-0.29014649999999997},{"$f":"'DF raw data'!N14/100","v":-1.3847094999999998},{"$f":"'DF raw data'!O14/100","v":16.286142999999999},{"$f":"'DF raw data'!S14/100","v":0.92573357000000001},{"$f":"'DF raw data'!T14/100","v":15.806913999999999},{"$f":"'DF raw data'!X14/100","v":-6.1537194999999993},{"$f":"'DF raw data'!Y14/100","v":-0.71828132999999994},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-28.009834000000001},{"$f":"'DF raw data'!E15/100","v":17.727802999999998},{"$f":"'DF raw data'!I15/100","v":-11.668025},{"$f":"'DF raw data'!J15/100","v":5.3018358999999995},{"$f":"'DF raw data'!N15/100","v":-9.4946631000000004},{"$f":"'DF raw data'!O15/100","v":1.6364941},{"$f":"'DF raw data'!S15/100","v":-18.958601000000002},{"$f":"'DF raw data'!T15/100","v":-26.156610999999998},{"$f":"'DF raw data'!X15/100","v":-17.975885999999999},{"$f":"'DF raw data'!Y15/100","v":-6.9624022999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-8.7334289999999992},{"$f":"'DF raw data'!E16/100","v":8.3716308999999995},{"$f":"'DF raw data'!I16/100","v":4.3152917000000004},{"$f":"'DF raw data'!J16/100","v":25.311620999999999},{"$f":"'DF raw data'!N16/100","v":-3.2140955000000004},{"$f":"'DF raw data'!O16/100","v":34.227705},{"$f":"'DF raw data'!S16/100","v":-1.5454625999999998},{"$f":"'DF raw data'!T16/100","v":20.065850000000001},{"$f":"'DF raw data'!X16/100","v":-1.9306641},{"$f":"'DF raw data'!Y16/100","v":0.68472655999999998}]
[6,{"$f":"'DF raw data'!D17/100","v":-12.672515000000001},{"$f":"'DF raw data'!E17/100","v":16.404140999999999},{"$f":"'DF raw data'!I17/100","v":-9.0605609000000005},{"$f":"'DF raw data'!J17/100","v":12.150469000000001},{"$f":"'DF raw data'!N17/100","v":4.0820553999999998},{"$f":"'DF raw data'!O17/100","v":2.0492383000000003},{"$f":"'DF raw data'!S17/100","v":-19.345554},{"$f":"'DF raw data'!T17/100","v":9.8375878999999991},{"$f":"'DF raw data'!X17/100","v":-11.272644},{"$f":"'DF raw data'!Y17/100","v":-13.17624}]
[8,{"$f":"'DF raw data'!D18/100","v":6.5364459000000004},{"$f":"'DF raw data'!E18/100","v":20.174472999999999},{"$f":"'DF raw data'!I18/100","v":-11.454801},{"$f":"'DF raw data'!J18/100","v":27.025273000000002},{"$f":"'DF raw data'!N18/100","v":-7.1752844000000007},{"$f":"'DF raw data'!O18/100","v":29.57959},{"$f":"'DF raw data'!S18/100","v":1.7385046},{"$f":"'DF raw data'!T18/100","v":4.2953320000000001},{"$f":"'DF raw data'!X18/100","v":-12.988968999999999},{"$f":"'DF raw data'!Y18/100","v":-9.4304784999999995}]
[10,{"$f":"'DF raw data'!D19/100","v":-21.874718999999999},{"$f":"'DF raw data'!E19/100","v":42.190263999999999},{"$f":"'DF raw data'!I19/100","v":-7.7320385999999992},{"$f":"'DF raw data'!J19/100","v":26.099735999999996},{"$f":"'DF raw data'!N19/100","v":-32.302236000000001},{"$f":"'DF raw data'!O19/100","v":28.991220999999999},{"$f":"'DF raw data'!S19/100","v":-13.084940999999999},{"$f":"'DF raw data'!T19/100","v":33.920811},{"$f":"'DF raw data'!X19/100","v":-5.1058093000000007},{"$f":"'DF raw data'!Y19/100","v":-3.4694141000000003}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":232},null,{"$f":"'DF raw data'!G4","v":233},null,{"$f":"'DF raw data'!L4","v":229},null,{"$f":"'DF raw data'!Q4","v":230},null,{"$f":"'DF raw data'!V4","v":231}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-6.7697853666666674},{"$f":"AVERAGE(C11:C13)","v":11.551087199999998},-1.6651326666666666,4.8622721333333336,-1.8165128666666668,9.589557366666666,-8.2411442099999999,3.2429426666666665,-11.757429166666668,1.9442187900000005]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16335 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16335 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16335 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16335 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16335 100 mm gap"]
["Date:","7. mar 2017","13:24:37",null,null,"Date:","7. mar 2017","13:41:31",null,null,"Date:","7. mar 2017","12:40:26",null,null,"Date:","7. mar 2017","12:55:23",null,null,"Date:","7. mar 2017","13:10:14"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",232,null,null,null,"Indice:",233,null,null,null,"Indice:",229,null,null,null,"Indice:",230,null,null,null,"Indice:",231]
["Background Indice:",232,null,null,null,"Background Indice:",233,null,null,null,"Background Indice:",229,null,null,null,"Background Indice:",230,null,null,null,"Background Indice:",231]
["Temperature:",20.99,null,null,null,"Temperature:",21.04,null,null,null,"Temperature:",20.95,null,null,null,"Temperature:",20.95,null,null,null,"Temperature:",20.98]
["Encoder:",386573,null,null,null,"Encoder:",487201,null,null,null,"Encoder:",587536,null,null,null,"Encoder:",787807,null,null,null,"Encoder:",2187952]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,5.1468600999999996,-6.7997499000000001,-854.79132000000004,-808.70507999999995,-1,7.5170002999999999E-2,-11.024990000000001,-3526.4868000000001,-383.37108999999998,-1,7.6790700000000003,-20.422661000000002,703.1825,-2640.1895,-1,-9.6972103000000001,-7.8865800000000004,-2486.8661999999999,-46.530270000000002,-1,-0.86945998999999996,-6.7975501999999999,-1189.7815000000001,227.74315999999999]
[-0.80000000999999998,0.66267997000000001,-4.6005402000000002,-2136.0214999999998,-1380.1602,-0.80000000999999998,3.8782499000000001,-10.466760000000001,332.91577000000001,120.66602,-0.80000000999999998,1.28687,-15.7202,724.09942999999998,-421.84473000000003,-0.80000000999999998,2.7096300000000002,-4.8853698000000003,-86.630981000000006,-144.95898,-0.80000000999999998,-2.3414999999999999,-8.1977100000000007,187.41759999999999,-1189.5352]
[-0.60000001999999997,-2.5943000000000001,-4.0891599999999997,-711.97790999999995,289.24511999999999,-0.60000001999999997,-1.19045,-3.3604701000000001,-338.11309999999997,1094.5752,-0.60000001999999997,-4.2116097999999997,-3.5898199000000002,-1419.0183999999999,524.77050999999994,-0.60000001999999997,-1.4494100000000001,2.4734001000000001,-902.46887000000004,1160.4844000000001,-0.60000001999999997,-2.7331300000000001,1.8844799999999999,-1268.7443000000001,406.27733999999998]
[-0.40000001000000002,1.05077,-3.3866700999999999,-2.2128899,996.34473000000003,-0.40000001000000002,-2.6096499,-9.7984895999999999,321.22046,-1467.5673999999999,-0.40000001000000002,-3.8149099,-1.1129100000000001,-354.85937999999999,868.63280999999995,-0.40000001000000002,-2.4322701000000002,1.7782100000000001,-413.20519999999999,1928.6846,-0.40000001000000002,-5.6578102000000001,-1.54958,-1650.54,250.52734000000001]
[-0.2,-0.11337,-5.8978000000000002,381.80786000000001,1239.7440999999999,-0.2,-1.76197,-6.3244300000000004,172.24994000000001,957.5127,-0.2,2.0547198999999998,-3.33846,542.98339999999996,1084.6034999999999,-0.2,1.6110199999999999,-5.3479999E-2,-669.05651999999998,2007.8525,-0.2,-2.8109099999999998,-4.6232901000000002,-1114.2682,1351.3340000000001]
[0,4.0902801000000002,-1.06423,388.23993000000002,452.80176,0,3.8299601000000001,-3.5168900000000001,495.01276000000001,-29.01465,0,-1.9351499999999999,0.48884000999999999,-138.47094999999999,1628.6143,0,-1.87862,-5.9078397999999996,92.573357000000001,1580.6913999999999,0,-5.9774498999999999,-9.3016299999999994,-615.37194999999997,-71.828132999999994]
[0.2,-0.75954001999999998,-3.8912599000000001,-2800.9834000000001,1772.7802999999999,0.2,-5.7591000000000001,-6.1579699999999997,-1166.8025,530.18358999999998,0.2,-4.3663401999999998,-4.20784,-949.46631000000002,163.64940999999999,0.2,-7.3482498999999999,-13.25789,-1895.8601000000001,-2615.6610999999998,0.2,-11.86341,-7.8213100000000004,-1797.5886,-696.24023]
[0.40000001000000002,5.3397198000000001,11.23719,-873.34289999999999,837.16309000000001,0.40000001000000002,2.3629999000000002,1.6575899999999999,431.52917000000002,2531.1621,0.40000001000000002,-0.77370000000000005,5.5411099999999998,-321.40955000000002,3422.7705000000001,0.40000001000000002,4.2098398000000001,6.3607798000000004,-154.54625999999999,2006.585,0.40000001000000002,-0.32512000000000002,-0.71147000999999999,-193.06640999999999,68.472656000000001]
[0.60000001999999997,-6.0624900000000004,5.8829298000000003,-1267.2515000000001,1640.4141,0.60000001999999997,-0.26941000999999998,1.7905,-906.05609000000004,1215.0469000000001,0.60000001999999997,0.54984999000000001,-1.3202499999999999,408.20553999999998,204.92383000000001,0.60000001999999997,-7.8935699000000001,2.4807301000000002,-1934.5554,983.75878999999998,0.60000001999999997,-4.4871100999999998,-2.09849,-1127.2644,-1317.624]
[0.80000000999999998,-2.3843200000000002,6.5428499999999996,653.64458999999999,2017.4473,0.80000000999999998,-1.0101500000000001,6.6595402000000004,-1145.4801,2702.5273000000002,0.80000000999999998,-3.5115699999999999,4.0817098999999999,-717.52844000000005,2957.9589999999998,0.80000000999999998,-1.9806299999999999,2.2995999,173.85046,429.53320000000002,0.80000000999999998,-10.90578,-6.0621400000000003,-1298.8969,-943.04785000000004]
[1,-3.0720000000000001,18.849139999999998,-2187.4719,4219.0263999999997,1,-1.83717,13.95645,-773.20385999999996,2609.9735999999998,1,-8.9072703999999998,11.881169999999999,-3230.2235999999998,2899.1221,1,-5.2515897999999996,13.952109999999999,-1308.4940999999999,3392.0810999999999,1,-2.8818901000000001,1.1548001000000001,-510.58093000000002,-346.94141000000002]