{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":83,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":388644},null,{"$f":"'DF raw data'!G7","v":489281},null,{"$f":"'DF raw data'!L7","v":589585},null,{"$f":"'DF raw data'!Q7","v":789909},null,{"$f":"'DF raw data'!V7","v":2189859},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-3.7245599999999999},{"$f":"C21","v":1.6266766333333333},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":14.554163300000001},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":4.4056699999999998}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":-1.0861934},{"$f":"E21","v":3.1427332666666667},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":14.554163300000001},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":12.424156566666667}]
[-10,{"$f":"'DF raw data'!B9","v":52.045990000000003},{"$f":"'DF raw data'!C9","v":18.52861},{"$f":"'DF raw data'!G9","v":41.032710999999999},{"$f":"'DF raw data'!H9","v":11.269159999999999},{"$f":"'DF raw data'!L9","v":37.182850000000002},{"$f":"'DF raw data'!M9","v":14.28046},{"$f":"'DF raw data'!Q9","v":23.509990999999999},{"$f":"'DF raw data'!R9","v":0.25845000000000001},{"$f":"'DF raw data'!V9","v":6.1549000999999999},{"$f":"'DF raw data'!W9","v":-5.5672797999999997},null,{"$f":"F4","v":20},{"$f":"F21","v":5.0998900666666662},{"$f":"G21","v":5.4897500000000008}]
[-8,{"$f":"'DF raw data'!B10","v":42.20937},{"$f":"'DF raw data'!C10","v":19.713989000000002},{"$f":"'DF raw data'!G10","v":40.653838999999998},{"$f":"'DF raw data'!H10","v":8.1293602000000007},{"$f":"'DF raw data'!L10","v":28.794170000000001},{"$f":"'DF raw data'!M10","v":4.9281001},{"$f":"'DF raw data'!Q10","v":24.495560000000001},{"$f":"'DF raw data'!R10","v":-3.3531499},{"$f":"'DF raw data'!V10","v":9.6855496999999993},{"$f":"'DF raw data'!W10","v":-12.142910000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":10.8296033},{"$f":"I21","v":6.0323466333333329},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":24.847401000000001},{"$f":"'DF raw data'!C11","v":21.394839999999999},{"$f":"'DF raw data'!G11","v":9.4712601000000003},{"$f":"'DF raw data'!H11","v":19.804970000000001},{"$f":"'DF raw data'!L11","v":19.687419999999999},{"$f":"'DF raw data'!M11","v":5.9604701999999996},{"$f":"'DF raw data'!Q11","v":17.336161000000001},{"$f":"'DF raw data'!R11","v":4.5468897999999998},{"$f":"'DF raw data'!V11","v":2.14852},{"$f":"'DF raw data'!W11","v":-4.3221698000000002},null,{"$f":"J4","v":100},{"$f":"J21","v":7.9189766666666666},{"$f":"K21","v":-6.3918099333333336},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":6.0981101999999998},{"$f":"'DF raw data'!C12","v":19.18965},{"$f":"'DF raw data'!G12","v":3.0897999},{"$f":"'DF raw data'!H12","v":11.991630000000001},{"$f":"'DF raw data'!L12","v":8.5551300000000001},{"$f":"'DF raw data'!M12","v":8.7559099000000007},{"$f":"'DF raw data'!Q12","v":11.92155},{"$f":"'DF raw data'!R12","v":9.1726799000000003},{"$f":"'DF raw data'!V12","v":-1.26485},{"$f":"'DF raw data'!W12","v":1.4119101000000001},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":1.34721},{"$f":"'DF raw data'!C13","v":-3.6117300999999999},{"$f":"'DF raw data'!G13","v":-5.1676102000000004},{"$f":"'DF raw data'!H13","v":11.45974},{"$f":"'DF raw data'!L13","v":9.9138002000000007},{"$f":"'DF raw data'!M13","v":8.4192400000000003},{"$f":"'DF raw data'!Q13","v":12.64889},{"$f":"'DF raw data'!R13","v":7.6207199000000001},{"$f":"'DF raw data'!V13","v":1.6166},{"$f":"'DF raw data'!W13","v":-8.7295798999999992},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":1.01112},{"$f":"'DF raw data'!C14","v":5.8788198999999999},{"$f":"'DF raw data'!G14","v":3.4217300000000002},{"$f":"'DF raw data'!H14","v":2.6669198999999999},{"$f":"'DF raw data'!L14","v":3.0504500999999999},{"$f":"'DF raw data'!M14","v":4.3760500000000002},{"$f":"'DF raw data'!Q14","v":13.29926},{"$f":"'DF raw data'!R14","v":3.0282098999999998},{"$f":"'DF raw data'!V14","v":7.7366299999999999},{"$f":"'DF raw data'!W14","v":-3.5502900999999998},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-13.53201},{"$f":"'DF raw data'!C15","v":2.6129400999999999},{"$f":"'DF raw data'!G15","v":-1.5126999999999999},{"$f":"'DF raw data'!H15","v":-4.6984601000000001},{"$f":"'DF raw data'!L15","v":2.3354199000000002},{"$f":"'DF raw data'!M15","v":3.6739600000000001},{"$f":"'DF raw data'!Q15","v":6.5406598999999996},{"$f":"'DF raw data'!R15","v":7.4481101000000001},{"$f":"'DF raw data'!V15","v":14.403700000000001},{"$f":"'DF raw data'!W15","v":-6.8955598},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-27.136339},{"$f":"'DF raw data'!C16","v":-15.071949999999999},{"$f":"'DF raw data'!G16","v":-7.5747999999999998},{"$f":"'DF raw data'!H16","v":-3.3517499000000002},{"$f":"'DF raw data'!L16","v":-3.6051299999999999},{"$f":"'DF raw data'!M16","v":5.3063001999999999},{"$f":"'DF raw data'!Q16","v":4.7175298000000003},{"$f":"'DF raw data'!R16","v":2.5804800999999999},{"$f":"'DF raw data'!V16","v":7.3000698000000002},{"$f":"'DF raw data'!W16","v":-4.3097801000000002}]
[6,{"$f":"'DF raw data'!B17","v":-24.696711000000001},{"$f":"'DF raw data'!C17","v":-20.446650000000002},{"$f":"'DF raw data'!G17","v":-14.366009999999999},{"$f":"'DF raw data'!H17","v":-23.87726},{"$f":"'DF raw data'!L17","v":-1.1813499999999999},{"$f":"'DF raw data'!M17","v":3.3218000000000001},{"$f":"'DF raw data'!Q17","v":3.2412801},{"$f":"'DF raw data'!R17","v":6.5943198000000001},{"$f":"'DF raw data'!V17","v":8.0713700999999993},{"$f":"'DF raw data'!W17","v":2.5059999999999998}]
[8,{"$f":"'DF raw data'!B18","v":-19.3673},{"$f":"'DF raw data'!C18","v":-21.799009000000002},{"$f":"'DF raw data'!G18","v":-21.306450000000002},{"$f":"'DF raw data'!H18","v":-2.4457200000000001},{"$f":"'DF raw data'!L18","v":-6.1117501000000001},{"$f":"'DF raw data'!M18","v":2.0776800999999998},{"$f":"'DF raw data'!Q18","v":8.3999003999999999},{"$f":"'DF raw data'!R18","v":6.5926999999999998},{"$f":"'DF raw data'!V18","v":7.6182299000000002},{"$f":"'DF raw data'!W18","v":1.75431}]
[10,{"$f":"'DF raw data'!B19","v":-13.510490000000001},{"$f":"'DF raw data'!C19","v":-15.65973},{"$f":"'DF raw data'!G19","v":-12.900040000000001},{"$f":"'DF raw data'!H19","v":-4.7103900999999997},{"$f":"'DF raw data'!L19","v":-3.7822198999999999},{"$f":"'DF raw data'!M19","v":-0.88490999000000004},{"$f":"'DF raw data'!Q19","v":6.5020699999999998},{"$f":"'DF raw data'!R19","v":10.50032},{"$f":"'DF raw data'!V19","v":5.8820499999999996},{"$f":"'DF raw data'!W19","v":-2.4812500000000002}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":329},null,{"$f":"'DF raw data'!G4","v":336},null,{"$f":"'DF raw data'!L4","v":331},null,{"$f":"'DF raw data'!Q4","v":335},null,{"$f":"'DF raw data'!V4","v":333}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-3.7245599999999999},{"$f":"AVERAGE(C11:C13)","v":1.6266766333333333},-1.0861934,3.1427332666666667,5.0998900666666662,5.4897500000000008,10.8296033,6.0323466333333329,7.9189766666666666,-6.3918099333333336]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":388644},null,{"$f":"'DF raw data'!G7","v":489281},null,{"$f":"'DF raw data'!L7","v":589585},null,{"$f":"'DF raw data'!Q7","v":789909},null,{"$f":"'DF raw data'!V7","v":2189859},null,null,{"$f":"B4","v":10},{"$f":"B21","v":-3.6202446000000008},{"$f":"C21","v":5.7224998666666664},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":17.075410166666664},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":9.4639192333333337}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":-9.5596837333333315},{"$f":"E21","v":12.584439999999999},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":17.075410166666664},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":27.628118333333333}]
[-10,{"$f":"'DF raw data'!D9/100","v":63.559813999999996},{"$f":"'DF raw data'!E9/100","v":38.492676000000003},{"$f":"'DF raw data'!I9/100","v":64.773706000000004},{"$f":"'DF raw data'!J9/100","v":18.922432000000001},{"$f":"'DF raw data'!N9/100","v":40.430664},{"$f":"'DF raw data'!O9/100","v":26.823437999999999},{"$f":"'DF raw data'!S9/100","v":30.047440999999999},{"$f":"'DF raw data'!T9/100","v":-1.1090723},{"$f":"'DF raw data'!X9/100","v":-3.8917748999999997},{"$f":"'DF raw data'!Y9/100","v":-9.0780566},null,{"$f":"F4","v":20},{"$f":"F21","v":-0.33699520000000005},{"$f":"G21","v":13.813293999999999}]
[-8,{"$f":"'DF raw data'!D10/100","v":66.902280000000005},{"$f":"'DF raw data'!E10/100","v":49.968896000000001},{"$f":"'DF raw data'!I10/100","v":41.317559000000003},{"$f":"'DF raw data'!J10/100","v":34.701309000000002},{"$f":"'DF raw data'!N10/100","v":34.925260999999999},{"$f":"'DF raw data'!O10/100","v":17.678379},{"$f":"'DF raw data'!S10/100","v":36.925398000000001},{"$f":"'DF raw data'!T10/100","v":8.7653613000000004},{"$f":"'DF raw data'!X10/100","v":6.3384308000000003},{"$f":"'DF raw data'!Y10/100","v":-17.767520000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":7.5157264333333345},{"$f":"I21","v":15.1864191}]
[-6,{"$f":"'DF raw data'!D11/100","v":26.465187999999998},{"$f":"'DF raw data'!E11/100","v":51.068477000000001},{"$f":"'DF raw data'!I11/100","v":29.668957999999996},{"$f":"'DF raw data'!J11/100","v":50.334902},{"$f":"'DF raw data'!N11/100","v":28.389462999999999},{"$f":"'DF raw data'!O11/100","v":27.251435999999998},{"$f":"'DF raw data'!S11/100","v":22.150739999999999},{"$f":"'DF raw data'!T11/100","v":19.723701000000002},{"$f":"'DF raw data'!X11/100","v":20.492280000000001},{"$f":"'DF raw data'!Y11/100","v":-2.6822851999999999},null,{"$f":"J4","v":100},{"$f":"J21","v":3.7113040666666666},{"$f":"K21","v":-12.441699233333333}]
[-4,{"$f":"'DF raw data'!D12/100","v":7.8854949999999997},{"$f":"'DF raw data'!E12/100","v":37.648651999999998},{"$f":"'DF raw data'!I12/100","v":-1.5410597000000001},{"$f":"'DF raw data'!J12/100","v":35.631230000000002},{"$f":"'DF raw data'!N12/100","v":9.3671135999999997},{"$f":"'DF raw data'!O12/100","v":27.537676000000001},{"$f":"'DF raw data'!S12/100","v":13.286891000000001},{"$f":"'DF raw data'!T12/100","v":8.9186426000000001},{"$f":"'DF raw data'!X12/100","v":-10.118615999999999},{"$f":"'DF raw data'!Y12/100","v":5.7172754000000001}]
[-2,{"$f":"'DF raw data'!D13/100","v":7.2343726000000004},{"$f":"'DF raw data'!E13/100","v":8.8763281000000003},{"$f":"'DF raw data'!I13/100","v":-8.4441057999999991},{"$f":"'DF raw data'!J13/100","v":50.354247999999998},{"$f":"'DF raw data'!N13/100","v":13.335844999999999},{"$f":"'DF raw data'!O13/100","v":26.44248},{"$f":"'DF raw data'!S13/100","v":16.906753000000002},{"$f":"'DF raw data'!T13/100","v":23.040898000000002},{"$f":"'DF raw data'!X13/100","v":-9.9874103000000005},{"$f":"'DF raw data'!Y13/100","v":-6.3283690999999997},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":-4.8664513999999999},{"$f":"'DF raw data'!E14/100","v":-9.7693944999999989},{"$f":"'DF raw data'!I14/100","v":-12.279628000000001},{"$f":"'DF raw data'!J14/100","v":-10.320469000000001},{"$f":"'DF raw data'!N14/100","v":-14.094918999999999},{"$f":"'DF raw data'!O14/100","v":-1.6057910000000002},{"$f":"'DF raw data'!S14/100","v":12.825718},{"$f":"'DF raw data'!T14/100","v":-4.4860547000000004},{"$f":"'DF raw data'!X14/100","v":-1.6941044999999999},{"$f":"'DF raw data'!Y14/100","v":-21.191172000000002},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":-13.228655000000002},{"$f":"'DF raw data'!E15/100","v":18.060565999999998},{"$f":"'DF raw data'!I15/100","v":-7.9553174000000002},{"$f":"'DF raw data'!J15/100","v":-2.280459},{"$f":"'DF raw data'!N15/100","v":-0.25191160000000001},{"$f":"'DF raw data'!O15/100","v":16.603193000000001},{"$f":"'DF raw data'!S15/100","v":-7.1852917000000005},{"$f":"'DF raw data'!T15/100","v":27.004414000000001},{"$f":"'DF raw data'!X15/100","v":22.815427},{"$f":"'DF raw data'!Y15/100","v":-9.8055565999999992},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-61.120380999999995},{"$f":"'DF raw data'!E16/100","v":-3.4206249999999998},{"$f":"'DF raw data'!I16/100","v":-36.824209000000003},{"$f":"'DF raw data'!J16/100","v":12.396582},{"$f":"'DF raw data'!N16/100","v":-21.510786},{"$f":"'DF raw data'!O16/100","v":14.761884999999999},{"$f":"'DF raw data'!S16/100","v":-13.862098000000001},{"$f":"'DF raw data'!T16/100","v":20.832997999999996},{"$f":"'DF raw data'!X16/100","v":-1.1939392},{"$f":"'DF raw data'!Y16/100","v":-5.8019628999999995}]
[6,{"$f":"'DF raw data'!D17/100","v":-43.870522000000001},{"$f":"'DF raw data'!E17/100","v":-14.066953000000002},{"$f":"'DF raw data'!I17/100","v":-25.878242},{"$f":"'DF raw data'!J17/100","v":-16.397969},{"$f":"'DF raw data'!N17/100","v":4.5518704000000003},{"$f":"'DF raw data'!O17/100","v":20.771747999999999},{"$f":"'DF raw data'!S17/100","v":4.6415462999999999},{"$f":"'DF raw data'!T17/100","v":19.195498000000001},{"$f":"'DF raw data'!X17/100","v":10.490098},{"$f":"'DF raw data'!Y17/100","v":7.6970410000000005}]
[8,{"$f":"'DF raw data'!D18/100","v":-34.997816999999998},{"$f":"'DF raw data'!E18/100","v":-10.73415},{"$f":"'DF raw data'!I18/100","v":-25.939951000000001},{"$f":"'DF raw data'!J18/100","v":3.9137401999999999},{"$f":"'DF raw data'!N18/100","v":-26.463117999999998},{"$f":"'DF raw data'!O18/100","v":19.251152000000001},{"$f":"'DF raw data'!S18/100","v":11.206737},{"$f":"'DF raw data'!T18/100","v":21.943622999999999},{"$f":"'DF raw data'!X18/100","v":9.9018268000000003},{"$f":"'DF raw data'!Y18/100","v":5.6301074000000009}]
[10,{"$f":"'DF raw data'!D19/100","v":-18.692444000000002},{"$f":"'DF raw data'!E19/100","v":-8.5642382999999995},{"$f":"'DF raw data'!I19/100","v":-16.066044999999999},{"$f":"'DF raw data'!J19/100","v":-3.9955761999999999},{"$f":"'DF raw data'!N19/100","v":13.647969},{"$f":"'DF raw data'!O19/100","v":-9.7033594000000001},{"$f":"'DF raw data'!S19/100","v":14.822085},{"$f":"'DF raw data'!T19/100","v":19.339140999999998},{"$f":"'DF raw data'!X19/100","v":35.868877000000005},{"$f":"'DF raw data'!Y19/100","v":-1.6450977}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":329},null,{"$f":"'DF raw data'!G4","v":336},null,{"$f":"'DF raw data'!L4","v":331},null,{"$f":"'DF raw data'!Q4","v":335},null,{"$f":"'DF raw data'!V4","v":333}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":-3.6202446000000008},{"$f":"AVERAGE(C11:C13)","v":5.7224998666666664},-9.5596837333333315,12.584439999999999,-0.33699520000000005,13.813293999999999,7.5157264333333345,15.1864191,3.7113040666666666,-12.441699233333333]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16342, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16342, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16342, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16342, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16342, 100 mm gap"]
["Date:","28. mar 2017","11:51:27",null,null,"Date:","28. mar 2017","13:59:48",null,null,"Date:","28. mar 2017","12:38:38",null,null,"Date:","28. mar 2017","13:41:03",null,null,"Date:","28. mar 2017","13:09:56"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",329,null,null,null,"Indice:",336,null,null,null,"Indice:",331,null,null,null,"Indice:",335,null,null,null,"Indice:",333]
["Background Indice:",329,null,null,null,"Background Indice:",336,null,null,null,"Background Indice:",331,null,null,null,"Background Indice:",335,null,null,null,"Background Indice:",333]
["Temperature:",21.3,null,null,null,"Temperature:",21.2,null,null,null,"Temperature:",21.25,null,null,null,"Temperature:",21.21,null,null,null,"Temperature:",21.22]
["Encoder:",388644,null,null,null,"Encoder:",489281,null,null,null,"Encoder:",589585,null,null,null,"Encoder:",789909,null,null,null,"Encoder:",2189859]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,52.045990000000003,18.52861,6355.9813999999997,3849.2676000000001,-1,41.032710999999999,11.269159999999999,6477.3706000000002,1892.2431999999999,-1,37.182850000000002,14.28046,4043.0664000000002,2682.3438000000001,-1,23.509990999999999,0.25845000000000001,3004.7440999999999,-110.90723,-1,6.1549000999999999,-5.5672797999999997,-389.17748999999998,-907.80565999999999]
[-0.80000000999999998,42.20937,19.713989000000002,6690.2280000000001,4996.8896000000004,-0.80000000999999998,40.653838999999998,8.1293602000000007,4131.7559000000001,3470.1309000000001,-0.80000000999999998,28.794170000000001,4.9281001,3492.5261,1767.8379,-0.80000000999999998,24.495560000000001,-3.3531499,3692.5398,876.53612999999996,-0.80000000999999998,9.6855496999999993,-12.142910000000001,633.84307999999999,-1776.752]
[-0.60000001999999997,24.847401000000001,21.394839999999999,2646.5187999999998,5106.8477000000003,-0.60000001999999997,9.4712601000000003,19.804970000000001,2966.8957999999998,5033.4902000000002,-0.60000001999999997,19.687419999999999,5.9604701999999996,2838.9463000000001,2725.1435999999999,-0.60000001999999997,17.336161000000001,4.5468897999999998,2215.0740000000001,1972.3701000000001,-0.60000001999999997,2.14852,-4.3221698000000002,2049.2280000000001,-268.22852]
[-0.40000001000000002,6.0981101999999998,19.18965,788.54949999999997,3764.8652000000002,-0.40000001000000002,3.0897999,11.991630000000001,-154.10597000000001,3563.123,-0.40000001000000002,8.5551300000000001,8.7559099000000007,936.71136000000001,2753.7676000000001,-0.40000001000000002,11.92155,9.1726799000000003,1328.6891000000001,891.86425999999994,-0.40000001000000002,-1.26485,1.4119101000000001,-1011.8616,571.72753999999998]
[-0.2,1.34721,-3.6117300999999999,723.43726000000004,887.63280999999995,-0.2,-5.1676102000000004,11.45974,-844.41057999999998,5035.4247999999998,-0.2,9.9138002000000007,8.4192400000000003,1333.5844999999999,2644.248,-0.2,12.64889,7.6207199000000001,1690.6753000000001,2304.0898000000002,-0.2,1.6166,-8.7295798999999992,-998.74103000000002,-632.83690999999999]
[0,1.01112,5.8788198999999999,-486.64514000000003,-976.93944999999997,0,3.4217300000000002,2.6669198999999999,-1227.9628,-1032.0469000000001,0,3.0504500999999999,4.3760500000000002,-1409.4919,-160.57910000000001,0,13.29926,3.0282098999999998,1282.5717999999999,-448.60547000000003,0,7.7366299999999999,-3.5502900999999998,-169.41045,-2119.1172000000001]
[0.2,-13.53201,2.6129400999999999,-1322.8655000000001,1806.0565999999999,0.2,-1.5126999999999999,-4.6984601000000001,-795.53174000000001,-228.04589999999999,0.2,2.3354199000000002,3.6739600000000001,-25.19116,1660.3193000000001,0.2,6.5406598999999996,7.4481101000000001,-718.52917000000002,2700.4414000000002,0.2,14.403700000000001,-6.8955598,2281.5427,-980.55565999999999]
[0.40000001000000002,-27.136339,-15.071949999999999,-6112.0380999999998,-342.0625,0.40000001000000002,-7.5747999999999998,-3.3517499000000002,-3682.4209000000001,1239.6582000000001,0.40000001000000002,-3.6051299999999999,5.3063001999999999,-2151.0785999999998,1476.1885,0.40000001000000002,4.7175298000000003,2.5804800999999999,-1386.2098000000001,2083.2997999999998,0.40000001000000002,7.3000698000000002,-4.3097801000000002,-119.39391999999999,-580.19628999999998]
[0.60000001999999997,-24.696711000000001,-20.446650000000002,-4387.0522000000001,-1406.6953000000001,0.60000001999999997,-14.366009999999999,-23.87726,-2587.8242,-1639.7969000000001,0.60000001999999997,-1.1813499999999999,3.3218000000000001,455.18704000000002,2077.1747999999998,0.60000001999999997,3.2412801,6.5943198000000001,464.15463,1919.5498,0.60000001999999997,8.0713700999999993,2.5059999999999998,1049.0098,769.70410000000004]
[0.80000000999999998,-19.3673,-21.799009000000002,-3499.7817,-1073.415,0.80000000999999998,-21.306450000000002,-2.4457200000000001,-2593.9951000000001,391.37401999999997,0.80000000999999998,-6.1117501000000001,2.0776800999999998,-2646.3117999999999,1925.1152,0.80000000999999998,8.3999003999999999,6.5926999999999998,1120.6737000000001,2194.3622999999998,0.80000000999999998,7.6182299000000002,1.75431,990.18268,563.01074000000006]
[1,-13.510490000000001,-15.65973,-1869.2444,-856.42382999999995,1,-12.900040000000001,-4.7103900999999997,-1606.6044999999999,-399.55761999999999,1,-3.7822198999999999,-0.88490999000000004,1364.7969000000001,-970.33594000000005,1,6.5020699999999998,10.50032,1482.2085,1933.9141,1,5.8820499999999996,-2.4812500000000002,3586.8877000000002,-164.50977]