{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":88,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":383456},null,{"$f":"'DF raw data'!G7","v":484029},null,{"$f":"'DF raw data'!L7","v":584349},null,{"$f":"'DF raw data'!Q7","v":784701},null,{"$f":"'DF raw data'!V7","v":2184518},null,null,{"$f":"B4","v":10},{"$f":"B21","v":2.7101600333333331},{"$f":"C21","v":-5.0316967999999997},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":3.4262001533333333},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":14.161143500000001}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":3.2639934433333333},{"$f":"E21","v":1.9100433366666667},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":4.4965735133333329},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":15.714556833333333}]
[-10,{"$f":"'DF raw data'!B9","v":-9.1125697999999993},{"$f":"'DF raw data'!C9","v":11.51989},{"$f":"'DF raw data'!G9","v":-3.5940799999999999},{"$f":"'DF raw data'!H9","v":12.57155},{"$f":"'DF raw data'!L9","v":-3.5534100999999998},{"$f":"'DF raw data'!M9","v":18.042819999999999},{"$f":"'DF raw data'!Q9","v":-10.040609999999999},{"$f":"'DF raw data'!R9","v":10.69922},{"$f":"'DF raw data'!V9","v":-4.1847601000000001},{"$f":"'DF raw data'!W9","v":10.05259},null,{"$f":"F4","v":20},{"$f":"F21","v":1.8349666666666666},{"$f":"G21","v":7.6105767000000002}]
[-8,{"$f":"'DF raw data'!B10","v":3.81724},{"$f":"'DF raw data'!C10","v":-0.79965001000000002},{"$f":"'DF raw data'!G10","v":9.1711597000000005},{"$f":"'DF raw data'!H10","v":6.1676102000000004},{"$f":"'DF raw data'!L10","v":3.1852901},{"$f":"'DF raw data'!M10","v":12.81185},{"$f":"'DF raw data'!Q10","v":-13.742319999999999},{"$f":"'DF raw data'!R10","v":10.203419999999999},{"$f":"'DF raw data'!V10","v":-3.91927},{"$f":"'DF raw data'!W10","v":4.6807698999999996},null,{"$f":"H4","v":30},{"$f":"H21","v":-0.16220670999999998},{"$f":"I21","v":9.1294467000000008},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":3.5302899000000001},{"$f":"'DF raw data'!C11","v":3.9225500000000002},{"$f":"'DF raw data'!G11","v":-2.83094},{"$f":"'DF raw data'!H11","v":2.3855200000000001},{"$f":"'DF raw data'!L11","v":-0.57270997999999995},{"$f":"'DF raw data'!M11","v":10.56222},{"$f":"'DF raw data'!Q11","v":-9.1785096999999993},{"$f":"'DF raw data'!R11","v":12.33338},{"$f":"'DF raw data'!V11","v":-1.5792200999999999},{"$f":"'DF raw data'!W11","v":7.45261},null,{"$f":"J4","v":100},{"$f":"J21","v":-1.23258007},{"$f":"K21","v":10.682860033333334},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-0.18592},{"$f":"'DF raw data'!C12","v":-2.1330399999999998},{"$f":"'DF raw data'!G12","v":-11.13043},{"$f":"'DF raw data'!H12","v":10.907590000000001},{"$f":"'DF raw data'!L12","v":-2.4865301},{"$f":"'DF raw data'!M12","v":2.1536601000000002},{"$f":"'DF raw data'!Q12","v":-9.2354202000000001},{"$f":"'DF raw data'!R12","v":16.418800000000001},{"$f":"'DF raw data'!V12","v":-6.4222397999999998},{"$f":"'DF raw data'!W12","v":6.1977900999999997},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":4.0454102000000001},{"$f":"'DF raw data'!C13","v":-4.4366202000000001},{"$f":"'DF raw data'!G13","v":0.59759002999999999},{"$f":"'DF raw data'!H13","v":2.2971001000000002},{"$f":"'DF raw data'!L13","v":5.1610000000000003E-2},{"$f":"'DF raw data'!M13","v":6.3981700000000004},{"$f":"'DF raw data'!Q13","v":-3.2609601000000001},{"$f":"'DF raw data'!R13","v":3.9005301000000001},{"$f":"'DF raw data'!V13","v":-2.4353001000000001},{"$f":"'DF raw data'!W13","v":12.23668},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":7.2287898000000004},{"$f":"'DF raw data'!C14","v":-6.5215601999999997},{"$f":"'DF raw data'!G14","v":9.8941803000000004},{"$f":"'DF raw data'!H14","v":2.6370699000000002},{"$f":"'DF raw data'!L14","v":2.3987400999999999},{"$f":"'DF raw data'!M14","v":12.903639999999999},{"$f":"'DF raw data'!Q14","v":0.79926997},{"$f":"'DF raw data'!R14","v":12.339130000000001},{"$f":"'DF raw data'!V14","v":-2.1643800999999998},{"$f":"'DF raw data'!W14","v":8.6991700999999999},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":-3.1437198999999998},{"$f":"'DF raw data'!C15","v":-4.1369100000000003},{"$f":"'DF raw data'!G15","v":-0.69979000000000002},{"$f":"'DF raw data'!H15","v":0.79596001000000005},{"$f":"'DF raw data'!L15","v":3.0545499},{"$f":"'DF raw data'!M15","v":3.5299201},{"$f":"'DF raw data'!Q15","v":1.9750700000000001},{"$f":"'DF raw data'!R15","v":11.148680000000001},{"$f":"'DF raw data'!V15","v":0.90193999000000002},{"$f":"'DF raw data'!W15","v":11.112730000000001},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":-0.54935997999999997},{"$f":"'DF raw data'!C16","v":-2.0408599000000001},{"$f":"'DF raw data'!G16","v":-5.8039097999999996},{"$f":"'DF raw data'!H16","v":7.6091398999999997},{"$f":"'DF raw data'!L16","v":-4.4956497999999998},{"$f":"'DF raw data'!M16","v":9.5803498999999999},{"$f":"'DF raw data'!Q16","v":8.1528396999999995},{"$f":"'DF raw data'!R16","v":17.58201},{"$f":"'DF raw data'!V16","v":-2.1038301000000001},{"$f":"'DF raw data'!W16","v":9.4197903000000007}]
[6,{"$f":"'DF raw data'!B17","v":-4.2107301000000001},{"$f":"'DF raw data'!C17","v":1.2239},{"$f":"'DF raw data'!G17","v":-2.1683699999999999},{"$f":"'DF raw data'!H17","v":6.9493799000000003},{"$f":"'DF raw data'!L17","v":-1.6348799000000001},{"$f":"'DF raw data'!M17","v":0.50765002000000004},{"$f":"'DF raw data'!Q17","v":5.8624301000000001},{"$f":"'DF raw data'!R17","v":13.508190000000001},{"$f":"'DF raw data'!V17","v":0.57176000000000005},{"$f":"'DF raw data'!W17","v":7.5638398999999996}]
[8,{"$f":"'DF raw data'!B18","v":-2.2623701000000001},{"$f":"'DF raw data'!C18","v":9.0809498000000008},{"$f":"'DF raw data'!G18","v":2.0325799},{"$f":"'DF raw data'!H18","v":11.49685},{"$f":"'DF raw data'!L18","v":2.6526898999999999},{"$f":"'DF raw data'!M18","v":7.8895001000000002},{"$f":"'DF raw data'!Q18","v":2.8015498999999999},{"$f":"'DF raw data'!R18","v":23.110558999999999},{"$f":"'DF raw data'!V18","v":-3.9771299},{"$f":"'DF raw data'!W18","v":11.427199999999999}]
[10,{"$f":"'DF raw data'!B19","v":8.9815102000000007},{"$f":"'DF raw data'!C19","v":12.038349999999999},{"$f":"'DF raw data'!G19","v":8.5057697000000001},{"$f":"'DF raw data'!H19","v":17.089728999999998},{"$f":"'DF raw data'!L19","v":4.8576999000000001},{"$f":"'DF raw data'!M19","v":19.740819999999999},{"$f":"'DF raw data'!Q19","v":11.407299999999999},{"$f":"'DF raw data'!R19","v":15.662739999999999},{"$f":"'DF raw data'!V19","v":-8.9100002999999997E-2},{"$f":"'DF raw data'!W19","v":5.8795500000000001}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":344},null,{"$f":"'DF raw data'!G4","v":345},null,{"$f":"'DF raw data'!L4","v":346},null,{"$f":"'DF raw data'!Q4","v":347},null,{"$f":"'DF raw data'!V4","v":348}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":2.7101600333333331},{"$f":"AVERAGE(C11:C13)","v":-5.0316967999999997},3.2639934433333333,1.9100433366666667,1.8349666666666666,7.6105767000000002,-0.16220670999999998,9.1294467000000008,-1.23258007,10.682860033333334]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":383456},null,{"$f":"'DF raw data'!G7","v":484029},null,{"$f":"'DF raw data'!L7","v":584349},null,{"$f":"'DF raw data'!Q7","v":784701},null,{"$f":"'DF raw data'!V7","v":2184518},null,null,{"$f":"B4","v":10},{"$f":"B21","v":10.739007233333332},{"$f":"C21","v":-0.18069000000000024},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":6.4777531033333329},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":22.351725333333334}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":13.586923233333335},{"$f":"E21","v":17.590358233333333},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":7.3623285033333321},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":22.351725333333334}]
[-10,{"$f":"'DF raw data'!D9/100","v":-28.641304000000002},{"$f":"'DF raw data'!E9/100","v":31.116093999999997},{"$f":"'DF raw data'!I9/100","v":-1.2290021},{"$f":"'DF raw data'!J9/100","v":24.494834000000001},{"$f":"'DF raw data'!N9/100","v":4.2022952},{"$f":"'DF raw data'!O9/100","v":30.490995999999999},{"$f":"'DF raw data'!S9/100","v":-11.573227999999999},{"$f":"'DF raw data'!T9/100","v":21.634931999999999},{"$f":"'DF raw data'!X9/100","v":0.55035831000000002},{"$f":"'DF raw data'!Y9/100","v":12.986190999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":15.696135436666665},{"$f":"G21","v":22.171035333333332}]
[-8,{"$f":"'DF raw data'!D10/100","v":19.067222000000001},{"$f":"'DF raw data'!E10/100","v":30.389092000000002},{"$f":"'DF raw data'!I10/100","v":10.031718},{"$f":"'DF raw data'!J10/100","v":28.109668000000003},{"$f":"'DF raw data'!N10/100","v":4.3670546999999997},{"$f":"'DF raw data'!O10/100","v":34.716270000000002},{"$f":"'DF raw data'!S10/100","v":-21.447593000000001},{"$f":"'DF raw data'!T10/100","v":34.049131000000003},{"$f":"'DF raw data'!X10/100","v":-5.5419745000000002},{"$f":"'DF raw data'!Y10/100","v":17.728545},null,{"$f":"H4","v":30},{"$f":"H21","v":9.2183823333333326},{"$f":"I21","v":13.454300066666667}]
[-6,{"$f":"'DF raw data'!D11/100","v":17.872244000000002},{"$f":"'DF raw data'!E11/100","v":22.846523000000001},{"$f":"'DF raw data'!I11/100","v":15.510134000000001},{"$f":"'DF raw data'!J11/100","v":26.558368999999999},{"$f":"'DF raw data'!N11/100","v":-4.0261999000000001E-3},{"$f":"'DF raw data'!O11/100","v":37.637764000000004},{"$f":"'DF raw data'!S11/100","v":-24.250532},{"$f":"'DF raw data'!T11/100","v":32.769424000000001},{"$f":"'DF raw data'!X11/100","v":4.6143758999999998},{"$f":"'DF raw data'!Y11/100","v":16.192354000000002},null,{"$f":"J4","v":100},{"$f":"J21","v":8.3338069333333333},{"$f":"K21","v":11.914352266666667}]
[-4,{"$f":"'DF raw data'!D12/100","v":9.7475451999999994},{"$f":"'DF raw data'!E12/100","v":11.915146},{"$f":"'DF raw data'!I12/100","v":-7.7668053999999991},{"$f":"'DF raw data'!J12/100","v":17.776592000000001},{"$f":"'DF raw data'!N12/100","v":2.3557334999999999},{"$f":"'DF raw data'!O12/100","v":7.9472753999999997},{"$f":"'DF raw data'!S12/100","v":-19.954027},{"$f":"'DF raw data'!T12/100","v":37.998682000000002},{"$f":"'DF raw data'!X12/100","v":1.4806725000000001},{"$f":"'DF raw data'!Y12/100","v":15.030996}]
[-2,{"$f":"'DF raw data'!D13/100","v":13.522788},{"$f":"'DF raw data'!E13/100","v":6.7900098},{"$f":"'DF raw data'!I13/100","v":8.9662750000000013},{"$f":"'DF raw data'!J13/100","v":31.050810999999999},{"$f":"'DF raw data'!N13/100","v":13.507233000000001},{"$f":"'DF raw data'!O13/100","v":28.065390999999998},{"$f":"'DF raw data'!S13/100","v":5.6266655999999999},{"$f":"'DF raw data'!T13/100","v":8.3690332000000005},{"$f":"'DF raw data'!X13/100","v":3.3858120999999999},{"$f":"'DF raw data'!Y13/100","v":17.108711},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":15.803317},{"$f":"'DF raw data'!E14/100","v":-11.357197000000001},{"$f":"'DF raw data'!I14/100","v":29.015752000000003},{"$f":"'DF raw data'!J14/100","v":1.1465137000000001},{"$f":"'DF raw data'!N14/100","v":-0.94123169000000007},{"$f":"'DF raw data'!O14/100","v":15.704502},{"$f":"'DF raw data'!S14/100","v":-8.8795165999999988},{"$f":"'DF raw data'!T14/100","v":10.299423999999998},{"$f":"'DF raw data'!X14/100","v":8.2233087000000005},{"$f":"'DF raw data'!Y14/100","v":2.0372167999999999},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":2.8909167},{"$f":"'DF raw data'!E15/100","v":4.0251172000000004},{"$f":"'DF raw data'!I15/100","v":2.7787427000000005},{"$f":"'DF raw data'!J15/100","v":20.57375},{"$f":"'DF raw data'!N15/100","v":34.522404999999999},{"$f":"'DF raw data'!O15/100","v":22.743213000000001},{"$f":"'DF raw data'!S15/100","v":30.907997999999999},{"$f":"'DF raw data'!T15/100","v":21.694443},{"$f":"'DF raw data'!X15/100","v":13.392300000000001},{"$f":"'DF raw data'!Y15/100","v":16.597128999999999},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":-13.209982},{"$f":"'DF raw data'!E16/100","v":8.1316699000000003},{"$f":"'DF raw data'!I16/100","v":-17.044294000000001},{"$f":"'DF raw data'!J16/100","v":40.607734000000001},{"$f":"'DF raw data'!N16/100","v":-5.9154181000000001},{"$f":"'DF raw data'!O16/100","v":9.4508594000000006},{"$f":"'DF raw data'!S16/100","v":4.2528310999999999},{"$f":"'DF raw data'!T16/100","v":34.717559000000001},{"$f":"'DF raw data'!X16/100","v":-5.6188549999999999},{"$f":"'DF raw data'!Y16/100","v":22.778359000000002}]
[6,{"$f":"'DF raw data'!D17/100","v":3.1956952000000003},{"$f":"'DF raw data'!E17/100","v":22.274756000000004},{"$f":"'DF raw data'!I17/100","v":5.5865844999999998},{"$f":"'DF raw data'!J17/100","v":31.963359000000001},{"$f":"'DF raw data'!N17/100","v":11.182836999999999},{"$f":"'DF raw data'!O17/100","v":7.8402246},{"$f":"'DF raw data'!S17/100","v":0.9283264200000001},{"$f":"'DF raw data'!T17/100","v":38.314774999999997},{"$f":"'DF raw data'!X17/100","v":10.366104},{"$f":"'DF raw data'!Y17/100","v":20.389530999999998}]
[8,{"$f":"'DF raw data'!D18/100","v":1.8951629999999999},{"$f":"'DF raw data'!E18/100","v":34.773662000000002},{"$f":"'DF raw data'!I18/100","v":-1.9614471},{"$f":"'DF raw data'!J18/100","v":42.238340000000001},{"$f":"'DF raw data'!N18/100","v":13.287988},{"$f":"'DF raw data'!O18/100","v":21.853057},{"$f":"'DF raw data'!S18/100","v":17.267548999999999},{"$f":"'DF raw data'!T18/100","v":58.662744000000004},{"$f":"'DF raw data'!X18/100","v":-3.3914263999999998},{"$f":"'DF raw data'!Y18/100","v":27.974229000000001}]
[10,{"$f":"'DF raw data'!D19/100","v":29.637287999999998},{"$f":"'DF raw data'!E19/100","v":29.113622999999997},{"$f":"'DF raw data'!I19/100","v":21.108328},{"$f":"'DF raw data'!J19/100","v":42.061454999999995},{"$f":"'DF raw data'!N19/100","v":20.252445999999999},{"$f":"'DF raw data'!O19/100","v":39.466045000000001},{"$f":"'DF raw data'!S19/100","v":32.49812},{"$f":"'DF raw data'!T19/100","v":37.592021000000003},{"$f":"'DF raw data'!X19/100","v":12.313101},{"$f":"'DF raw data'!Y19/100","v":8.0708007999999989}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":344},null,{"$f":"'DF raw data'!G4","v":345},null,{"$f":"'DF raw data'!L4","v":346},null,{"$f":"'DF raw data'!Q4","v":347},null,{"$f":"'DF raw data'!V4","v":348}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":10.739007233333332},{"$f":"AVERAGE(C11:C13)","v":-0.18069000000000024},13.586923233333335,17.590358233333333,15.696135436666665,22.171035333333332,9.2183823333333326,13.454300066666667,8.3338069333333333,11.914352266666667]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16343, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16343, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16343, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16343, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16343, 100 mm gap"]
["Date:","29. mar 2017","14:49:24",null,null,"Date:","29. mar 2017","15:06:33",null,null,"Date:","29. mar 2017","15:21:45",null,null,"Date:","29. mar 2017","15:36:07",null,null,"Date:","29. mar 2017","15:51:33"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",344,null,null,null,"Indice:",345,null,null,null,"Indice:",346,null,null,null,"Indice:",347,null,null,null,"Indice:",348]
["Background Indice:",344,null,null,null,"Background Indice:",345,null,null,null,"Background Indice:",346,null,null,null,"Background Indice:",347,null,null,null,"Background Indice:",348]
["Temperature:",21.43,null,null,null,"Temperature:",21.36,null,null,null,"Temperature:",21.32,null,null,null,"Temperature:",21.3,null,null,null,"Temperature:",21.29]
["Encoder:",383456,null,null,null,"Encoder:",484029,null,null,null,"Encoder:",584349,null,null,null,"Encoder:",784701,null,null,null,"Encoder:",2184518]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,-9.1125697999999993,11.51989,-2864.1304,3111.6093999999998,-1,-3.5940799999999999,12.57155,-122.90021,2449.4834000000001,-1,-3.5534100999999998,18.042819999999999,420.22951999999998,3049.0996,-1,-10.040609999999999,10.69922,-1157.3227999999999,2163.4931999999999,-1,-4.1847601000000001,10.05259,55.035831000000002,1298.6190999999999]
[-0.80000000999999998,3.81724,-0.79965001000000002,1906.7221999999999,3038.9092000000001,-0.80000000999999998,9.1711597000000005,6.1676102000000004,1003.1718,2810.9668000000001,-0.80000000999999998,3.1852901,12.81185,436.70546999999999,3471.627,-0.80000000999999998,-13.742319999999999,10.203419999999999,-2144.7593000000002,3404.9131000000002,-0.80000000999999998,-3.91927,4.6807698999999996,-554.19745,1772.8544999999999]
[-0.60000001999999997,3.5302899000000001,3.9225500000000002,1787.2244000000001,2284.6523000000002,-0.60000001999999997,-2.83094,2.3855200000000001,1551.0134,2655.8368999999998,-0.60000001999999997,-0.57270997999999995,10.56222,-0.40261998999999998,3763.7764000000002,-0.60000001999999997,-9.1785096999999993,12.33338,-2425.0531999999998,3276.9423999999999,-0.60000001999999997,-1.5792200999999999,7.45261,461.43759,1619.2354]
[-0.40000001000000002,-0.18592,-2.1330399999999998,974.75451999999996,1191.5146,-0.40000001000000002,-11.13043,10.907590000000001,-776.68053999999995,1777.6592000000001,-0.40000001000000002,-2.4865301,2.1536601000000002,235.57335,794.72753999999998,-0.40000001000000002,-9.2354202000000001,16.418800000000001,-1995.4027000000001,3799.8681999999999,-0.40000001000000002,-6.4222397999999998,6.1977900999999997,148.06725,1503.0996]
[-0.2,4.0454102000000001,-4.4366202000000001,1352.2788,679.00098000000003,-0.2,0.59759002999999999,2.2971001000000002,896.62750000000005,3105.0810999999999,-0.2,5.1610000000000003E-2,6.3981700000000004,1350.7233000000001,2806.5391,-0.2,-3.2609601000000001,3.9005301000000001,562.66656,836.90332000000001,-0.2,-2.4353001000000001,12.23668,338.58121,1710.8711000000001]
[0,7.2287898000000004,-6.5215601999999997,1580.3317,-1135.7197000000001,0,9.8941803000000004,2.6370699000000002,2901.5752000000002,114.65137,0,2.3987400999999999,12.903639999999999,-94.123169000000004,1570.4502,0,0.79926997,12.339130000000001,-887.95165999999995,1029.9423999999999,0,-2.1643800999999998,8.6991700999999999,822.33087,203.72167999999999]
[0.2,-3.1437198999999998,-4.1369100000000003,289.09167000000002,402.51172000000003,0.2,-0.69979000000000002,0.79596001000000005,277.87427000000002,2057.375,0.2,3.0545499,3.5299201,3452.2404999999999,2274.3213000000001,0.2,1.9750700000000001,11.148680000000001,3090.7997999999998,2169.4443000000001,0.2,0.90193999000000002,11.112730000000001,1339.23,1659.7129]
[0.40000001000000002,-0.54935997999999997,-2.0408599000000001,-1320.9982,813.16699000000006,0.40000001000000002,-5.8039097999999996,7.6091398999999997,-1704.4294,4060.7734,0.40000001000000002,-4.4956497999999998,9.5803498999999999,-591.54181000000005,945.08594000000005,0.40000001000000002,8.1528396999999995,17.58201,425.28311000000002,3471.7559000000001,0.40000001000000002,-2.1038301000000001,9.4197903000000007,-561.88549999999998,2277.8359]
[0.60000001999999997,-4.2107301000000001,1.2239,319.56952000000001,2227.4756000000002,0.60000001999999997,-2.1683699999999999,6.9493799000000003,558.65845000000002,3196.3359,0.60000001999999997,-1.6348799000000001,0.50765002000000004,1118.2837,784.02246000000002,0.60000001999999997,5.8624301000000001,13.508190000000001,92.832642000000007,3831.4775,0.60000001999999997,0.57176000000000005,7.5638398999999996,1036.6104,2038.9530999999999]
[0.80000000999999998,-2.2623701000000001,9.0809498000000008,189.5163,3477.3661999999999,0.80000000999999998,2.0325799,11.49685,-196.14471,4223.8339999999998,0.80000000999999998,2.6526898999999999,7.8895001000000002,1328.7988,2185.3056999999999,0.80000000999999998,2.8015498999999999,23.110558999999999,1726.7548999999999,5866.2744000000002,0.80000000999999998,-3.9771299,11.427199999999999,-339.14263999999997,2797.4229]
[1,8.9815102000000007,12.038349999999999,2963.7287999999999,2911.3622999999998,1,8.5057697000000001,17.089728999999998,2110.8328000000001,4206.1454999999996,1,4.8576999000000001,19.740819999999999,2025.2446,3946.6044999999999,1,11.407299999999999,15.662739999999999,3249.8119999999999,3759.2021,1,-8.9100002999999997E-2,5.8795500000000001,1231.3100999999999,807.08007999999995]