{"text_format":"1","workbook":"xlsx","n_sheets":3,"sheets":["1st Int","2nd Int","DF raw data"],"not_rendered":{"charts":2,"images":0,"macros":1,"external_links":0,"defined_names":194,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"1st Int","index":0,"n_rows":30,"n_cols":19,"rows_kept":30,"cut":false,"hidden":false,"merged":[]}
["SXR 1st Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap nominal",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Ix [muTm]","Iy [muTm]",null,"Gap interval","Ix [muTm]","Iy [muTm]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386307},null,{"$f":"'DF raw data'!G7","v":486954},null,{"$f":"'DF raw data'!L7","v":587324},null,{"$f":"'DF raw data'!Q7","v":787626},null,{"$f":"'DF raw data'!V7","v":2187630},null,null,{"$f":"B4","v":10},{"$f":"B21","v":9.8369067333333344},{"$f":"C21","v":-8.9408199966666668},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":5.5326129333333327},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":4.9055403366666681}]
["x pos [mm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]","Ix [muTm]","Iy [muTm]",null,{"$f":"D4","v":15},{"$f":"D21","v":15.369519666666667},{"$f":"E21","v":-10.1102367},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":9.0826930666666676},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":9.045483633333335}]
[-10,{"$f":"'DF raw data'!B9","v":8.1423702000000002},{"$f":"'DF raw data'!C9","v":-13.99343},{"$f":"'DF raw data'!G9","v":-7.6044698000000004},{"$f":"'DF raw data'!H9","v":-6.3994799000000002},{"$f":"'DF raw data'!L9","v":8.4387597999999997},{"$f":"'DF raw data'!M9","v":-22.915371},{"$f":"'DF raw data'!Q9","v":14.13165},{"$f":"'DF raw data'!R9","v":-14.360060000000001},{"$f":"'DF raw data'!V9","v":-6.2310299999999996},{"$f":"'DF raw data'!W9","v":-3.3148700999999998},null,{"$f":"F4","v":20},{"$f":"F21","v":14.393370366666666},{"$f":"G21","v":-10.026013333333333}]
[-8,{"$f":"'DF raw data'!B10","v":0.19045000000000001},{"$f":"'DF raw data'!C10","v":-7.2811399000000003},{"$f":"'DF raw data'!G10","v":9.1924495999999998},{"$f":"'DF raw data'!H10","v":-5.6423101000000004},{"$f":"'DF raw data'!L10","v":18.319299999999998},{"$f":"'DF raw data'!M10","v":-8.4756899000000008},{"$f":"'DF raw data'!Q10","v":0.40864},{"$f":"'DF raw data'!R10","v":-8.1246004000000003},{"$f":"'DF raw data'!V10","v":10.92963},{"$f":"'DF raw data'!W10","v":-21.325710000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":12.518829933333334},{"$f":"I21","v":-13.846360333333335},null,"Accept criteria:"]
[-6,{"$f":"'DF raw data'!B11","v":-0.29017000999999998},{"$f":"'DF raw data'!C11","v":-1.9674700000000001},{"$f":"'DF raw data'!G11","v":15.50127},{"$f":"'DF raw data'!H11","v":-9.7119502999999998},{"$f":"'DF raw data'!L11","v":11.019970000000001},{"$f":"'DF raw data'!M11","v":-16.367719999999998},{"$f":"'DF raw data'!Q11","v":12.43378},{"$f":"'DF raw data'!R11","v":-16.908809999999999},{"$f":"'DF raw data'!V11","v":3.9573100000000001},{"$f":"'DF raw data'!W11","v":-31.089399},null,{"$f":"J4","v":100},{"$f":"J21","v":6.2868266000000004},{"$f":"K21","v":-4.8008766999999999},null,"Max for 10-30:",15,15]
[-4,{"$f":"'DF raw data'!B12","v":-8.7370596000000003},{"$f":"'DF raw data'!C12","v":-1.2584599999999999},{"$f":"'DF raw data'!G12","v":2.81758},{"$f":"'DF raw data'!H12","v":-21.298189000000001},{"$f":"'DF raw data'!L12","v":10.430059999999999},{"$f":"'DF raw data'!M12","v":-19.436501},{"$f":"'DF raw data'!Q12","v":11.77398},{"$f":"'DF raw data'!R12","v":-11.913119999999999},{"$f":"'DF raw data'!V12","v":-2.7963499999999999},{"$f":"'DF raw data'!W12","v":-16.779240000000001},null,null,null,null,null,"Test result",{"$f":"IF(R5<R9,\"Pass\",\"Fail\")","v":"Pass"},{"$f":"IF(S5<S9,\"Pass\",\"Fail\")","v":"Pass"}]
[-2,{"$f":"'DF raw data'!B13","v":8.1936903000000001},{"$f":"'DF raw data'!C13","v":0.38256001000000001},{"$f":"'DF raw data'!G13","v":12.85463},{"$f":"'DF raw data'!H13","v":-4.0197000999999997},{"$f":"'DF raw data'!L13","v":8.4867600999999997},{"$f":"'DF raw data'!M13","v":-11.379960000000001},{"$f":"'DF raw data'!Q13","v":15.450760000000001},{"$f":"'DF raw data'!R13","v":-11.29477},{"$f":"'DF raw data'!V13","v":3.4091898999999999},{"$f":"'DF raw data'!W13","v":-3.1791999},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!B14","v":7.0873898999999998},{"$f":"'DF raw data'!C14","v":-16.867629999999998},{"$f":"'DF raw data'!G14","v":18.677019000000001},{"$f":"'DF raw data'!H14","v":-13.08187},{"$f":"'DF raw data'!L14","v":15.30214},{"$f":"'DF raw data'!M14","v":-7.4789399999999997},{"$f":"'DF raw data'!Q14","v":7.9587998000000004},{"$f":"'DF raw data'!R14","v":-10.727690000000001},{"$f":"'DF raw data'!V14","v":10.526160000000001},{"$f":"'DF raw data'!W14","v":-5.6685800999999998},null,"Values within: +/-",17.5,17.5]
[2,{"$f":"'DF raw data'!B15","v":14.22964},{"$f":"'DF raw data'!C15","v":-10.337389999999999},{"$f":"'DF raw data'!G15","v":14.57691},{"$f":"'DF raw data'!H15","v":-13.229139999999999},{"$f":"'DF raw data'!L15","v":19.391210999999998},{"$f":"'DF raw data'!M15","v":-11.219139999999999},{"$f":"'DF raw data'!Q15","v":14.146929999999999},{"$f":"'DF raw data'!R15","v":-19.516621000000001},{"$f":"'DF raw data'!V15","v":4.9251298999999999},{"$f":"'DF raw data'!W15","v":-5.5548501000000003},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!B16","v":19.894010999999999},{"$f":"'DF raw data'!C16","v":-1.4725999999999999},{"$f":"'DF raw data'!G16","v":17.713301000000001},{"$f":"'DF raw data'!H16","v":-17.839808999999999},{"$f":"'DF raw data'!L16","v":12.50811},{"$f":"'DF raw data'!M16","v":-22.097460000000002},{"$f":"'DF raw data'!Q16","v":10.96805},{"$f":"'DF raw data'!R16","v":-18.441680999999999},{"$f":"'DF raw data'!V16","v":3.069},{"$f":"'DF raw data'!W16","v":-17.651520000000001}]
[6,{"$f":"'DF raw data'!B17","v":25.126418999999999},{"$f":"'DF raw data'!C17","v":-9.5381899000000008},{"$f":"'DF raw data'!G17","v":20.553128999999998},{"$f":"'DF raw data'!H17","v":-14.90456},{"$f":"'DF raw data'!L17","v":9.1611499999999992},{"$f":"'DF raw data'!M17","v":-13.01342},{"$f":"'DF raw data'!Q17","v":2.8752},{"$f":"'DF raw data'!R17","v":-5.4107098999999996},{"$f":"'DF raw data'!V17","v":4.5773602000000002},{"$f":"'DF raw data'!W17","v":-5.2077298000000001}]
[8,{"$f":"'DF raw data'!B18","v":21.135019},{"$f":"'DF raw data'!C18","v":-24.269919999999999},{"$f":"'DF raw data'!G18","v":28.608720999999999},{"$f":"'DF raw data'!H18","v":-21.695129000000001},{"$f":"'DF raw data'!L18","v":21.212129999999998},{"$f":"'DF raw data'!M18","v":-17.268459},{"$f":"'DF raw data'!Q18","v":10.065939999999999},{"$f":"'DF raw data'!R18","v":-15.859030000000001},{"$f":"'DF raw data'!V18","v":8.2519798000000009},{"$f":"'DF raw data'!W18","v":-15.176769999999999}]
[10,{"$f":"'DF raw data'!B19","v":38.212090000000003},{"$f":"'DF raw data'!C19","v":-25.317221},{"$f":"'DF raw data'!G19","v":25.916401},{"$f":"'DF raw data'!H19","v":-22.457599999999999},{"$f":"'DF raw data'!L19","v":21.681761000000002},{"$f":"'DF raw data'!M19","v":-12.551069999999999},{"$f":"'DF raw data'!Q19","v":21.202840999999999},{"$f":"'DF raw data'!R19","v":-27.665849999999999},{"$f":"'DF raw data'!V19","v":1.2774399999999999},{"$f":"'DF raw data'!W19","v":-7.4210099999999999}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":531},null,{"$f":"'DF raw data'!G4","v":532},null,{"$f":"'DF raw data'!L4","v":533},null,{"$f":"'DF raw data'!Q4","v":534},null,{"$f":"'DF raw data'!V4","v":535}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":9.8369067333333344},{"$f":"AVERAGE(C11:C13)","v":-8.9408199966666668},15.369519666666667,-10.1102367,14.393370366666666,-10.026013333333333,12.518829933333334,-13.846360333333335,6.2868266000000004,-4.8008766999999999]
[]
[]
[]
[]
[]
[]
[]
[]
[null,null,null,null,null,null,null,null,null,null,null,null,null," "]

{"sheet":"2nd Int","index":1,"n_rows":21,"n_cols":19,"rows_kept":21,"cut":false,"hidden":false,"merged":[]}
["SXR 2nd Integral SW measurements"]
["Background has been subtracted",null,null,null,null,null,null,null,null,null,null,null,"Center values",null,"Average of 3 points x = -2 to x = 2"]
["y pos",0,"mm"]
["Gap",10,"mm",15,"mm",20,"mm",30,"mm",100,"mm",null,"Gap","Jx [muTm^2]","Jy [muTm^2]",null,"Gap interval","Jx [muTm^2]","Jy [muTm^2]"]
["Encoder",{"$f":"'DF raw data'!B7","v":386307},null,{"$f":"'DF raw data'!G7","v":486954},null,{"$f":"'DF raw data'!L7","v":587324},null,{"$f":"'DF raw data'!Q7","v":787626},null,{"$f":"'DF raw data'!V7","v":2187630},null,null,{"$f":"B4","v":10},{"$f":"B21","v":27.830695466666665},{"$f":"C21","v":9.7923893333333325},null,"10-30",{"$f":"MAX(N5:N8)-MIN(N5:N8)","v":7.3170128000000041},{"$f":"MAX(O5:O8)-MIN(O5:O8)","v":26.2701563}]
["x pos [mm]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]","Jx [muTm^2]","Jy [muTm^2]",null,{"$f":"D4","v":15},{"$f":"D21","v":27.922864800000003},{"$f":"E21","v":-6.0998631666666681},null,"10-100",{"$f":"MAX(N5:N9)-MIN(N5:N9)","v":18.642986533666672},{"$f":"MAX(O5:O9)-MIN(O5:O9)","v":27.41309893}]
[-10,{"$f":"'DF raw data'!D9/100","v":27.396677},{"$f":"'DF raw data'!E9/100","v":-4.5253900999999999E-2},{"$f":"'DF raw data'!I9/100","v":-32.289023},{"$f":"'DF raw data'!J9/100","v":8.7984960999999995},{"$f":"'DF raw data'!N9/100","v":19.742146000000002},{"$f":"'DF raw data'!O9/100","v":-37.907363000000004},{"$f":"'DF raw data'!S9/100","v":27.917663999999998},{"$f":"'DF raw data'!T9/100","v":-10.061289},{"$f":"'DF raw data'!X9/100","v":-6.5538116000000004},{"$f":"'DF raw data'!Y9/100","v":12.294267999999999},null,{"$f":"F4","v":20},{"$f":"F21","v":24.442460999999998},{"$f":"G21","v":-4.1981673433333331}]
[-8,{"$f":"'DF raw data'!D10/100","v":11.613325999999999},{"$f":"'DF raw data'!E10/100","v":-44.499257999999998},{"$f":"'DF raw data'!I10/100","v":11.958725999999999},{"$f":"'DF raw data'!J10/100","v":-5.4346190999999999},{"$f":"'DF raw data'!N10/100","v":34.922663999999997},{"$f":"'DF raw data'!O10/100","v":-21.862910000000003},{"$f":"'DF raw data'!S10/100","v":4.7389114000000001},{"$f":"'DF raw data'!T10/100","v":-18.550370999999998},{"$f":"'DF raw data'!X10/100","v":19.605383},{"$f":"'DF raw data'!Y10/100","v":-23.724395000000001},null,{"$f":"H4","v":30},{"$f":"H21","v":20.605851999999999},{"$f":"I21","v":-16.477766966666668}]
[-6,{"$f":"'DF raw data'!D11/100","v":21.505675999999998},{"$f":"'DF raw data'!E11/100","v":12.767988000000001},{"$f":"'DF raw data'!I11/100","v":27.648312999999998},{"$f":"'DF raw data'!J11/100","v":17.340322},{"$f":"'DF raw data'!N11/100","v":26.087439},{"$f":"'DF raw data'!O11/100","v":-25.314492000000001},{"$f":"'DF raw data'!S11/100","v":24.861430999999996},{"$f":"'DF raw data'!T11/100","v":-36.351835999999999},{"$f":"'DF raw data'!X11/100","v":-1.3584453999999999},{"$f":"'DF raw data'!Y11/100","v":-47.090781},null,{"$f":"J4","v":100},{"$f":"J21","v":9.2798782663333323},{"$f":"K21","v":10.935331963333333}]
[-4,{"$f":"'DF raw data'!D12/100","v":-13.488284},{"$f":"'DF raw data'!E12/100","v":21.243456999999999},{"$f":"'DF raw data'!I12/100","v":12.933523000000001},{"$f":"'DF raw data'!J12/100","v":-42.323379000000003},{"$f":"'DF raw data'!N12/100","v":22.552389999999999},{"$f":"'DF raw data'!O12/100","v":-34.342246000000003},{"$f":"'DF raw data'!S12/100","v":19.959702},{"$f":"'DF raw data'!T12/100","v":-13.073438000000001},{"$f":"'DF raw data'!X12/100","v":14.254073},{"$f":"'DF raw data'!Y12/100","v":-18.581503999999999}]
[-2,{"$f":"'DF raw data'!D13/100","v":47.984804999999994},{"$f":"'DF raw data'!E13/100","v":31.883593999999999},{"$f":"'DF raw data'!I13/100","v":8.0112994000000004},{"$f":"'DF raw data'!J13/100","v":5.6594628999999994},{"$f":"'DF raw data'!N13/100","v":20.490590999999998},{"$f":"'DF raw data'!O13/100","v":-6.2985253999999999},{"$f":"'DF raw data'!S13/100","v":29.362251000000001},{"$f":"'DF raw data'!T13/100","v":-9.1565039000000006},{"$f":"'DF raw data'!X13/100","v":19.604755999999998},{"$f":"'DF raw data'!Y13/100","v":3.9935547000000002},null,"Accept criteria:"]
[0,{"$f":"'DF raw data'!D14/100","v":2.7490633999999998},{"$f":"'DF raw data'!E14/100","v":-17.131240000000002},{"$f":"'DF raw data'!I14/100","v":65.204136000000005},{"$f":"'DF raw data'!J14/100","v":-3.8299414000000001},{"$f":"'DF raw data'!N14/100","v":12.861955999999999},{"$f":"'DF raw data'!O14/100","v":-7.1233496000000001},{"$f":"'DF raw data'!S14/100","v":12.21022},{"$f":"'DF raw data'!T14/100","v":-18.56335},{"$f":"'DF raw data'!X14/100","v":8.2792583999999998},{"$f":"'DF raw data'!Y14/100","v":0.42761718999999998},null,"Values within: +/-",42.5,42.5]
[2,{"$f":"'DF raw data'!D15/100","v":32.758217999999999},{"$f":"'DF raw data'!E15/100","v":14.624813999999999},{"$f":"'DF raw data'!I15/100","v":10.553159000000001},{"$f":"'DF raw data'!J15/100","v":-20.129111000000002},{"$f":"'DF raw data'!N15/100","v":39.974836000000003},{"$f":"'DF raw data'!O15/100","v":0.82737296999999999},{"$f":"'DF raw data'!S15/100","v":20.245085},{"$f":"'DF raw data'!T15/100","v":-21.713447000000002},{"$f":"'DF raw data'!X15/100","v":-4.4379600999999998E-2},{"$f":"'DF raw data'!Y15/100","v":28.384823999999998},null,"Test result",{"$f":"IF(MAX(N$5:N$9)<N$12,IF(MIN(N$5:N$9)>-1*N12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"},{"$f":"IF(MAX(O$5:O$9)<O$12,IF(MIN(O$5:O$9)>-1*O12,\"Pass\",\"Fail\"),\"Fail\")","v":"Pass"}]
[4,{"$f":"'DF raw data'!D16/100","v":25.287424000000001},{"$f":"'DF raw data'!E16/100","v":36.060693000000001},{"$f":"'DF raw data'!I16/100","v":40.939526000000001},{"$f":"'DF raw data'!J16/100","v":-22.605302999999999},{"$f":"'DF raw data'!N16/100","v":18.349036000000002},{"$f":"'DF raw data'!O16/100","v":-40.436318},{"$f":"'DF raw data'!S16/100","v":22.535212000000001},{"$f":"'DF raw data'!T16/100","v":3.230791},{"$f":"'DF raw data'!X16/100","v":5.7660248000000003},{"$f":"'DF raw data'!Y16/100","v":-24.603975000000002}]
[6,{"$f":"'DF raw data'!D17/100","v":-0.78550346000000004},{"$f":"'DF raw data'!E17/100","v":25.218036999999999},{"$f":"'DF raw data'!I17/100","v":29.561347999999999},{"$f":"'DF raw data'!J17/100","v":-1.0545800999999999},{"$f":"'DF raw data'!N17/100","v":12.456656000000001},{"$f":"'DF raw data'!O17/100","v":-9.0098438000000005},{"$f":"'DF raw data'!S17/100","v":16.487074},{"$f":"'DF raw data'!T17/100","v":16.552714999999999},{"$f":"'DF raw data'!X17/100","v":-3.6596378000000001},{"$f":"'DF raw data'!Y17/100","v":14.035195000000002}]
[8,{"$f":"'DF raw data'!D18/100","v":31.161316000000003},{"$f":"'DF raw data'!E18/100","v":-24.334229000000001},{"$f":"'DF raw data'!I18/100","v":42.379092},{"$f":"'DF raw data'!J18/100","v":-20.712197},{"$f":"'DF raw data'!N18/100","v":39.198818000000003},{"$f":"'DF raw data'!O18/100","v":-11.078173999999999},{"$f":"'DF raw data'!S18/100","v":48.751356999999999},{"$f":"'DF raw data'!T18/100","v":8.8626757999999999},{"$f":"'DF raw data'!X18/100","v":18.324617},{"$f":"'DF raw data'!Y18/100","v":-11.832861000000001}]
[10,{"$f":"'DF raw data'!D19/100","v":60.829883000000002},{"$f":"'DF raw data'!E19/100","v":-25.002627000000004},{"$f":"'DF raw data'!I19/100","v":30.461995000000002},{"$f":"'DF raw data'!J19/100","v":-19.806885000000001},{"$f":"'DF raw data'!N19/100","v":40.976698999999996},{"$f":"'DF raw data'!O19/100","v":-5.0360059000000001},{"$f":"'DF raw data'!S19/100","v":6.6188007000000004},{"$f":"'DF raw data'!T19/100","v":-60.892002000000005},{"$f":"'DF raw data'!X19/100","v":-0.13209650000000001},{"$f":"'DF raw data'!Y19/100","v":-9.0320409999999995}]
[]
["DF indice",{"$f":"'DF raw data'!B4","v":531},null,{"$f":"'DF raw data'!G4","v":532},null,{"$f":"'DF raw data'!L4","v":533},null,{"$f":"'DF raw data'!Q4","v":534},null,{"$f":"'DF raw data'!V4","v":535}]
[]
["Average",{"$f":"AVERAGE(B11:B13)","v":27.830695466666665},{"$f":"AVERAGE(C11:C13)","v":9.7923893333333325},27.922864800000003,-6.0998631666666681,24.442460999999998,-4.1981673433333331,20.605851999999999,-16.477766966666668,9.2798782663333323,10.935331963333333]

{"sheet":"DF raw data","index":2,"n_rows":19,"n_cols":25,"rows_kept":19,"cut":false,"hidden":false,"merged":[]}
["Comment:","2nd Int scan, PS 16351, 10 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16351, 15 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16351, 20 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16351, 30 mm gap",null,null,null,"Comment:","2nd Int scan, PS 16351, 100 mm gap"]
["Date:","8. jun 2017","15:44:02",null,null,"Date:","8. jun 2017","15:58:21",null,null,"Date:","8. jun 2017","16:13:44",null,null,"Date:","8. jun 2017","16:29:17",null,null,"Date:","8. jun 2017","16:41:55"]
["Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0,null,null,null,"Z-Pos:",0]
["Indice:",531,null,null,null,"Indice:",532,null,null,null,"Indice:",533,null,null,null,"Indice:",534,null,null,null,"Indice:",535]
["Background Indice:",531,null,null,null,"Background Indice:",532,null,null,null,"Background Indice:",533,null,null,null,"Background Indice:",534,null,null,null,"Background Indice:",535]
["Temperature:",21.35,null,null,null,"Temperature:",21.52,null,null,null,"Temperature:",21.45,null,null,null,"Temperature:",21.26,null,null,null,"Temperature:",21.49]
["Encoder:",386307,null,null,null,"Encoder:",486954,null,null,null,"Encoder:",587324,null,null,null,"Encoder:",787626,null,null,null,"Encoder:",2187630]
["X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb","X","IXb","IZb","JXb","JZb"]
[-1,8.1423702000000002,-13.99343,2739.6677,-4.5253901000000001,-1,-7.6044698000000004,-6.3994799000000002,-3228.9023000000002,879.84960999999998,-1,8.4387597999999997,-22.915371,1974.2146,-3790.7363,-1,14.13165,-14.360060000000001,2791.7664,-1006.1289,-1,-6.2310299999999996,-3.3148700999999998,-655.38116000000002,1229.4268]
[-0.80000000999999998,0.19045000000000001,-7.2811399000000003,1161.3326,-4449.9258,-0.80000000999999998,9.1924495999999998,-5.6423101000000004,1195.8725999999999,-543.46190999999999,-0.80000000999999998,18.319299999999998,-8.4756899000000008,3492.2664,-2186.2910000000002,-0.80000000999999998,0.40864,-8.1246004000000003,473.89114000000001,-1855.0371,-0.80000000999999998,10.92963,-21.325710000000001,1960.5382999999999,-2372.4395]
[-0.60000001999999997,-0.29017000999999998,-1.9674700000000001,2150.5675999999999,1276.7988,-0.60000001999999997,15.50127,-9.7119502999999998,2764.8312999999998,1734.0322000000001,-0.60000001999999997,11.019970000000001,-16.367719999999998,2608.7438999999999,-2531.4492,-0.60000001999999997,12.43378,-16.908809999999999,2486.1430999999998,-3635.1835999999998,-0.60000001999999997,3.9573100000000001,-31.089399,-135.84453999999999,-4709.0780999999997]
[-0.40000001000000002,-8.7370596000000003,-1.2584599999999999,-1348.8284000000001,2124.3456999999999,-0.40000001000000002,2.81758,-21.298189000000001,1293.3523,-4232.3379000000004,-0.40000001000000002,10.430059999999999,-19.436501,2255.239,-3434.2246,-0.40000001000000002,11.77398,-11.913119999999999,1995.9702,-1307.3438000000001,-0.40000001000000002,-2.7963499999999999,-16.779240000000001,1425.4073000000001,-1858.1504]
[-0.2,8.1936903000000001,0.38256001000000001,4798.4804999999997,3188.3593999999998,-0.2,12.85463,-4.0197000999999997,801.12994000000003,565.94628999999998,-0.2,8.4867600999999997,-11.379960000000001,2049.0590999999999,-629.85253999999998,-0.2,15.450760000000001,-11.29477,2936.2251000000001,-915.65039000000002,-0.2,3.4091898999999999,-3.1791999,1960.4756,399.35547000000003]
[0,7.0873898999999998,-16.867629999999998,274.90634,-1713.124,0,18.677019000000001,-13.08187,6520.4135999999999,-382.99414000000002,0,15.30214,-7.4789399999999997,1286.1956,-712.33496000000002,0,7.9587998000000004,-10.727690000000001,1221.0219999999999,-1856.335,0,10.526160000000001,-5.6685800999999998,827.92583999999999,42.761718999999999]
[0.2,14.22964,-10.337389999999999,3275.8218000000002,1462.4813999999999,0.2,14.57691,-13.229139999999999,1055.3159000000001,-2012.9111,0.2,19.391210999999998,-11.219139999999999,3997.4836,82.737296999999998,0.2,14.146929999999999,-19.516621000000001,2024.5084999999999,-2171.3447000000001,0.2,4.9251298999999999,-5.5548501000000003,-4.4379600999999997,2838.4823999999999]
[0.40000001000000002,19.894010999999999,-1.4725999999999999,2528.7424000000001,3606.0693000000001,0.40000001000000002,17.713301000000001,-17.839808999999999,4093.9526000000001,-2260.5302999999999,0.40000001000000002,12.50811,-22.097460000000002,1834.9036000000001,-4043.6318000000001,0.40000001000000002,10.96805,-18.441680999999999,2253.5212000000001,323.07909999999998,0.40000001000000002,3.069,-17.651520000000001,576.60248000000001,-2460.3975]
[0.60000001999999997,25.126418999999999,-9.5381899000000008,-78.550346000000005,2521.8036999999999,0.60000001999999997,20.553128999999998,-14.90456,2956.1347999999998,-105.45801,0.60000001999999997,9.1611499999999992,-13.01342,1245.6656,-900.98437999999999,0.60000001999999997,2.8752,-5.4107098999999996,1648.7074,1655.2715000000001,0.60000001999999997,4.5773602000000002,-5.2077298000000001,-365.96377999999999,1403.5195000000001]
[0.80000000999999998,21.135019,-24.269919999999999,3116.1316000000002,-2433.4229,0.80000000999999998,28.608720999999999,-21.695129000000001,4237.9092000000001,-2071.2197000000001,0.80000000999999998,21.212129999999998,-17.268459,3919.8818000000001,-1107.8173999999999,0.80000000999999998,10.065939999999999,-15.859030000000001,4875.1356999999998,886.26757999999995,0.80000000999999998,8.2519798000000009,-15.176769999999999,1832.4617000000001,-1183.2861]
[1,38.212090000000003,-25.317221,6082.9883,-2500.2627000000002,1,25.916401,-22.457599999999999,3046.1995000000002,-1980.6885,1,21.681761000000002,-12.551069999999999,4097.6698999999999,-503.60059000000001,1,21.202840999999999,-27.665849999999999,661.88007000000005,-6089.2002000000002,1,1.2774399999999999,-7.4210099999999999,-13.20965,-903.20410000000004]